## Slide 1
Diamond and converter based Timepix3 detectors
for fast and thermal neutron measurements in nuclear fusion
Antonella Tamburrino1,2,3 & G. Claps2,3
1DIAEE - Dipartimento di Ingegneria Astronautica, Elettrica ed Energetica, Sapienza Università di Roma, Piazzale Aldo Moro 5, 00185 Roma, Italia
2INFN - Laboratori Nazionali di Frascati (LNF), via E. Fermi 40, 00044 Frascati, Italia
3ENEA (FSN-PLAS-PAX) - C. R. Frascati, via E. Fermi 45, 00044 Frascati, Italia
Meeting nTOF, Trieste, 9 – 10 novembre 2023

## Slide 2
SCIENTIFIC BACKGROUND: Fast neutron detection in nuclear fusion
NEUTRON DIAGNOSTIC
Plasma physics, machine protection and control issues
Total emission strength, spatial emissivity and energy spectra
MAIN PARAMETERS PROVIDED BY NEUTRON DIAGNOSTIC:
Fusion reaction rates, fuel ion temperature and density
Beam plasma interaction
Neutron load on components (vacuum chamber and divertor)
REQUIREMENTS OF A NEUTRON DETECTION SYSTEM:
Diagnostic apparatus (collimation and shielding)
Resistance to harsh environment conditions
Detector requirements
THE PROBLEM OF BACKGROUND CONTRIBUTION:
SCATTERED NEUTRON + GAMMA
It needs to be understood and controlled!
Neutron discrimination from gammas rays
Good discrimination of the direct component
Time resolution
Rate capability

## Slide 3
Detector 2D
Active area: 14 x 14 µm2
Pixel arrangement: 256 x 256 pixels
Pixel size: 55 x 55 µm2
Minimum Timing resolution: 1.6 ns
ACQUISITION MODES OF TPX3
TIMEPIX3 chip
TIMEPIX3 chip
  KATHERINE MODULE + TIMPEPIX3
Pixel by pixel, TPX3 provides a simultaneous measure of :
EVEN COUNTING
TIME (ToA mode)
CHARGE (ToT mode)
CHARACTERISTIC:

## Slide 4
Diamondpix
10 x 10 mm2
DIAMOND PROPERTIES
It exploits specific reaction channels with charge products
DETECTOR 2D
SENSOR: CVD diamond
active area: 10 x 10 mm2
500 μm thick CVD diamond
300 nm thick Au layer
Bias: 300 V
DETECTOR LAYOUT
12C(n, n)12C
12C(n, α)9Be
12C(n, α)2α
Main reactions in 12C :
The new Diamond detector “DIAMONDPIX”
PROPERTY | SILICON | DIAMOND | DIAMOND
Material | MCz, FZ, epi | Polycrystalline | Single crystal
Displacement [eV] | 13-20 | 43 | 43
Energy gap [eV] | 1.12 | 5.5 | 5.5
Ebreakdown [V/cm] | 3×106 | 107 | 107
µe [cm2/Vs] | 1450 | 1800 | >1800
µh [cm2/Vs] | 450 | 1200 | >1200
Saturate drift velocity [cm/s] | 0.8×107 | 2.2×107 | 2.2×107
Atomic number Z | 14 | 6 | 6
Relative dielectric constant (εr) | 11.9 | 5.7 | 5.7
e-h energy [eV] | 3.6 | 13.0 | 13.0
Density [g/cm3] | 2.33 | 3.51 | 3.51
e-h/µm for mips | ~ 80 | 36 | 36
Max ccd [µm] | >500 | 300 | 800
Thermal conductivity [W/(mK)] | 150 | ~1900 | ~1900
14 x 14 mm2
10 x 10 mm2

## Slide 5
Diamondpix for neutron Time-of-Flight measurements: n_ToF (CERN, 2022)
EXPERIMENTAL SET-UP
DIAMONDPIX
ACQUISITION PROGRAM
BEAM DUMP
DIAMONDPIX
nTOF facility
TPX3 has been checked with the katherine module by MMTrack program
DATA-DRIVEN MODE
ACQUISITION MODE: TOA&TOT
IKRUM: 5
DURATION: 150 ms neutron energy range (1 GeV – 10 meV)
  A single particle interaction turns on a group of pixels (cluster) and, for each pixel, the ToT acquisition provides a measure of a fraction of the released charge.
1000 triggers acquisition (no cut image)

## Slide 6
Track analysis
gamma background discrimination
Gamma rejection has been validated through a set of laboratory gamma sources.
10 triggers acquisition
SOME CLUSTER PARAMETERS:
Cluster Size (CS): the sum of all pixels forming the cluster
ToT volume (ToTv): sum of all pixel ToT values of the cluster
Roundness (Rnd):

## Slide 7
CS > 3
time @ gamma flash
10509 ns
ToF neutron spectrum measured by Diamondpix
The measured N_TOF spectrum has been obtained through the minimum of ToA times in each cluster and correcting the delay due to time-walk
ToF spectrum reconstruction
The ToA measures highlights clearly the gamma flash peak, used as reference time to obtain the energy spectrum.
Some tof peaks are observed in correspondence of cross section resonances and validate the spectrum reconstruction
ToA measurements

## Slide 8
Neutron tracks discrimination and efficiency estimation at 2.5 and 14 MeV (FNG facility)
A neutron flux scan (from 6.0 × 108 to 1.6 × 1010 n/s) on FNG allowed an efficiency estimation obtaining a value of 2.5 ± 0.1 ‰
A neutron flux scan (from 2.1 × 108 to 2.5 × 108 n/s) on FNG allowed an efficiency estimation obtaining a value of 6.8 ± 0.5 ‰

## Slide 9
Energy calibration of Diamondpix
Depending to the neutron energy, the mean CS values of all cluster follow a linear trend tha can be used to define the CS cut on the plane CS/Rnd.
Study of Cluster Size and Roundness as a function of neutron energy
Energy calibration of Diamondpix with low neutron energies
Diamondpix has been calibrated exploiting the Carbon recoil for neutron energies lower than 5 MeV.

## Slide 10
Calibration curve
The calibration curve has been applied to evaluate the deposited energy on charge responces of Diamondpix

## Slide 11
Comparison of the Diamondpix charge responce at 2.5 and 14 MeV between FNG and n_TOF

## Slide 12
Neutrons - 12C REACTIONS: |  | 
Reactions | Q-value (MeV) | Threshold energy (MeV)
12C(n, n)12C | 0.0 | 0.0
12C(n, n’ γ)12C | 4.44 | 0.0
12C(n, γ)13C | 4.95 | 0.0
12C(n, α)9Be | -5.70 | 6.18
12C(n, n’ 2α)4He | -7.27 | 7.89
12C(n, n’ α)8Be | -7.37 | 7.99
12C(n, 2α)5He | -8.16 | 8.85
12C(n, p)12B | -12.59 | 13.64
12C(n,d)11B | -13.73 | 14.89
Analysis of charge profiles and comparison with Monte Carlo simulations

## Slide 13
Thermal neutron detection by means of Timepix3
Thermal neutron detection in Radiation Protection & Nuclear reactors
TPX3 with 6LiF CONVERTER
material | converter | Natural Isotopic Abundance | Neutron reactions | Reaction products energetic
LiF | 6Li enrichment 95 % | 6Li (7.5 %) 7Li (92.5 %) | 6L(n, α)3H | α (2.05 MeV) 3H (2.73 MeV)
  6LiF converter enriched with 95% 6Li
  Film deposited on a Silicon substrate
  Thicknesses : 330 μm Si, 5 μm 6LiF
6LiF CONVERTER

## Slide 14
HOTNES
(HOmogeneous Thermal NEutron Source)
FACILITY
(ENEA FRASCATI)
L’istogramma 2D del ToT volume in funzione della Cluster Size mostra due distribuzioni: una dovuta alla presenza simultanea di tracce da alfa e tritoni e l’altra relativa alla presenza di sole tracce da tritoni.
  Source: 241Am – B
  Flux(Ref.point): 763  11 n/(cm2 s)
Experimental measurements at HOTNES facility & results
Simultaneous presence of alpha and triton tracks
Presence of only triton tracks
The 1D histogram of ToTv highlights the presence of the two energy contributions.

## Slide 15
ToT  1050
E (3H)  2.3 MeV
ToT energy calibration procedure & results
Energy response at different distances of the alphas source

## Slide 16
Estimated efficiency
(Tritons)
Estimated efficiency (Tritons & alphas)
nToF (neutron Time of Flight) FACILITY (CERN, Ginevra)
L’istogramma 2D del ToT volume in funzione della Cluster Size mostra due distribuzioni: una dovuta alla presenza simultanea di tracce da alfa e tritoni e l’altra relativa alla presenza di sole tracce da tritoni.
Energy range: 5 meV – 0.5 eV
10 frame
1000 frame
Alpha and Triton tracks
CdTe TPX3
CS > 5
Rnd > 70
X cent: 10-100
Y cent: 10-100
ToT: 0 - 4500
CS > 5
Rnd > 70
X cent: 10-100
Y cent: 10-100
ToT > 450
  Neutron Flux
Estimated measured efficiency & measurements at n_TOF facility
WORK IN PROGRESS…
Time-of-flight measurements, after appropriately selecting the reference times for thermal neutrons:

## Slide 17
Thanks for attention
Dedicated to Fabrizio Murtas