## Slide 1
Piergiorgio Cerello
INFN, Sezione di Torino, Torino, Italy
Optimising Hadrontherapy: challenges and prospects
Treatment planning
- Nuclear fragmentation
Treatment delivery
- Ion gantries, minibeams
Treatment monitoring
- Beam monitoring & Range verification, radioactive beams
New approaches
- Boron Neutron Capture Therapy
- FLASH therapy

## Slide 2
reduced toxicity​
lower secondary malignancy risk​
increased tumor local control
Particle Therapy: the concept
Compared to standard
radiotherapy
Radiotherapy
Particle therapy
Dose maps

## Slide 3
Well established
Very effective
At the end of 2023:
Clinical Centres
  ~ 150
Treated Patients:
  ~ 350000 (protons)
  ~ 56000 (carbon)
  ~ 3500 (helium)
Particle Therapy: Worldwide statistics
https://ptcog.site
Creative commons 3.0

## Slide 4
Particle Therapy: the standard
https://www.oncolink.org/cancer-treatment/radiation/types-of-radiation-therapy/proton-therapy/overviews-of-proton-therapy/proton-therapy-behind-the-scenes
Treatment planning
  Nuclear fragmentation
Treatment delivery
  Ion Gantries
  Minibeams
Treatment monitoring
  Beam monitoring
  Range verification
  Radioactive beams

## Slide 5
Optimising Hadrontherapy: challenges and prospects
Treatment planning
- Nuclear fragmentation

## Slide 6
Piergiorgio Cerello
Text
Particle therapy: E < 500 MeV/u
the contribution of tissue fragmentation in the entrance channel was not measured (yet)
12C, 16O
p
12C
Inverse kinematics
FragmentatiOn Of Target (FOOT)

## Slide 7
Piergiorgio Cerello
Large acceptance
Low statistics
Emulsion Spectrometer
FragmentatiOn Of Target (FOOT)
Electronic Spectrometer
Forward acceptance (~ 11 degrees)
High statistics
Forthcoming data takings (2025)
16O beam at GSI
12C beam at CNAO

## Slide 8
Optimising Hadrontherapy: challenges and prospects
Treatment delivery
- Ion gantries, minibeams

## Slide 9
Piergiorgio Cerello
Courtesy of Elisa Fiorina, Marco Prioli
Study and development of a next generation ion gantry of “only” 50 tons for non-coplanar ion irradiation
Carbon ions up to 430 MeV/u
14 m long, ~50 tons weight
4 T curved superconducting dipoles
Superconducting spool piece quadrupoles
Downstream scanning magnet system
Dose Delivery and Range Verification Systems for adaptive cancer treatments
Superconducting Ion Gantry
CERN, CNAO, INFN, MedAustron

## Slide 10
Minibeams
P. Lansonneur et al, Scientific Reports (2020) 10:7025 https://doi.org/10.1038/s41598-020-63975-9
Eclipse
4mm ctc
6mm ctc
MC
Dose profile - at 3 cm depth - in the tumour volume
Brass collimator provides
beam intensity modulation
with high, thin peaks and
valleys (ctc: centre-to-centre)
Proton minibeam radiation therapy (pMBRT): a novel dose delivery method based on spatial dose fractionation

## Slide 11
Simulation of Dose comparison for glioma (left) and menigioma (right)
Minibeams
P. Lansonneur et al, Scientific Reports (2020) 10:7025 https://doi.org/10.1038/s41598-020-63975-9
pMBRT provided:
- reduced side effects
- superior tumour control
in high-grade glioma-bearing rats compared to standard irradiation
Next step: clinical studies

## Slide 12
Optimising Hadrontherapy: challenges and prospects
Treatment monitoring
- Beam monitoring & Range verification, radioactive beams

## Slide 13
Beam monitoring
Where (x,y) is the beam fired?
What intensity?
Is it compliant with the treatment plan?

## Slide 14
Piergiorgio Cerello
Courtesy of Roberto Sacchi
Collection times ~ 100 µs
Sensitivity ~ 104 protons
Time resolution ~ no/poor
IONIZATION CHAMBERS
Not suitable for fast scanning modalities and timing applications
SOLID STATE DETECTORS
~ ns
single protons
< 100 ps
proton counting
timing applications
Main issues at
φ = 1010 p/cm2s
Signal pile-up
→ fast sensors & readout
→ segmentation
Radiation tolerance
→ manufacturing strategies
→ damage compensation
Detector area
→ at least 20x20 cm2
Beam monitoring

## Slide 15
Piergiorgio Cerello
30 ns
Single proton identification and timing
Carbon ions, 399 MeV/u
Courtesy of Roberto Sacchi
Beam monitoring with LGAD silicon detectors
ABACUS ASIC (24 ch)
2.7×2.7 cm2 active area (144 strips)

## Slide 16
Range uncertainty estimation for robust treatment plans
2.5-3% + 1-3 mm
If range well under control,
only
Particle Range Uncertainty

## Slide 17
charged particles
𝛄 511 keV
𝛄 511 keV
neutrons
𝛄 prompt
In vivo range verification
What depth (z) is the beam reaching?
Is it compliant with the treatment plan?

## Slide 18
M.G. Bisogni at al, JMI 2017; Ferrero et al., Sci.Rep 2018; Traini et al., Phys. Med., 2017
The INSIDE bimodal system

## Slide 19
J Pawelke et al., Proceedings IBIBAM, 26.-29.09.2007, Heidelberg
Zhu X, Fakhri GE. Theranostics. 2013;3(10):731-740.
Protons
Carbon ions
Isotope | T1/2
11C | ~ 20 m
15O | ~ 2 m
8B | 770 ms
12N | 0.011 s
9C | 126.5 ms
13O | 8.58 ms
Positron emitters
Activity distribution and Range are correlated
Main contributions from 11C and 15O

## Slide 20
Pz6
Mettere I campi del pz 6
Forse meglio logY?
Coincidence event rate
1st field
270° IEC
2nd field
0° IEC
3rd field
180° IEC
In-beam PET data patterns
In-spill
vs.
Inter-spill

## Slide 21
Ferrero et al., Sci.Rep 2018; Fiorina et al., EJMP, 2018
2016: First in-vivo measurement

## Slide 22
ClinicalTrials.gov NCT03662373
ACC
4
3
2
1
SCC
chordoma
Chondrosarcoma
5
Meningioma
MMM
chordoma
Chondrosarcoma
ITAC
ACC
proton patients
carbon ion patients
GOALS
Evaluate the INSIDE sensitivity in a clinical environment
Test the INSIDE capability to spot interfractional morphological changes
Phase 1 (2019-2020):
9 proton pts + 10 carbon ion pts
Phase 2 (2024-2025):
10 proton pts + 10 carbon ion pts
5 monitored patients underwent interfractional morphological changes
2 patients needed replanning
The INSIDE clinical trial

## Slide 23
In-beam PET: proton treatments
in-beam PET image - 2 fx
in-beam PET image - 21 fx
Patient ID: 006P: female, 39y, affected by ACC, IMPT, CTV dose: 70 GyE in 35 fx (2.0 GyE/fx, 5 fx/week), two beams (15° - 175° IEC)
Fiorina et al., Fron. Phys., 2021; Moglioni et al., Front. Onc. 2022
Measure the range variation along the beam direction:
- BEV: Beam Eye View method
- Most Likely Shift method
planning CT
control CT - 20 fx

## Slide 24
C, 212 MeV/u
Arbitrary units
Penetration depth [mm]
Ferrero, Bersani et al., poster IEEE MIC-NSS 2022
CTV
In-beam PET: carbon treatments
~ 1/100 of proton treatments statistics!!!

## Slide 25
Piergiorgio Cerello
Courtesy of Alessio Sarti
8 planes composed of 2 orthogonally oriented layers of plastic scintillating fibres
SiPM read-out
Detection of charged secondary fragments emitted @ large angle wrt the beam direction
Dose profiler: carbon treatments

## Slide 26
Piergiorgio Cerello
3D emission map of fragments collected during the treatment delivery
Gamma test (9mm/10%) has been used for voxel to voxel comparison
beam
gamma map
ClinicalTrials.gov NCT03662373
Courtesy of Alessio Sarti
Dose profiler: carbon treatments

## Slide 27
Radioactive ion beams (i.e., 11C)
1/6 of a Siemens Biograph mCT clinical PET scanner with custom-modified detectors.
Two detector panels installed above and below the phantom each at a distance of 30 cm
RIB stops in the PMMA phantom
PET range uncertainty after delivering 1.6 x 107 ions in 4 spills (19.2 s total):
0.04 mm, 0.7mm, and 1.3 mm for 10C, 11C, and 12C
Kostyleva et al. Phys. Med. Biol. 2022
Courtesy of Marco Durante

## Slide 28
Magalhaes Martins, P. et al., Sci Rep 10, 7007 (2020)
Prompt photons
Several techniques to detect prompt photons:
Collimated gamma camera
Compton Camera
Prompt Gamma Timing

## Slide 29
TOF = t_stop(𝛾) – t0 (beam particle) = t(beam in air) + t(beam in target) + t(𝛾)
Beam monitor
F. Pennazio et al., PMB 67 (2022) 065005, doi: 10.1088/1361-6560/ac5765
Prompt Gamma Timing

## Slide 30
Piergiorgio Cerello
Text
Courtesy of Veronica Ferrero
Measurement of the EneRgy Loss for IN-vivo Optimization in particle therapy
Multi-detector approach relying on the SER-PGT reconstruction algorithm
(F. Pennazio et al., Phys Med Biol 2022)
The Bragg curve (in-vivo!!!)
t (ns)
Prompt Gamma Timing: Stopping power!

## Slide 31
Piergiorgio Cerello
Courtesy of Francesco Pennazio
First I3PET proton beam test (2021):
primary-secondary radiation coincidences identified!!!
- Trigger on PET detector
- Joint DAQ
I3PET module
signal
UFSD
signal
5 ns
beam
MoVeIT UFSD
I3PET modules
Primary beam
Secondary radiation
monitoring of Positron Emitters & Prompt Photons with the same PET detector
Prompt Gamma Timing

## Slide 32
Piergiorgio Cerello
Text
UFSD 8 strips
Beam nozzle
t (ns)
t (ns)
Data vs. simulation
Data
with(out) hole
Prompt Gamma Timing
227 MeV p
on PMMA
107 pps
LaBr3 – UFSD time coincidences
t (ns)
Courtesy of Veronica Ferrero

## Slide 33
Optimising Hadrontherapy: challenges and prospects
New approaches
- Boron Neutron Capture Therapy
- FLASH therapy

## Slide 34: Boron Neutron Capture Therapy (BNCT)
Piergiorgio Cerello
Text
- neutron flux on 10B-enriched target tissue
- high LET* secondaries produced in 10B(n,α)7Li
- very short range (μm)
highly effective at cell level
*Linear Energy Transfer
Several facilities built / planned worldwide
Most advanced status:
- clinical trial (Japan)

## Slide 35
Photon detectors
Piergiorgio Cerello
in vivo BNCT dosimetry by single photon detection
accelerator-based neutron source
”therapeutic” dose dD(x,y,z) ~ nB10(x,y,z) . Φ(x,y,z)dV ~ dIγ(x,y,z)
Real time measurement of 10B reaction rate
multiple detectors = multiple projections
BNCT-SPECT or Compton Camera
voxel size < 1 cm3
statistical uncertainty < 10% @478 keV
10B reaction rate > 106 cm-3s -1
Challenges:
mechanical collimator effective @ 478 keV
intense n+γ background (2.2 MeV γ rays from 1H captures)
compact and portable system to adapt to patient’s position
Courtesy of Nicoletta Protti

## Slide 36: The FLASH effect
> 40 Gy/s for a total irradiation time < 100 ms
Instantaneous dose rate up to 109 Gy/s
Irradiation with ultra-high dose rate
28 Gy 31Gy 34 Gy
(Vozenin et al. 2019, Clin. Canc. Res. )
Decrease of the normal tissue response
Preservation of the tumor responses
CONV
FLASH
Courtesy of Alessio Sarti
(V. Favaudon et al. 2014, Sci. Transl. Med.)

## Slide 37: What if FLASH can help on tumours too?
First time it is demonstrated that FLASH can help on hypoxic tumours!
U-87 MG human GBM cells implanted in subcutaneous of female mice
20 Gy single fraction delivered with:
CONV (left)
FLASH (right)
radiation therapy
different oxygenation conditions (physioxia – hypoxia – hyperoxia)
Oriatron 6e, 6-MeV electron beam linear accelerator (LINAC)
Courtesy of Alessio Sarti
Leavitt RJ et al. Int J Radiat Oncol Biol Phys. 2024 Feb 20:S0360-3016(24)00320-1. doi: 10.1016/j.ijrobp.2024.02.015
Relative tumour volume
Relative tumour volume
Days after irradiation
Days after irradiation

## Slide 38: What is missing for clinical implementation?
… pretty much “everything” …
Proper explanation of the mechanism and modelling
Tools for the treatment planning
A technology suitable for delivering FLASH treatments meeting the clinical / dose conformity prescriptions for p, 𝛾 & e- machines
A dosimeter/beam monitor technology that can be routinely used in the clinical workflow
Courtesy of Alessio Sarti

## Slide 39
Beam delivery: ion gantries, minibeams
Beam monitoring: silicon detectors radiation resistance and size
Range Verification: radioactive beams, in-vivo Bragg peak measurement
Target Fragmentation: cross section measurements (p, C beams on C, O, Ca targets)
Boron Neutron Capture Therapy: from clinical trials to clinical implementation
FLASH therapy: «all of it» - but very promising
Piergiorgio Cerello
Text
Particle Therapy key developments

## Slide 40
Piergiorgio Cerello
Text
Unless things are much simpler…
… a sieve to modulate protons,
and a fan to remove neutrons…

## Slide 41
Beam delivery: ion gantries, minibeams
Beam monitoring: silicon detectors radiation resistance and size
Range Verification: radioactive beams, in-vivo Bragg peak measurement
Target Fragmentation: cross section measurements (p, C beams on C, O, Ca targets)
Boron Neutron Capture Therapy: from clinical trials to clinical implementation
FLASH therapy: «all of it» - but very promising
Piergiorgio Cerello
Text
Particle Therapy key developments