## Slide 1: Microstrip sensor production at FBK
Benno Margesin

## Slide 2: Introduction
Initial specifications
  Frontside electrodes with 20 µm minimum feature size
  Backside metallization with only coarse alignment requirements
  Metallization in aluminum or nickel
  Substrate: sapphire 110 µm and 150 µm
Possible challenge
  Transparent wafer vs automatic handling
Real challenge
  Substrate size 2” only manual processing and limits on available machines and processes

## Slide 3: First fabrication run DZ1
The device and the wafer layout was prepared with Tanner Tools™
A wafer handling jig was fabricated by UniPd used as wafer holder and shadow mask for the back metallization
Summary process flow
  Rinsing
  Evaporation of 300 nm aluminum
  Lithography of the microstrips
    Manual resist processing
    Litho with mask aligner
  Wet aluminum etching
  Ashing
  Evaporation of 200 nm aluminum on the back with shadow mask
  5 wafers processed, one without sintering, 3 with sintering and one lost

## Slide 4: First fabrication run DZ1 cont.
Processing issues
  Shift of the wafer during exposure. Alignment disc is not precise (mask aligner cannot align on the wafer itself)
  Rotation and shift in the jig for shadow evaporation
Outcome and problems
  Definition of the structures was ok as expected
  Main defects were shorts caused by particles on the wafer due to poor cleanliness of the substrates

## Slide 5: First fabrication run DZ1 cont.
Removal of shorts with FIB
  On a few wafers an attempt was made to remove the shorts using a Focused Ion Beam
  This was partially successful
    Hard to find the defects in this large dies
    Some shorts were very persistent

## Slide 6: 2nd Fabrication run DZ2
Summary process flow
    Piranha only clean
    Oxygen plasma clean
    Evaporation of 30/300 nm of Cr/Al
    Lithography
    Wet etching & Ashing
    Evaporation of 30/200 nm Ti/Al on the back with shadow mask
    Sintering at 200 °C 30 min
    1 wafer processed
Results
  Definition is Ok
  Improved adhesion
  Much reduced shorts due to better wafer cleaning
  Opens due to air bubbles in the photoresist

## Slide 7: Third fabrication run DZ3
Summary process flow
  First clean with air gun
  Piranha only clean followed
  Oxygen plasma clean
  Evaporation of 30/300 nm of Cr/Al
  Lithography
  Wet etching & Ashing
  Evaporation of 30/200 nm Ti/Al on the back with shadow mask
  Sintering at 200 °C 30 min
  2 wafer processed
Results
  Similar outcome as with DZ2
  Air gun cleaning removes more particles
  Issues with air bubbles in the photoresist still persisted

## Slide 8: Conclusions and Suggestions
Major issues encountered
  Centering of the mask on the 2” wafer is difficult
    This affects mainly the overlap with the back metal
  Initial cleanliness of the wafers affects the lithography generating shorts
  Bubbles in the resist (responsible for opens) can hardly be avoided due to the manual resist processing
In summary not being fully equipped for working on 2” substrates is a major handicap
  We can not fully exploit our capabilities
  The whole work is more cumbersome
The use of standard 6”(150 mm) substrates could substantially improve fabrication
  Better cleanliness
  Much improved resist processing
  Possibility of using dry etching
  Substantial higher throughput