## Slide 1
L. Palumbo
Attività Sapienza su acceleratori e PNRR: stato e prospettive
D. Alesini, G. Bisogni, F. Bosco, F. Cardelli, M. Carillol, E. Chiadroni, G. Cuttone, D. De Arcangelis, A. Di Gregorio, R. Di Raddo, L. Faillace, L. Ficcadenti, G. Franciosini, G. Franzini, A. Gallo, L. Giuliano, V. Lollo, M. Magi, M. Migliorati, A. Mostacci, P. Cirrone, L. Palumbo, V. Patera, L. Piersanti, S. Pioli, R. Remetti, A. Sarti, B. Spataro, G.Torrisi, A. Vannozzi

## Slide 2
Flash therapy is a revolutionary technique
in the perspective of cancer cure
it spares healthy tissues from the damage of the ionizing radiation maintaining the tumor control as efficient as in conventional radiotherapy.
to allow the implementation of the revolutionary FLASH therapy concept into actual clinical use, electron linear accelerators are required to deliver very high dose rate per pulse (> 106 Gy/s) in very short time of irradiation (100 ms).
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 3: Is the evidence robust?
First evidence on animals….. Much less damage to standard tissue for same dose release
34 Gy 31 Gy 28 Gy
Conventional therapy
Mouse skin treated with electron beam
DOSE MODIFYING FACTOR
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 4: Is the evidence robust?
The evidence seems robust, even if the exact features are yet to be explored. The first patients has been already treated!!!
Electron Beam 5.6 MeV, 15 Gy in 90 ms
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 5
TCP = Tumor Control Probability
NTCP = Normal Tissue Complication Probability
FLASH EFFECT
Probability
Probability
Dose
Dose
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 6
CHUV
Oriatron e6
Curie
EF
Stanford Varian clinac 21ex
Oxford
PisaEF
Rossendorf
ELBE
Anversa EF
ELECTRON FLASH LINACS
Lund
Elekta SL75
CLEAR
CERN
Roma C-b
Catania EF
Tsinghua CN
Atlanta Univ. USA
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 7: What does it mean in terms of dose delivery?
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 8
CONVENTIONAL VS FLASH
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 9
2020 Electron ELECTRON-FLASH 7 MeV – SIT Company
2023

## Slide 10
TOWARD VHEE LINACS
FOR DEEP TUMORS
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 11
JOINT
STUDY GROUP
SAPIENZA – INFN
2021
2022
2022
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 12: What range of energy ?
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 13: PDD Integrated Dose – 9 MeV, 80 MeV, 130 MeV
,
9 MeV
80 MeV
130 MeV
Integrated Relative Dose
1
cm
1
10
20
30
40
50
10
100%
Beam spot 3 cm
Never introduced such VHEE in clinical RT till now!
Percentage Depth Dose curve

## Slide 14: 1. VHEE BETTER THAT PHOTONS EVEN WITHOUT FLASH
2. VHEE COMPARABLE WITH PROTONS IN CASE OF FLASH
SIMULATIONS WITH MONTECARLO CODE SHOW THAT
(V. PATERA)
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 15
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 16: SCHEME FOR 100 MeV Linac
Standing wave Injector
Traveling wave
Frequency | 5.712 GHz
Beam Energy | 65 - 100 MeV
RF Repetition rate | 100 Hz
Current | 100 mA
C-band average accelerating gradient | 45 MV/m
RF pulse duration | 1.5 – 2.5 ms
In pulse dose rate | > 10^6 Gy/s
Average dose rate | > 100 Gy/s
Dose per pulse | >> 1 Gy

## Slide 17
* Inizialmente insieme a INFN (G. Cuttone)

## Slide 18
Asymmetric power splitter
Phase shifter
Attenuator
SW Injector
TW Structure
5.5 MW
5 ms
Pulse compressor 17.5 MW, 1.6 ms
15 MW
REALIZZATO PRESSO SIT, TRASFERITO ALLA SAPIENZA A FINE PROGRAMMA
OBIETTIVO INIZIALE
FLASH FUNDED WITH BUDGET 1.7 ME

## Slide 19
SIT ha ridotto la sua partecipazione
Budget Trasferito a Sapienza
Laboratorio FLASH presso SBAI
Lavori per adeguamento laboratorio da iniziare in autunno

## Slide 20
SCANDINOVA
5.5 MW
5 ms

## Slide 21: DESIGN AND PROTOTYPING OF SW injector
Electron source: DC thermionic gun at 15 kV.
Cavity type: standing wave, bi-periodic, on-axis coupling cells
Electromagnetic field: π/2 accelerating mode
Nose - cone geometry: high electric field at the noses zone
Bunching section for low-energy beam capture (> 45%)
Beam energy up to ≈ 10 MeV and 100 mA beam peak current
DC gun
Power in
Bunching section
Single cell parameters [mm]
D | g | L | Rb | Ro | α
38.25 | 15.6 | 24.4 | 3 | 6 | 30°
10 MeV
100 mA
Input
2023

## Slide 22
Input power
E field
H field
SW biperiodic 𝜋/2 structure prototype
Resulting profile of the on-axis accelerating electric field

## Slide 23: Mechanical drawing of the SW linac (SIT)
prototyping on going

## Slide 24
Asymmetric power splitter
Phase shifter
Attenuator
SW Injector
TW Structure
5.5 MW
5 ms
Pulse compressor 17.5 MW, 1.6 ms
15 MW

## Slide 25
DESIGN AND PROTOTYPE TW HIGH GRADIENT (2/3)𝜋
Single Cell design
Single cell RF parameters in mm
a | b | d | R0 | t | r1/r2
5 | 21.03 | 17.5 | 6 | 3 | 1.25
Shunt impedance R = 95 MW/m for r1/r2 = 1.25 and t=3
IRIS OPTIMIZATION

## Slide 26
Prototyping phase 1
Brazing alloy
Two pre-prototypes on 5-cells without couplers to test the brazing procedure, vacuum sealing and the in-house mechanical design.
Main contributors: D. Alesini, R. Di Raddo, L. Faillace, L. Giuliano, M. Magi, M. Migliorati
In house building of the accelerating cavities
Screws to join the cells
Water hole
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 27
3D Electric field on beam axis
Power in
Power out
Dual-feed different input and output RF power couplers
Single Cell design
a | b | d | R0 | t | r1/r2
5 | 21.03 | 17.5 | 6 | 3 | 1.25
Asymmetric cell to facilitate in-house mechanical processing.
Beam axis
INFN Acceleratori - LNF, 4-5 aprile 2024
Prototyping phase 2

## Slide 28
Prototyping phase 2
A prototype of 12 cells with couplers has been brazed @INFN LNF –FRASCATI oven to perform low-power RF tests.
Main contributors: D. Alesini, R. Di Raddo, L. Faillace, L. Giuliano, M. Magi, M. Migliorati
In house building of the accelerating cavities
Structure in the INFN LNF –FRASCATI oven ready to be brazed
INFN Acceleratori - LNF, 4-5 aprile 2024

## Slide 29
Measurement of electric field with “bead-pull”
A Radio Frequency (RF) signal is applied by the Network Analyzer (VNA)
A perturbing object (glue bead) is attached to a horizontal fish wire.
RF cavity under test
Step- motor
Network Analyzer
Pc
RF cavity under test
RF signal
1
2
3
4
3. Using a step-by-step motor, the wire goes through the structure and measures the electric field perturbations caused by the bead
4. All the data are stored in a PC

## Slide 30
BeadPull measurements after tuning
After tuning the structure presents a average phase advance is 120,38° , the petals are superimposed in the RF phase diagram
The electric field presented still a stationary pattern: final coupler needed to be tuned
Others contribution: F. Cardelli, D. Alesini,L.Faillace, L.Ficcadenti A. Mostacci
Tuning by deformation
Input coupler
Output coupler
tuners

## Slide 31
INFN Acceleratori - LNF, 4-5 aprile 2024
Prototyping phase 3
To be tested in power at LNF

## Slide 32
PULSE COMPRESSOR (G. Torrisi LNS)

## Slide 33: FUNDINGS & SCHEDULE

## Slide 34
TOWARD VHEE LINACS
FOR DEEP TUMORS
INFN PROJECT 2021
FRIDA
PNRR NATIONAL PROJECT 2022
A compact C-band Linac for FLASH therapy: accelerator and dosimetry study (2020)
Commissione Scientifica Ateneo
Sapienza medie attrezzature
SBAI Contracts
260 kE
Basic VHEE Prototype
SAPIENZA 1.500 kE
SIT 230 KE
R&D RF Structure 120 kE
Pulse compressor 70 kE

## Slide 35
PROJECT TIMELINE

## Slide 36
COLLABORAZIONE SAPIENZA-INFN

## Slide 37
ROMA1
LNF
LNS

## Slide 40
PE6 : FLASH-LAB
Referente Sapienza PNRR
Vincenzo Patera
Project Coordinator
Luigi Palumbo
Treatment Plans VHEE
A. Sarti
RF Design : L. Giuliano
LINAC VHEE-FLASH
M. Migliorati
RF prototype: A. Mostacci
Beam Dynamics : L. Giuliano
Beam Delivery : A. Vannozzi (INFN) & A. Trigilio
Diagnostics : A. Mostacci
Beam measur. : E. Chiadroni
Linac Parameters : M. Migliorati & L. Giuliano
e-beam/dose character L. Giuliano
Linac Controls : A. Mostacci
Magnets: A. Vannozzi (INFN)
Online Beam Monitor : M. Marafini
Dose engine: G. Traini & G. Franciosini
Optimization: A. Schiavi & A. Di Gregorio
Radioprotection: A. Di Gregorio & G. Franciosini, F. Perondi
Mech. machining: M. Magi
WEB, Repository & Network M. Coppola
SIT

## Slide 41
RF Design : L Faillace, Alesini
RF prototype: Faillace, Alesini
Beam Dynamics : L. Faillace
Diagnostics : G. Franzini
RF Linac Controls : S. Gallo
Magnets: A. Vannozzi
LNF Collaboration
Beam delivery : A. Vannozzi
C power test: Cardelli/Pioli
G. Cuttone
P. Cirrone
G. Torrisi (Pulse compresor)
LNS Collaboration
L. Ficcadenti
Facilities (off. Meccanica)
……….
Roma1 Collaboration
Controls: Piersanti
Brasatura: Lollo, Di Raddo