## Slide 1: Silicon sensors in Medical Physics
Simona Giordanengo
INFN Torino

## Slide 2: Outline
Radiation detectors in medical physics:
  where and why?
Silicon sensors for:
  beam monitoring in particle therapy,
  timing applications,
  Flash radiotherapy (FLASH RT),
  proton/ion imaging.
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
Notes: In my talk I’ll start with a short review of radiation detectors applications in medical imaging and therapy, then the focus of the talk will be on new beam monitors for PT and FLASH RT followed by an overview of silicon sensors development for dosimetry, micro-dosimetry and radiobiology

## Slide 3: Introduction
Radiotherapy with external beams
Medical Imaging techniques
X-ray Imaging
Computed Tomography (CT)
Positron Emission Tomography (PET)
Single Photon Emission Computed Tomography (SPECT)
Image fusion PET-CT
Photons (6 - 18 MV)
Electrons (6 - 18 MeV)
LINAC for radiotherapy
Protons (60-250 MeV/u)
Carbon ions (115-400 MeV/u)
Cyclotron for protons and Synchrotron for protons&light ions
Benefits from new technologies
 improve image resolution and sensitivity
Radiation detectors in medical physics: where and why?
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
Notes: This slice is trivial but essential to highlight the crucial role that detet

## Slide 4: Radiation detectors for radiotherapy
Radiotherapy with external beams
Photons (6 - 18 MV)
Electrons (6 - 18 MeV)
LINAC for radiotherapy
Protons (60-250 MeV/u)
Carbon ions (115-400 MeV/u)
Cyclotron for protons and Synchrotron for protons&light ions
Beam commissioning
  - Before facility startup
Beam monitoring
  - Online to guide and control the beam delivery
Quality Assurance
- daily/weekly/monthly
Dosimetry
- daily
Microdosimetry
Radiobiology
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024

## Slide 5: The role of beam monitors in the nozzle
CNAO horizontal beam line
Phys. Med. Biol. 68 (2023) 105013
PSI – Gantry 1
Nucl. Instr. Met. A 698 (2013) 202
Nozzle equipped with particle detectors integrated with the accelerator control system to guide and control the beam in real time.
Charged particles
𝛄 511 keV
𝛄 511 keV
neutrons
𝛄 prompt
beam
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
Notes: Well, you all know that to guide and control the beam delivery, every nozzle is equipped with beam monitors integrated with the accelerator control system.
which measure in real time the beam fluence and position, mandatory and if needed.
One easy question: are they, or the most advanced beam monitors able to control the beam energy online?
No
Why?
First because the beam energy is guaranteed by the accelerators, but the crucial reason is that it’s not trivial the energy measurement in the limited space and time constrained by clinical requirements. We will see this topic later
detectors

## Slide 6
Pencil Beam
Tumor
Target dimension: 1 ÷ 40 cm
Beam range in water for protons&ions  | 330 cm
Beam energy range (protons) [MeV, β, MIP] Beam energy range (C ions) | 60250 MeV, β 0.34-0.59 , 5-2 MIP 120400 MeV/u (β 0.46-0.71)
Beam intensity (particles/s) | 1081010 p/s 4×1074×108 Cions/s
Spot size (FWHM in air at the isocenter) | FWHM 720 protons /48 mm Cions (highestlowest) Energy
Field size at the isocenter | 20x20/40×40 cm2 (min-max)
Beam characteristics for particle therapy
Not to scale
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024

## Slide 7: Beam parameters measured in the nozzle
Exit window
Pencil Beam
FWHM  3 ÷ 10 mm
FWHMx
FWHMy
Tumor
Not to scale
Beam flux and fluence (beam Intensity – dose rate)
Accepted uncertainty 1-2 %
Transversal Beam positions
Accepted uncertainty ±0.5 mm
Transversal Beam shape (FWHMs)
Accepted uncertainty ±1 mm
To be measured during treatment, along the beam line just before the patient …
Beam range or energy
 Accepted uncertainty 1 mm (0.5 – 1 MeV)
Currently not online!!
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024

## Slide 8: Pencil beams are driven by beam monitors
En
Monitor on-line the beam fluence and position (mandatory), FWHM (optional)
Set the beam position spot by spot through the scanning magnets power supplies
Correct on-line the beam position (feed-back operations)
Stop the beam when something is wrong (beam parameters out of range)
Monitor 2
Treated spots
ionization chambers
2 : Integral chambers
2 : Strip chambers
Scanning magnets
Monitor 1
y
x
z
Isocenter
Scanned Field
y
x
z
Isocenter
Scanned Field
Sinergy with scanning system
First slice
(min energy)
Last slice
(max energy)
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
Notes: ,

## Slide 9: Online beam delivery progress
First slice
(min energy)
Last slice
(max energy)
Dose Delivery System GUI in the CNAO Local Control Room
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024

## Slide 10: Robustness vs Performances in medical applications
Robustness
Performances
and accuracy
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024

## Slide 11
Beam monitor must deal with different beam temporal structures
Synchrotron
The cyclotron provides proton therapy reliably and at low cost (main vendors on the market: IBA, Varian, Mevion, Hitachi).
Energy modulation with degraders.
Multi-ion accelerator (more flexible than cyclotrons).
Beam delivered in spills
1 s
3-5 s
Cyclotron
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
11
Beam monitor
integrated charge

## Slide 12
Gas used:
Air
N2,
Ar-CO2
E = 1 kV/cm
Sequence of Parallel Plate ICs:
 single large electrode for FLUENCE
 electrodes segmented in strips for BEAM POSITION and SHAPE
Thickness ~ 1 mm of water
Reliability & long term (years) stability
Large (up to 40x40 cm2) sensitive area
Simple to use
Deeply studied manufacture
A few mm water equivalent thickness
Radiation resistant
“Currently ionization chambers(IC) are predominately used (in fact they are legally required in most countries). “
Roadmap: proton therapy physics and biology
Paganetti et al. Phys. Med. Biol. 66 (2021) 05RM01
ADVANTAGES
State-of-the-art  large area gas ionization chambers
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
12

## Slide 13: State-of-the-art  large area gas ionization chambers
Gas used:
Air
N2,
Ar-CO2
E = 1 kV/cm
Sequence of Parallel Plate ICs:
 single large electrode for FLUENCE
 electrodes segmented in strips for BEAM POSITION and SHAPE
Thickness ~ 1 mm of water
Long collection times ~ 100 µs
Low Sensitivity ~ 104 protons
Poor time resolution ~ no/poor
Deviation from linearity @ high dose rates
Not suitable for
fast scanning modalities
timing applications
high dose rates (FLASH)
LIMITATIONS
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
13

## Slide 14
New challenges for beam monitoring
New accelerators (different timing structures and treatment procedures)
Advanced treatment modalities (multiple rescanning, online image-guided, …)
Hipofractionation up to FLASH
Fast pencil beam scanning
Proton Computed Tomography
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
14

## Slide 15
INFN MoveIT project (2017-2021)
Solid state detectors for p-beam monitoring
Collection times ~ 100 µs
Sensitivity ~ 104 protons
Time resolution ~ no/poor
IONIZATION CHAMBERS
Not suitable for
fast scanning modalities
timing applications
SOLID STATE SENSORS: LGAD 45-60 µm thickness
int. gain ~ 10
~ ns
single protons
<< 100 ps
direct counting # protons
timing applications
Main issues at φ = 1010 p/cm2s
Signal pile-up
→ fast sensors & readout
→ segmentation
Radiation tolerance
→ manufacturing strategies
→ damage compensation
Large sensitive area
→ technological challenge
Increased complexity
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
15

## Slide 16
Thin Low Gain Avalanche Detectors
thickness of sensitive volume < 50 µm
internal charge multiplication ~ 10
+
Enhanced signal of very small duration
Time resolution of tens of ps
Strip detectors (strip area ~ 3 mm2, active thickness 45µm)
Detectors for particle counting
Large area (2.7×2,7 cm2)
144 strips
Detectors for timing applications
Smaller size, 11 strips
Strip area 2.2active thickness 45 total thickness 615 )
Si- substrate removed to reduce total thickness to 70 µm
Beam spot
1 cm FWHM
Pads for large ionization rate studies
4 active thicknesses (15/20/30/45 µm)
5 pad sizes (0,125/0,25/0,56/1/2 mm2)
Internal gain | 
yes | no
 | 
 | 
 | 
Use:
Protons
Use:
C-ions
FLASH
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
16

## Slide 17
Proton counting
117 MeV protons
Data collected with strip sensors + digitizer
Short peak duration
Large Landau fluctuations
Single particle discrimination
Signal pile-up
threshold
Pile-up correction studies
Mohammadian-Behbahani M, et. al., NIM A 1040 (2022) 167195
Monaco V, et. al., Phys. Med. Biol. 68 (2023) 235009
Strip area 2mm2
CNAO beam flux
5 x 108 p/cm2s
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024

## Slide 18: Technologies for multi-channels silicon detectors
Proton pencil
beam
ABACUS
LabVIEW interface
3 FPGA
LGAD strips
1 cm FWHM
108 p/s*cm2
146 strips
6 ABACUS
24x6 channels
Kintex-7
Gbit ethernet interface
ESA-ABACUS
Custom Frontend board
Custom ASIC
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
Notes: Slide per sottolineare la complessità di un rivelatore e le tecnologie necessarie

## Slide 19: 2.7×2.7 cm2 particle counter (ESA ABACUS)
Silicon strip sensor
Strips width: 114 μm
Pitch: 180 μm
Active thickness: ~50 μm
Capacitance: ~7 pF
The ABACUS chip (x6)
110 nm CMOS technology, chip area = 2 × 5 mm2, 24 channels, 144 in total. CSA dynamic range: 4 fC – 150 fC. Dead time : < 10 ns.
LVDS
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024

## Slide 20
Particle counter test at CNAO to study the counting efficiency with protons and carbon ions
Runs with different energies in the clinical energy range
  60 - 230 MeV (protons)
  115 - 400 MeV/u (carbon ions)
Different beam fluence rate:
  20, 50, 100 % of maximum fluence rate
beam direction
x
y
Z

## Slide 21
  Three different energies, 60 - 230 MeV (protons), 115 - 400 MeV/u (carbon ions)
  Three different beam fluence: 20, 50, 100 % of the maximum fluence rate.
  Counting efficiency: 50% protons, > 90 % Carbon ions
Protons
Carbon Ions
Measured beam profiles and fluence at CNAO
Protons per spill
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024

## Slide 22: Perspectives about online beam fluence and position control
Technology ready for fix pencil beam characterization.
Improved front-end readout will overcome noise issues encountered with protons.
Current LGAD radiation tolerance will guarantee 1 year of safe operation
The challenge
Increase the sensitive area (> 24x24 cm2)
  Pixel segmentation needed
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
22

## Slide 23: Exploiting timing performances of fast silicon detectors
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
23

## Slide 24: Timing application (i): beam energy detector
beam
LGAD sensors
L
Beam telescope
Beam energy/range from time-of-flight
range can be measured with ±1 mm uncertainty;
need few ms of active acquisition time;
requires precise calibration of the system.
A. Vignati, et. al., Phys. Med. Biol. 65 (2020) 215030
A. Vignati, et. al., Med. Phys. 50 (2023) 5817-5827
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
24

## Slide 25
Two Si sensors with the same characteristics placed in a telescope configuration 30 cm apart.
Beam energies (MeV/u): 115.2 and 398.8.
Bias voltage: 50V, 100V, and 200 V.
σt (ps) = σ() /
Time difference
Coincidences
Time resolution values for single ion crossing were less than 26 ps
(relative error < 2%)
Time resolution for ion crossing with thin silicon sensor
CARBON IONS
115 MeV/u
400 MeV/u

## Slide 26: Perspectives about online beam energy control
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
26
Today
Current treatments do not need online beam energy control
 the beam energy is guaranteed by the accelerators control system
Future
 Clinical treatments with LINAC accelerator will provide fast energy modulation with online beam energy control.
The challenge
 Energy measurement in the limited space and time constrained by clinical requirements.

## Slide 27: Particle crossing time measurements CERN picoTDC evaluation board (64 input channels)
3ps or 12ps binning
very low jitter (<1ps)
High rate capability
Readout through FPGA
Successfully integrated with 8 channels of ESA-ABACUS
Conversion efficiency 100%
Tested up to 150 MHz freq.
15 ps time resolution
Qin = 48 fC
f = 1 MHz
1 channel
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
27

## Slide 28
ASIC 5
ASIC 4
CTRL 3-4
ASIC 3
ASIC 6
ASIC 1
ASIC 2
CTRL 1-2
CTRL 5-6
ABACUS digital outputs
Digital control signals
Kintex 7 KC705 FPGA
Kintex 7 FPGA
CERN PicoTDC
Acquisition PC
picoTDC integration
F. Pennazio

## Slide 29
Timing application (ii): beam time microstructure with picoTDC
362 ns
Carbon ion energy = 398 MeV/u
Carbon ion energy = 115 MeV/u
The bunch structure and the accelerator radiofrequency were observed for all the energies.
Preliminary results
Preliminary results
Time difference between 2 c-ions
3 ps bin size
20 µs time window
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
29

## Slide 30: Timing application (iii): Start time for Prompt Gamma Timing
IC chambers
PET-PGT detector
(γ position and time)
patient
PET-PGT detector
(γ position and time)
LGAD detector
START for
Prompt Gamma Timing
F. Pennazio’s talk
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
30
Notes: The focus is on beam monitors, followed by the PET-PGT system and charge particles detectors, all these instrumentation operate online, in the treatment room during the delivery.
Accurate and safe treatment need several detectors
[23] Muraro S et al. Front Oncol6:177, 2016
[24] Mattei I et al. Nucl Instrum Methods Phys Res A 954:161229, 2020
[25] Félix-Bautista R et al. Phys Med Biol 64:175019, 2019
[26] Henriquet P et al. Phys Med Biol 57:4655, 2012
[27] Fischetti M et al. Sci Rep 10:20735, 2020
Scintillating crystals coupled to Silicon PhotoMultiplier (SiPM) measure the secondary particles produced by the interaction between the primary beam and human tissues to provide in-vivo treatment delivered verification.
Both annihilation photons (for PET modality) and nuclear de-excitation photons (multi-detector PGT modality) are collected within the same detector (PET-PGT detector).
SiPMs readout signals are integrated with the information of the primary particle transit time, measured by new beam monitors based on thin, segmented and fast planar silicon detectors readout through high-performance DAQ system, based on fast TDC and FPGAs.
Secondary charged particles [23] are detected with very high efficiency as a nearly background free signal. Their production positions, through the Interaction Vertex Imaging (IVI) technique, can be used to spot
inter-fractional morphological changes and beam range variations with respect to the planning [24] or to check the beam position in the transverse plane inside the patient [25]. Silicon pixel detectors [26] and scintillating
fibres devices have been proposed and, recently, the latter have been tested on patients treated with carbon ions within the INSIDE project [27]. However, to fully exploit IVI capabilities, the detected tracks of secondary
charged particles have to be maximised by reducing the detector dead time and enlarging the active area, getting as close as possible to the patient.

## Slide 31: Exploiting thin silicon sensors for FLASH RT
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
31

## Slide 32: FLASH radiotherapy
FLASH EFFECT:
Does not induce classical radiation induced toxicity in normal tissues.
Retains antitumor efficacy compared to standard RT
FLASH RT delivers radiation (electrons, photons, particles) at ultra-high dose rate (UHDR, average dose rate > 40 Gy/s) in < 200 ms.
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
32

## Slide 33: FLASH radiotherapy
A crucial role: dose delivery time structure (parameters need to be kept under control )
The most of the pre-clinical studies using electron beams (by LINACs with E<20 MeV)
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
33

## Slide 34
New field for BM: FLASH RT
Interrupting a FLASH treatment quickly enough to avoid a misadministration is clearly a non-trivial task and no publicly available testing data exist showing a commercial system is able to stop delivery of a FLASH dose quickly enough to address this need. [Taylor et al. Med Phys. 2022]
Ultra High Dose Rate and high dose deliveries --> Reproducibility and interlock activation
Main Challenges
Main Requirements
Linearity and no saturation
Radiation hardness
Only the dependence on the average dose rate and on the duration of the entire irradiation has been clearly observed so far, but the roles of…
Dose per pulse
Instantaneous dose per pulse
Pulse duration and frequency
.. still remain to be understood.
[Di Martino 2020 frontiers]
Introduction of correction factors
Optimization of the existing technologies
Investigation of novel detection methods
Possible solutions
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
34

## Slide 35: Solid State Devices for Beam Monitoring
Ultra-thin ~10, segmented, high polarized
High sensitivity, spatial res, developed technology
Unkown factors: linearity with dose-rate, recombination effect, radiation resistance
Radiation hardness, high resistivity, large saturated carrier velocity
Challenging issues: dose-rate linearity, possibility of straddle areas several cm2
High electrical stiffness, speed of charges, melting T, thermal diffusivity, industrual maturity
Preliminary sim: dose-rate linearity up to 1011 Gy/s on X-ray beams for SiC membrames (2 thick)
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
35

## Slide 36: Silicon sensors for FLASH RT beam monitoring
Deviation from linearity @ high dose rates | Less recombination @ high dose rates 102 × E field 102 × charge mobility 10-1 × thickness
Not suitable for high dose rates (FLASH) | New applications Monitoring high intensity and pulsed beams (high dose rates needed for FLASH RT)
Ionization chambers
Main issues
High dose rates (FLASH)
→ 103 × dose rates
→ plasma effects in silicon
Radiation tolerance
→ manufacturing strategies
→ damage compensation
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
36
Ultra-thin ~10, segmented, high polarized
High sensitivity, spatial res, developed technology
Unkown factors: linearity with dose-rate, recombination effect, radiation resistance
Conventional IC used in LINACs

## Slide 37
Monitoring of FLASH UHDR electron beams
13mm solid water slab
Flash diamond and silicon sensor in same conditions
ElectronFlash accelerator (CFR - Pisa)
9 MeV electrons pulsed beam
Beam current: 1-100 mA
Pulse duration: 4 𝜇s
Pulse frequency: 5 Hz
Uniform fields using 3 cm PMMA plastic applicator
f
45/ 30𝜇m thickness
2/1/0.25 mm2 area
Bias voltage: 10V ÷ 200V
Dose/Pulse 0 ÷ 10Gy
Sensors tested
INFN-FRIDA project
2021-2024
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
37

## Slide 38
Monitoring of FLASH UHDR electron beams
Collected charge/pulse scales with pad area and sensor thickness
Ratios between different area/thickness independent from dose/pulse
Charge collected (nC/pulse)
Measured dose (Gy/pulse)
Measured dose (Gy/pulse)
4 𝜇s pulse duration
4 𝜇s pulse duration
FLASH regime
FLASH regime
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
38

## Slide 39: 4D-tracking with LGADs for proton Radiography&CT
Measurement of residual proton energy from ToF in pCT applications
Calorimeter/
Range monitor
proton
patient
Tracking
detectors
Tracking
detectors
Residual proton energy
ToF
Phys. Med. Biol. 67 (2022) 095005
Schematic of a TOF calorimeter based on LGAD detectors.
The TOF is measured between two timing stations (T1/T2 and
T3/T4), each consisting of two generic LGAD planes
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
39

## Slide 40
Summary and Conclusions
> 24x24 cm2 Si-based beam monitor in clinical nozzle
Specific optimization of
sensors,
frontend readout,
high performance DAQ,
tailored on specific applications
(counting, timing, imaging, dosimetry, radiobiology, …)
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024
40

## Slide 41
Silicon detectors offer interesting features for new developments in radiation therapy
For beam monitoring in PT: integrating counting and timing in the same device seem possible with state-of-the-art TDCs
Good linearity with dose per pulse was demonstrated in FLASH e- beams
Conclusions
Silicon sensors in Medical Physics - S. Giordanengo, INFN Torino – TREDI Torino - 22/02/2024

## Slide 42
Thank you for your attention!

## Slide 43: Spare slide

## Slide 44: New detector requirements
HIGH SENSITIVITY
   GOOD SIGNAL TO NOISE RATIO
   FAST SIGNAL CREATION
FINE GRANULARITY
TRANSPARENT TO THE BEAM
RESISTANT TO RADIATION
LARGE AREA ( > 20x20 cm2)
DEAD TIME FREE
Increase the charge generated for each proton
W small
Hgher density
→ S larger
Short collection time
Reduced thickness
High Electric field E
,
105 e- /h in 300 μm → Q=16 fC
For Silicon ….

## Slide 45: Hybrid microdosimeter exploits silicon sensors tracking capability
https://doi.org/10.3389/fphy.2020.578444
A Novel Hybrid Microdosimeter for Radiation Field Characterization Based on the Tissue Equivalent Proportional Counter Detector and Low Gain Avalanche Detectors Tracker: A Feasibility Study
Notes: In the proposed
setup, the TEPC will provide the energy deposition ϵ directly in a tissue-equivalent medium while the LGADs will offer information about particle spatial distribution with a precision of about 200 or 300 μm, depending on the chosen configuration. Together with the TEPC that provides energy depositions, LGADs constitute the tracker stage of HDM.
to reconstruct the particle path inside the TEPC, from which we could estimate the real track length.
Distribuzione di energia divisa per la lunghezza media, con il tracciamento misuri dove è passata la particelle e dividi per una lunghezza vera e non media (rivelatori devono essere molto precisi per energie basse dei secondari che devono attraversare tutti I layer del rivelatore. Trasparenti.
Per fare tracking ti aiuta a misurare in modo più preciso la

## Slide 46: Imaging with protons and ions  proton/ion radiography proton/ion Computed Tomography  needed to reduce uncertainty in the water equivalent path, density  measure the residual proton energy  measure directly the Stopping Power 3D distribution will take advantage from 4D particle tracking through silicon sensors

## Slide 47: Imaging with X-rays : state-of-the-art
MAMMOGRAFIA
ANGIOGRAFIA
TOMOGRAFIA COMPUTERIZZAZA
Film-based Detectors
Scintillation Detectors
Solid-state Detectors  amorphous silicon (a-Si) or cadmium telluride (CdTe) coupled with thin-film transistor (TFT) arrays.
Advantages: high spatial resolution, fast readout times, and low noise levels.
Used in Direct digital radiography (DR) systems.
Charge-Coupled Devices (CCDs) and CMOS detectors  not commonly used in medical X-ray imaging
Amorphous Silicon Flat-Panel Detectors (FPDs)  amorphous silicon layer coupled with an array of thin-film transistors (TFTs).
FPDs offer high spatial resolution, fast readout speeds, and low radiation dose requirements.
Notes: Film-based Detectors: Traditional X-ray imaging initially used photographic films as detectors. X-ray photons expose the film, causing a chemical change that creates a latent image. The film is then developed to produce a visible image. While still used in some applications, film-based detectors are increasingly being replaced by digital detectors due to their lower sensitivity and longer processing times.
Scintillation Detectors: Scintillation detectors use materials called scintillators that emit light when exposed to X-rays. Common scintillator materials include cesium iodide (CsI), gadolinium oxysulfide (GOS), and amorphous selenium (a-Se). Photodetectors, such as photomultiplier tubes (PMTs) or photodiodes, then convert the emitted light into an electrical signal. Scintillation detectors are used in both indirect and direct digital X-ray imaging systems.
Solid-state Detectors: Solid-state detectors directly convert X-ray photons into electrical signals without the need for a scintillator. These detectors typically consist of materials like amorphous silicon (a-Si) or cadmium telluride (CdTe) coupled with thin-film transistor (TFT) arrays. Solid-state detectors offer advantages such as high spatial resolution, fast readout times, and low noise levels. They are commonly used in direct digital radiography (DR) systems.
Charge-Coupled Devices (CCDs) and Complementary Metal-Oxide-Semiconductor (CMOS) Detectors: CCDs and CMOS detectors, which are commonly used in digital cameras and consumer electronics, can also be adapted for X-ray imaging. These detectors capture the light emitted by scintillators or directly detect X-rays and convert it into electrical signals. While not as commonly used in medical X-ray imaging, they find applications in research, industrial, and astronomy X-ray imaging systems.
Amorphous Silicon Flat-Panel Detectors (FPDs): FPDs are a type of solid-state detector commonly used in modern digital X-ray systems. They consist of an amorphous silicon layer coupled with an array of thin-film transistors (TFTs). X-ray photons interact directly with the silicon, generating electron-hole pairs that produce an electrical signal. FPDs offer high spatial resolution, fast readout speeds, and low radiation dose requirements.
If interpreted as waves, X-rays-just like visible light-experience a phase shift in matter, and this-if exploited correctly-can produce a new class of X-ray images, which then depict the wave interactions of X-rays with matter, rather than only the attenuating properties, as done until now.

## Slide 48: Imaging with e+ Arrays of photon detectors around the patient to detect the 2 γ opposite directions
For diagnostic purpose
Therapeutic hadron beams produce β+ emitters in the body
Offline and in-beam PET provide the range of particles and the dose distribution
PET-PGT detector
(γ position and time)
PET-PGT detector
(γ position and time)
Notes: Positron Emission Tomography (PET) uses detectors specifically designed to capture the emitted positrons from radiopharmaceuticals administered to the patient. These detectors are crucial components of PET scanners, enabling the reconstruction of images showing the distribution of the radiotracer within the body. Here are the main types of detectors used in PET:
Scintillation Crystals: These crystals are often made of materials like lutetium oxyorthosilicate (LSO), lutetium-yttrium oxyorthosilicate (LYSO), or gadolinium orthosilicate (GSO). When a positron annihilates with an electron, it produces two gamma photons moving in opposite directions. These photons interact with the scintillation crystals, producing flashes of light (scintillations) that are detected by photodetectors.
Photodetectors: Photomultiplier tubes (PMTs) or avalanche photodiodes (APDs) are commonly used as photodetectors in PET scanners. They convert the scintillation light emitted by the scintillation crystals into electrical signals that can be processed by the PET system.
Coincidence Detection Circuitry: PET detectors are arranged in pairs or arrays around the patient. When two detectors detect simultaneous gamma photons (resulting from positron annihilation events occurring near-simultaneously), it is termed a coincidence event. Coincidence detection circuitry is used to identify and record these events, which are then used to reconstruct the PET image.
Time-of-Flight (TOF) Detectors: Some modern PET scanners incorporate TOF detectors, which measure the time difference between the detection of two coincident gamma photons. This information can improve image quality and reconstruction by providing more accurate localization of the radiotracer uptake within the body.
Detector Arrays: PET scanners consist of arrays of detectors positioned around the patient. These arrays can have different configurations, such as rings or panels, depending on the scanner design. The detectors in these arrays work together to capture gamma photons emitted from the patient and produce data for image reconstruction.
These detector components work in concert to detect the gamma photons emitted during positron annihilation events, allowing PET scanners to create detailed images of biochemical processes in living tissues. Advances in detector technology continue to improve the sensitivity, resolution, and overall performance of PET imaging systems.

## Slide 49: Advanced Imaging with Phase-Contrast X-ray imaging for medical imaging
Grating Interferometry
can be performed both at Synchrotron Radiation (SR) facilities and using laboratory X-ray sources. Until now, GI has been applied in a laboratory setting
mostly with low-brilliance X-ray tubes, by using either long or compact interferometer
arrangements [1,3,4]. Liquid-metal-jet sources, combined with a phase and an analyzer grating,
improve the resolution and sensitivity of PCI, thanks to a higher photon flux [5]. GI systems at
25 keV [6] and 9.25 keV [7] yield reduced scan times or an improved signal to noise ratio;
The use of a high granularity detector pushes the resolution down to the micron scale.
X-ray sCMOS camera
If interpreted as waves, X-rays-just like visible light-experience a phase shift in matter, and this-if exploited correctly-can produce a new class of X-ray images, which then depict the wave interactions of X-rays with matter, rather than only the attenuating properties, as done until now.
Notes: Film-based Detectors: Traditional X-ray imaging initially used photographic films as detectors. X-ray photons expose the film, causing a chemical change that creates a latent image. The film is then developed to produce a visible image. While still used in some applications, film-based detectors are increasingly being replaced by digital detectors due to their lower sensitivity and longer processing times.
Scintillation Detectors: Scintillation detectors use materials called scintillators that emit light when exposed to X-rays. Common scintillator materials include cesium iodide (CsI), gadolinium oxysulfide (GOS), and amorphous selenium (a-Se). Photodetectors, such as photomultiplier tubes (PMTs) or photodiodes, then convert the emitted light into an electrical signal. Scintillation detectors are used in both indirect and direct digital X-ray imaging systems.
Solid-state Detectors: Solid-state detectors directly convert X-ray photons into electrical signals without the need for a scintillator. These detectors typically consist of materials like amorphous silicon (a-Si) or cadmium telluride (CdTe) coupled with thin-film transistor (TFT) arrays. Solid-state detectors offer advantages such as high spatial resolution, fast readout times, and low noise levels. They are commonly used in direct digital radiography (DR) systems.
Charge-Coupled Devices (CCDs) and Complementary Metal-Oxide-Semiconductor (CMOS) Detectors: CCDs and CMOS detectors, which are commonly used in digital cameras and consumer electronics, can also be adapted for X-ray imaging. These detectors capture the light emitted by scintillators or directly detect X-rays and convert it into electrical signals. While not as commonly used in medical X-ray imaging, they find applications in research, industrial, and astronomy X-ray imaging systems.
Amorphous Silicon Flat-Panel Detectors (FPDs): FPDs are a type of solid-state detector commonly used in modern digital X-ray systems. They consist of an amorphous silicon layer coupled with an array of thin-film transistors (TFTs). X-ray photons interact directly with the silicon, generating electron-hole pairs that produce an electrical signal. FPDs offer high spatial resolution, fast readout speeds, and low radiation dose requirements.
If interpreted as waves, X-rays-just like visible light-experience a phase shift in matter, and this-if exploited correctly-can produce a new class of X-ray images, which then depict the wave interactions of X-rays with matter, rather than only the attenuating properties, as done until now.

## Slide 50: Beam delivery techniques
«The possibility of volumetric rescanning and other advanced techniques require the BDS to be able to deliver efficiently with fast energy modulation.»
REVIEW
Frontiers - 2021
doi: 10.3389/fonc.2021.780025
Phys. Med. Biol. 63 (2018) 145006
Scan paths for: spot scanning, raster scanning and line scanning
Iso-energy slice (E = 151 MeV)
Beam weights were optimized for a total field dose of 0.606 Gy
Notes: Different scanning patterns are available. For spot scanning, the beam is turned off between movement to the next spot but remains on for the whole delivery in raster scanning. The control is on MU and the beam is Spot scanning and raster scanning are driven by the dose, it means that the beam is delivered in each spot without moving position and the main difference is
Spot, raster and continuous line scanning patterns for PBS delivery. Spots and solid lines indicate beam-on irradiation and dashed lines indicate
movement with beam-off. For
continuous line scanning, the dose is delivered across a linear path (rather than as spots) and may be turned off between movement to subsequent lines.
For spot and raster scanning techniques the delivery is done in dose-driven mode: the beam stays on until a certain preset dose amount of charge has been collected in a beam monitor.

## Slide 51: Protons - beam projection on y axis
The lower the energy and the wider the beam because of multiple scattering of protons with air
Curves → Gaussian fit
0 cm
2.7 cm
We are loosing the
tails of the beam for
62 MeV
Strip pitch = 180 µm
Expected FWHM at isocenter from Mirandola et Al 10.1118/1.4928397
Protons per spill
Strip number
FWHM [cm]
Energy [MeV]

## Slide 52: C-ions counter
100%
50%
Beam Intensitiy 20%
3% difference between low-high energy
Tails of beam profile
3%
1% difference between low-high intensity
Pile-up inefficiency reduced to 0.5% after corrections
1%
Spots of 5*107 C ions
Comparison of beam FWHM with GAFCHROMIC films
0 cm
2.7 cm

## Slide 53: Beam monitor systems
Conventional IC used in LINACs
PTW 60019 microDiamond
Ultra-Thin Ionization chamber (UTIC)
Beam Current Transformers (BCT)
DOSIMETERS
BEAM MONITOR
Continuous check of beam parameters
IC CONV: Gas-filled IC  IC UHDR : high rate of recombination, too slow
Need of new beam monitoring device to stop delivery of a FLASH dose quickly enough
    High temporal resolution
    High spatial resolution
    Beam transparency
    Large response dynamic range
    Large sensitive area
    Radiation hardness
Existing technologies

## Slide 54: Experimental setup
TERA08 measurements
45 𝜇m thickness, 2mm2 area
RC circuit to extend signal duration and not exceed 256𝜇A for 64 chns
RC connected to TERA08 and NI module
Bias voltage 200 V
Increasing dose-per-pulse (DPP) from 0 to ~10Gy/pulse
Oscilloscope measurements
45𝜇m / 30𝜇m thickness, 2mm2, 1mm2 ,0.25 mm2 area
3 pads connect to 3 oscilloscope channels
Bias voltage: 10V, 50V, 100V, 150V, 200V
Increasing DPP (from 0 to ~10Gy/pulse)
Compare different areas/thickness charge generation
National Instruments PXIe-1071 PXI Chassis
TERA08
Keysight infiniuum S-series DSOS254A
45𝜇m
30/45𝜇m
f
RC

## Slide 55: Readout system: TERA08
Readout with TERA08 (64 equal CHNs)
In each CHN current-to-frequency converter (each digital pulse = fixed input charge quantum)
Converter based on recycling integrator architecture
Chip structure
DAQ Period () | (fC) | Max conversion freq per chn | Max conversion (total) | Max current (for 64 CHNs)
1e4 (0.01 s) | 200 fC | 20 MHz | 1280 MHz | ± 256 𝜇A
 |  | Max conversion freq per chn | Max conversion (total) | Max current (for 64 CHNs)
1e4 (0.01 s) | 200 fC | 20 MHz | 1280 MHz | ± 256 𝜇A

## Slide 56: TERA08 and oscilloscope comparison
45 𝜇m thickness, 2mm2 area
200V bias voltage
Good linearity (R2 > 99%) up to dose-rates >10Gy/pulse (1Gy/pulse is already FLASH regime)
Good correlation of charge measured with TERA08 and oscilloscope
FLASH regime
UNPUBLISHED RESULTS!

## Slide 57: Electric Field distortion
200 V
At bias < 150 V (where the sensor is completely depleted) a shortening of the signal was observed: electric field distortion at high dose rates?
TCAD Sentaurus simulations ongoing
UNPUBLISHED RESULTS!

## Slide 58: Next steps: TERA09
Frontend chip based on 64 charge recycling CHNs
Extended current range with respect to TERA08 (preliminary design and test phase): 12 µA / chn with 200 fC.
Larger sensor (Area 2.7⨉2.7 cm2 and 146 strips) to cover all beam spot area (~ cm2)
Strip based / pad based system: Online control of beam shape and dose after one single shot
Detector board
TERA09
Current range | 100 pA-100 μA
Max conv freq | 62.5 MHz
Large sensor
[ Designed to cover proton beam spot ]
WORK IN PROGRESS!

## Slide 59: Electric Field distortion
50 V
UNPUBLISHED RESULTS!
At bias < 150 V (where the sensor is completely depleted) a shortening of the signal was observed: electric field distortion at high dose rates?
TCAD Sentaurus simulations ongoing