## Slide 1: Hybrid detectors for soft X-ray photon science using iLGAD sensors
Viktoria Hinger :: Postdoc :: PSI Center for Photon Science :: Detectors Group
19th TREDI Workshop on Advanced Silicon Radiation Detectors :: 21 February 2024

## Slide 2: Soft X-ray applications at synchrotrons and FELs
L-edges of 3d transition metals
  Fe, Cu, Zn, …
    Batteries
    Quantum materials
    Superconductors …
“Water window” and light elements
  C and O K-edges (282 – 543 eV), B, F, Mg, …
    Biological samples
    Catalysis
    Photosynthetic water splitting …
Magnetic contrast ptychography of BiFeO3
T. A. Butcher et al 2023 arXiv:2308.13465
Time resolved resonant inelastic X-ray scattering (RIXS)
R. Abela et al 2019 J. Synchrotron Rad. 26, 1073-1084
BN nanobamboo STXM & ptychography imaging, 192 eV
J. Vijayakumar et al 2023 J. Synchrotron Rad. 30, 746-757
STXM
Amplitude
Phase
TEM

## Slide 3: Limitations of current systems: frame rate and area
Charge coupled device (CCD)
  Low noise ( 1 e-)
  High spatial resolution
  Slow readout ( 1 Hz )
  Shutter required
  Limited area ( 2 × 2 cm2)
Scientific CMOS image sensor
  Low noise (1-2 e-)
  High spatial resolution
  Moderate frame rates (≤ 70 Hz)
  Limited dynamic range
  Limited area ( 2 × 2 cm2)
Electron multiplying CCDs at SLS SAXES
(RIXSCam, https://www.xcam.co.uk), courtesy of T. Schmitt
GSENSE400BSI (https://gpixel.com)
K. Desjardins et al 2020 J. Synchrotron Rad. 27, 1577-1589
pnCCD: 3e-, 150 Hz, 4 × 8 cm2
M. Kuster et al 2021 J. Synchrotron Rad. 28, 576-587.

## Slide 4: Hybrid detectors can tackle these challenges …
  High frame rates (≥ kHz)
  Large dynamic range ( 104-105 photons charge integrating, “unlimited” for photon counters)
  Customizable area (e.g., 4 × 8 cm2 single module, tileable)
Notes: Could put pictures

## Slide 5: … if we resolve limitations in the soft X-ray range
  Current limitations
  Readout noise (≥ 20 e-)
  Low quantum efficiency at soft X-ray energies
  We can optimize sensor and chip separately!

## Slide 6: JUNGFRAU – low noise, high dynamic range
G0
G1
G2
Charge integrating
75 x 75 µm2 pixel
max. 2.2 kHz frame rate
for continuos readout
Noise below Poisson limit over the full dynamic range
@ 12 keV
34 e- r.m.s ENC (was 52 e- in JF1.0)
Low-noise ASIC version JUNGFRAU 1.1
SNR 10 @ 1.2 keV (detection limit for JUNGFRAU with standard sensors)
Dynamic gain switching
Jungfrau mountain
Wire bonded ASIC
Courtesy of A. Mozzanica

## Slide 7: Thin entrance window improves quantum efficiency
Std. sensor
Al/Si interface
Passivated interface
Thin passivation
M. Carulla et al 2024 Sensors 24, 942
62% at 250 eV: comparable to standard CCD and CMOS image sensors

## Slide 8: Inverse low gain avalanche diode (iLGAD) sensors for soft X-rays
Courtesy of F. Baruffaldi

## Slide 9: iLGAD depth-dependent multiplication factor
electrons
holes
-
-
-
+
-
+
+
+
-
-
-
-
+
-
+
+
+
-
Courtesy of M. Carulla

## Slide 10: iLGAD depth-dependent multiplication factor
n+
gain layer
gain layer
n+
Standard gain layer
(Ultra) Shallow gain layer
Courtesy of M. Carulla
Mh
Me
Mh
Me
A. Liguori et al 2023 JINST 18 P12006
E1 < E2
E1 < E2
Mh/Me 0.25
Mh/Me 0.5

## Slide 11: The first JUNGFRAU-iLGAD prototype
G0
G1
G2
Standard gain layer
JUNGFRAU 1.1

## Slide 12: Characterization using X-ray fluorescence
200 V bias
Four metal targets
  Al (Kα: 1.5 keV)
  Ti (Kα: 4.5 keV)
  Cr (Kα: 5.4 keV)
  Fe (Kα: 6.4 keV)
-22°C to 15°C
5 µs to 100 µs integration time
10 µs, -22°C, 2 × 2 clusters, cumulative spectra across 80 × 50 pixels
V. Hinger et al 2024 Front. Phys. (forthcoming)

## Slide 13: Border pixels measure non-amplified signal…

## Slide 14: …allow extracting iLGAD multiplication factor M
Ti Kα-line (4.5 keV)
Multiplication

## Slide 15: Two domains of iLGAD multiplication factor M
10% variation
-22°C, 10 µs integration time
Notes: 3.5% pixel-by-pixel gain variations

## Slide 16: M is strongly influenced by the temperature
Customized simulation program (by J. Zhang):
Solving the 1D Poisson equation
Incorporating doping profiles and concentrations
Calculating electric field distribution
Impact ionization coefficients modelled with Okuto-Crowell and updated parameters by Currás Rivera and Moll (IEEE Trans. Electron Devices 2023 70 2919–2926)
Simulated difference: 1.25% integrated dose

## Slide 17: Cooling is required to reduce leakage current
From the slope of the pedestal as a function of integration time:

## Slide 18: Effective detector noise profits from increasing gain
5.5 e- r.m.s.
SNR = 51 (25 for 2 × 2 cluster)
Gain increase, leakage current decrease
SNR = 5
200 eV: SNR > 5 (2 × 2 cluster)
For photon energies < 500 eV:
Hole multiplication (Mh/Me 0.25) dominates
Improvements are available!
Shallow gain layer (Mh/Me 0.5)
Decrease of ASIC noise with decreasing temperature

## Slide 19: Applications

## Slide 20: Soft X-ray ptychography
F. Baruffaldi et al 2023, in preparation
Magnetic contrast ptychography of BiFeO3
T. A. Butcher et al 2023 arXiv:2308.13465
Single-photon-counting iLGAD-EIGER detector
at SLS SIM beamline
Scanning Transmission X-ray Microscopy (STXM)
+
Coherent Diffraction Imaging (CDI)
Wavelength limited resolution

## Slide 21: Resonant Inelastic X-ray Scattering (RIXS)
hRIXS at EuXFEL SCS,
courtesy of N. Duarte
Energy resolution
Spatial resolution
Rectangular pixels (“strixels”)
JUNGFRAU-iLGAD module
R. Abela et al 2019
J. Synchrotron Rad. 26,
1073-1084

## Slide 22: Takeaway messages

## Slide 23: iLGADs with thin entrance window enable hybrid detectors to operate in the soft X-ray range
Improved quantum efficiency
Signal amplification

## Slide 24: By cooling the detectors, we can resolve single photons down to 200 eV
Multiplication factor depends strongly on temperature
SNR > 5 at 200 eV with T < -50°C, improvements are realistic
1 keV: SNR = 200  200 eV: SNR > 5 (2 × 2 cluster)
SNR = 5

## Slide 25: The technological advancements of X-ray light sources are driving the development of detectors
Magnetic contrast ptychography of BiFeO3
T. A. Butcher et al 2023 arXiv:2308.13465
Time resolved resonant inelastic X-ray scattering (RIXS)
R. Abela et al 2019 J. Synchrotron Rad. 26, 1073-1084

## Slide 26: Acknowledgements
PSI Center for Photon Science Detectors Group
R. Barten, F. Baruffaldi, A. Bergamaschi, B. Braham, M. Brückner, M. Carulla, R. Dinapoli, S. Ebner, K. D. Ferjaoui, E. Fröjdh, D. Greiffenberg, S. Hasanaj, J. Heymes, V. Hinger, T. King, P. Kozłowski, C. Lopez Cuenca, D. Mezza, K. Moustakas, A. Mozzanica, C. Ruder, K. A. Paton, B. Schmitt, P. Sieberer, D. Thattil, X. Xie, J. Zhang
Fondazione Bruno Kessler
A. Bisht, G. Borghi, M. Boscardin, M. Centis Vignali, F. Ficorella, O. Hammad Ali, G. Paternoster, S. Ronchin

## Slide 27: Backup

## Slide 28: Sensors for soft X-rays: Thin entrance window
Al/Si interface
J. Zhang et al 2022 JINST 17 C11011

## Slide 29: Sensors for soft X-rays: First application
Full-size 4M system at SwissFEL Maloja
V. Hinger et al 2022 JINST 17 C09027
Single-particle diffraction on Xe clusters,
courtesy of K. A. Schnorr
Courtesy of A. Mozzanica

## Slide 30: Application: Resonant Inelastic X-ray Scattering
Method
Scan incoming x-ray beam over electronic resonance
Measure transfered energy and momentum
Impact
Probe complex electronic-structural information
Element, chemical, and orbital site specific
Energy resolution
Spatial resolution
Detector
Spherical grating
Sample
Energy dispersion

## Slide 31: “Strixels”: High resolution in one dimension
75 µm
75 µm
25 µm
225 µm
High spatial resolution by interpolation