## Slide 1: Enhancing spatial resolution in MÖNCH detectors for electron microscopy via deep learning
Xiangyu Xie :: Postdoc :: Paul Scherrer Institute
TREDI 2024, Turino, Feb 20-22, 2024

## Slide 2: Direct Electron Detector for Electron Microscope, the “resolution revolution”
Monolithic Active Pixel Senor (MAPS)
  Widely applied for (cryo-)imaging and diffraction measurements
    K2, Falcon…
  Good spatial resolution
    Pixel size usually
  Limited frame rate
  Poor radiation hardness
Sensor
Electronics
Sensor
Electronics
Hybrid Pixel Detector (HPD)
  Getting popular for diffraction measurements
    Jungfrau, Medipix…
  High dynamic range and frame rate
  Good radiation hardness
  Poor resolution
    Large pixel size usually
    Electron multiple scattering
Chart: 
| Resolution | Radiation hardness | Frame rate | Data quality | Data accessiblity
CCD | 5 | 3 | 4 | 4 | 5
HPD | 2 | 5 | 5 | 5 | 5
Data availability
Resolution
Radiation hardness
Dynamic range
Frame rate
MAPS
HPD

## Slide 3: MÖNCH, towards a universal detector for diffraction and imaging
MÖNCH, a HPD with much smaller pixel
  pitch, pixels per tile
  Charge integration mode, up to
While for multiple scattering
  Poor resolution for with
Deep learning to reconstruct position
  From the complex pattern to learn
    Nature of energy deposition
    Charge diffusion
  Lead an equivalent smaller pixel size
Deep learning?
Chart: 
| Resolution | Radiation hardness | Frame rate | Data quality | Data accessiblity
CCD | 5 | 3 | 4 | 4 | 5
HPD | 2 | 5 | 5 | 5 | 5
HPD+ML | 4 | 5 | 5 | 5 | 5
Resolution
Radiation hardness
Frame rate
Dynamic range
Data availability
MAPS
HPD

## Slide 4: Project members
PSI
Photon Science Division Detector group
PSI
Electron Microscope Facility group
Swiss Data Science Center (SDSC) and
SDSC hub @ PSI
Funded as a Swiss Data Science Center(SDSC) project (link)
Dr. Carl Remlinger

## Slide 5: Project overview
Deep learning development
  Neural network candidate models
    A baseline model (results discussed in this report)
    Auto-encoder
    …
  Samples preparation
    Simulation and measurement
    Key: the ground truth impact point
Characterization
  Spatial resolution
  Detective quantum efficiency
  Standard sample imaging
Deployment & dissemination
  Data processing pipeline

## Slide 6: Baseline neural network model
Input: readout pattern and features
Backbone model for readout patterns
  7 Convolutional Neural Network(CNN)
    Kernel size 3; 64 output channels
  One Fully Connected(FC) layer
Feature model for manually extracted features
  One FC layer
    7 features: cluster size, nominal energy …
Model size: parameters
  Training takes a few hours on a RTX4090 GPU
Optimizer: Stochastic Gradient Descent implemented in PyTorch
Loss function: Squared Error Loss
Readout pattern
CNN
CNN
Features
FC
FC
Impact
point
FC
CNN
…
CNN
Backbone model
Feature model
Simulation for a electron

## Slide 7: Simulation-based: samples analysis
Simulation setup
  Geant4 + charge diffusion model [link]
    silicon sensor
    Start with electrons
  Total samples:
    for training/testing
Simulation analysis
  backscattered electrons
  Charge centroid gives
Charge centroid method
Backscattered electrons

## Slide 8: Simulation-based: training and results
Training setup
  epochs/iterations
  Data augmentation
    Sample rotation and flipping
    Training label smoothing
    Synthetic noise ()
Simulation-based results
  Deep learning gives pixel
    Fully deposited :
    Backscattered
Knife-edge measurements
  Data processed by the trained model
  Super resolution ( physical pitch)
  pixel
Testing dataset
Square Error Loss

## Slide 9: Measurement-based: data taking setup
EM: Jeol JEM ARM200F
  beam
  Customized alignments
  Perpendicular narrow beam ()
Detector: MÖNCH v0.3
  One tile; pixels
Ground truth impact points obtaining
  EM randomly scans with longer dwell time ()
  scan points frames
  Unbiased ground truth position:
    With negligible uncertainty
Overlay of the scanning

## Slide 10: Experiment-based: sample analysis
Frame level
  Mask malfunction region
  Pedestal subtraction
  Energy conversion
Cluster finding and selection
  Adjacent pixels over a threshold of
  Outlier filtered
Total samples:
  Charge centroid gives
Charge centroid method

## Slide 11: Experiment-based: deep learning results
Training setup
  epochs
  Data augmentation
    Sample rotation and flipping
    Training label smoothing
Experiment-based results
  Deep learning give
    Fully deposited :
    Backscattered
Knife-edge measurements
  Data processed by the trained model
  Super resolution ( physical pitch)
Testing dataset
Square Error Loss

## Slide 12: Preliminary MTF and DQE results ()
Electron flux

## Slide 13: Results from the measurement data
TEM sample with continuous carbon and gold nanoparticles
With deep learning, we can see the ring corresponding to atom arrangements
Charge centroid
Deep learning

## Slide 14: Summary and outlook
Significant improvements in spatial resolution on MÖNCH, via deep learning
Promising performance for energies lower than
Outlook
  Further detector calibration
  Further simulation development
  Data processing pipeline
MÖNCH spatial resolution | Charge centroid | Deep learning + super resolution | 
 |  | Simulation based model | Measurement based model
[ pixel] | 1.83 | 0.66 | 0.60
MÖNCH spatial resolution | Charge centroid | Deep learning + super resolution | 
 |  | Simulation based model | Measurement based model
 | 1.83 | 0.66 | 0.60
DOI 10.1088/1748-0221/19/01/C01020
Notes: E.g., pixel-wise gain and pedestal

## Slide 15: Backup: early-stage simulation-based results

## Slide 16: Backup: Fitting function for ESF
Double-Erf term
  Corresponding to different behaviors of the backscattered electrons and fully-deposited electrons
Exponential term
  To model the bump around the edge due to single-electron-cluster selection

## Slide 17: Backup: Deep learning workflow
Neural network model
Training samples
Loss calculation & backward propagation
Samples consist of readout pattern and truth impact position
Model structure designed in prior;
Parameters get updated during training
Loss: distance between predictions and labels;
Loss updates model parameters via gradient descent
Neural network model
Testing samples
Results compared with label
Training samples
Loss calculation & backward propagation