## Slide 1
Experience on the interferometry and polarimetry for fusion power plants
7th International Conference Frontiers in Diagnostic Technologies (ICFDT7)
ENEA Frascati, 22nd Oct 2024
Alexandru Boboc PhD CEng CPhys

## Slide 2
Outline
Few words on JET Far Infrared Diagnostic(FIR)
Strategy on diagnostics development approach
How to …. ?
Examples
Conclusions

## Slide 3
Intro - JET Far InfraRed Interferometer/Polarimeter Diagnostic System
Hybrid interferometer and polarimeter
Essential for JET Plasma Operation
Measures electron plasma density and magnetic fields via Faraday Rotation angle
Number of channels 10 (8 probing the plasma)
Four lateral channels have double passage due to in-vessel mirrors
Uses Far Infrared Lasers(FIR) (wavelength 100 microns 200mW power)
Number of optical components ~ 500 over 3 buildings
Optical path length 80 meters x16 branches
Largest single component: Diagnostic tower (14 metres high, 70 tonnes weight)
Operation time: 16 hr/day, mostly unattended
Age: ~40 years old
Why Far Infrared (THz) ?
Plasma frequencies (GHz) => plasma is a transparent medium for FIR
Note 1: Interferometry is based on plasma optical refractive index
Note 2: Polarimetry is based on plasma optical activity and birefringence

## Slide 4
Lessons learned from FIR Interferometry/Polarimetry on JET alone
JET Far Infrared Diagnostic Status before Deuterium-Tritium Experiment Campaigns no. 2 and no. 3(DTE2 & DTE3)
Triple redundancy for laser availability
Triple redundancy on measurements for density control
Double redundancy in signal detection + hot spare
Monitoring from remote designed for non-expert (e.g. Shift Technicians, Diagnostic co-ordinators, Responsible Diagnostic Experts)
Three Data Acquisition Systems with 100% channel reliability in terms of signal quality (C40, PowerPC and FPGA architecture)
Spare parts/Ancillary equipment commissioned/checked and ready to use/swap

## Slide 5
How to design power-plant diagnostics based on past experience?

## Slide 6
Lesson learned
Learn from the mistakes done in last 70 years
Learn from the positives done in last 70 years
Developments were multi-generational

## Slide 7
Systems developed to maximise physics output and not reliability
Prototype, one-off, one measurement - one diagnostic
Long time for developing a system (small team) => “old design” when required
Current development approach*
“Comfortable” approach of copying existing validated systems
Care for meeting the minimum requirements only => no built-in resilience
Not an agile approach with not much innovation => risk of obsolescence
No clear direction with respect Human Resources required for developing, commissioning, operating the system
Safety considered mostly locally within the system or close area
Quality control and tracking to various levels not unitary with the entire plan
* The list is not exhaustive

## Slide 8
Plasma duration from seconds to non-stop
Neutron fluences higher order of magnitudes (up to 7 !)
No human access to most of the parts of the machine for decades
Current machine versus Fusion power plants*
* The list is not exhaustive

## Slide 9
Fusion power plants approach
Fully engineering approach
Integrated with other plants and machine design
Strong Safety Systems
From day one of development…..
Diagnostics are for control and monitoring – Physics is optional
Clear parameters to control and what measurements (sensors) are required
Each measurement will have a functional role for machine control, monitoring and protection

## Slide 10
Think how to keep it alive

## Slide 11
Plasma
Partial access
Non-stop operation
Turn-key devices
Intermediate area
Access every 5-15 years
Very limited manual entry
Zero Manual access
Bioshield
Vacuum vessel - No access Parts embedded in first wall
Bioshield
Plasma
Full access
Day operation
Full access when not in plasma operation
Easy access for manual emergency
Regular access allowed
Bioshield
Vacuum vessel –parts replaceable during shutdown
Bioshield
MAST-U
EAST
ASDEX-U
WEST
Plasma
Partial access
Day operation
Up to 8 hr shift
Semi-automation
Full access during long shutdowns
Easy access for people entry in an emergency
Partial access
Bioshield
Vacuum vessel –parts replaceable during shutdown
Bioshield
JET
ITER
SPARC?
STEP
BEST
DEMO
ARC
Access considerations
Think how to keep it alive
Nobody
done it yet !
Interspace area
Diagnostic room
Torus(reactor) hall

## Slide 12
Diagnostic room
FIR Lasers
Ancillary Equipment
Electronics
HV Power supplies
Cryogenics
Pressure systems
Hot surfaces
Service areas
Access
Basement
Opto-mechanics
Access
Dry air
Torus Hall
In-vessel
Mirrors (not replaceable)
Vacuum-Windows (replaceable once every 10 years or so)
Bioshield
Bioshield
Think how to keep it alive
Lessons learned from FIR Interferometry/Polarimetry on JET alone
Tritium barrier
Tritium barrier
Nickel in-vessel mirror hiding
behind limiter 30 years later
Ex-Vessel
Diagnostic-Tower
Access
Opto-mechanics
TPX 90mm thick pellicle windows acting as tritium barrier to the basement
A failed cryostat will take minimum one week to recommission
Losing dry air requires 3 days for restoring signal power as basement duct account for 50m(80 in total) of the optical path (rate of air 5l/min for a 25mx4mx2m box)
Interspace area
Diagnostic room
Torus(reactor) hall

## Slide 13
Think about signal losses and errors
Lessons learned from FIR Interferometry/Polarimetry on JET alone

## Slide 14
Think about signal losses and errors
Laser power and power stability
Laser pointing stability
Polarisation quality & settings
Risk associated with operation/maintenance
Long term operation cost
Lasers
Better a low power but stable
Essential for 100m optical path
Unnecessary losses 100% avoidable
High Voltage, heat, pressure, laser
Time required for recovery after fault
Cost of spares, people time, number of people
Lessons learned from FIR Interferometry/Polarimetry on JET alone

## Slide 15
Vacuum Window transmission
Tritium barrier losses
Waveguides
Environmental controls
In-vessel reflectors
Transmission(T) losses
T of new double vac window: 75%, 56% after 30 years in visible, T for FIR 24% day one (4cm thick z-cut quartz)
Total transmission level(24%x50%x50%x50%) = > 3 %T in the ideal case scenario so 97% losses !!!
JET FIR system
T 50%
T 50%
Reflection 100% when new (Nickel mirrors)
unknown now
T difficult to predict let’s assume T 50% here
Lessons learned from FIR Interferometry/Polarimetry on JET alone
Think about signal losses and errors
Channel 5 (L1) 31-years old in-vessel mirror (Survey from 2024)
(“Tokamakium” deposits)

## Slide 16
Depolarisations due to the transmission lines
Hot air turbulence (3 mm laser – first compensation laser on JET FIR )
Moisture in air (all FIR lasers)
Electronics signal distortion(capacitors dying after 10 years)
Windows under high magnetic field – new thing in ITER and beyond
Bad or degrading alignment
Phase shift errors -very hard to model and eliminate
Lessons learned from FIR Interferometry/Polarimetry on JET alone
Think about signal losses and errors

## Slide 17
Wrong wavelength - DCN laser emits in 190 and 195micron leading with 2.5% error if wrongly selected
Polarimetry Torus hall encoders(potentiometers) degradation leading to wrong conversion Voltage-Faraday Angle =>solved by regular offline calibration (remote)
Polarimetry wrong half-wave plate position – Signal-to-noise issue
Mutual interaction of Faraday Angle- Cotton Mouton effects (all papers use cold plasma approximation)
1-D Calibration (Faraday- linear polarisation) for 2-D real problem (Faraday and Cotton-Mouton –elliptical polarisation)
Calibration errors
Lessons learned from FIR Interferometry/Polarimetry on JET alone
Think about signal losses and errors

## Slide 18
All big systems require alignment, in power-plants environment that could be possible only partially every decade - E.g. JET FIR System
Integrating alignment scheme at early design phase is crucial
Integrate alignment requirement with rest of the plant (including shared space with other diagnostics for example)
Choose the most resilient parts where there is no access
Use good practices and good old examples
Choose right components
Alignment
Lessons learned from FIR Interferometry/Polarimetry on JET alone
Think about signal losses and errors
Opto-mechanics
support plate
f12mm location pins
with 50mm clearance
Torus Hall non-magnetic mirror assembly
Find new means (some very cheap)
to align
25m straight beamline visible laser
(2-3mm => 0.005 degrees error)
Basement optics
Bad practice
“Spacer” made of M8 nuts and
washers pointing down
Half-wave plate
assembly

## Slide 19
Think for what do you use the measurements
Gas feedback and control of plasma density in real-time
Additional heating-systems safety interlocks in real-time
Reference diagnostic for many systems (E.g. High-Resolution Thomson Scattering)
Physics analysis (80% of Physics Data depends on interferometry)
Interferometry
Safety Factor control, including real-time (via Equinox)
Gas feedback and control of density (backup)
Additional heating-systems safety interlocks(backup)
Physics analysis (EFIT constrains)
Secondary measurement for density validation
Polarimetry
Lessons learned from FIR Interferometry/Polarimetry on JET alone

## Slide 20
Think what changes to the machine impact to your measurements
Example- Divertor introduction impact to the FIR interferometer:
Vacuum window 60mm diameter
Ch 2 slot in divertor plate
9mmx 44mm x 1 m
FIR Beam
diameter
35mm
Additional Line-integrated density component
in the divertor not included in EFIT
Ch1 tangent
to plasma
2 channels lost for any H-mode pulse with NBI on

## Slide 21
*Extract from STEP Interferometer/polarimeter preliminary study
Essential for plasma breakdown and ramp-up control for scenario of plasma forming close to inner limiter (very low density, high resolution required)
First external adjustable reflector ~10 meters from plasma
Key problem
1cm of beam displacement due to refraction over 25m requires an angular precision of tracking of 0.001 degrees or 17 microradians
What is the function of your measurement ?
Example: STEP Interferometer Tangential Channel 4 *
Tangential plane (Z=0)
Ch 4 FIR laser beam (RED colour) close to central column (~10cm)
Distance between centre of machine to input mirror ~10 meters
~10m from centre of
machine to exit mirror

## Slide 22
A word about RAMI
RAMI (Reliability, Availability, Maintainability and Inspectability) is one of the most severe challenges of Fusion Technologies, and it is probably the engineering branch where the fusion community is the least competent.
Fusion people know/want this*
What is required is this*
*Extract from ITER Dispersion Interferometer RAMI Analysis Report

## Slide 23
Magnetic information
Magnetic sensors/coils
Polarimetry
Fiber Optical Current Sensor (FOCS)
Real-time profile controller (AI driven)
STEP
Think about measurement (not-diagnostic) redundancy
System requirements
Integration with other diagnostics
Wall protection
Visible and Infrared cameras
Pyrometers
Laser diffusion spectroscopy
Thermocouples
Real-time controller (Machine Learning)
Electron Density
Interferometry
Polarimetry
Microwave diagnostics
Visible spectroscopy
Realtime profile controller (AI driven)
RAMI for measurement ?

## Slide 24
Scenario: Catastrophic failure of in-vessel internal magnetic sensors for vertical stability due to radiation damage of cables/coils
Polarimetry is one possible viable alternative
EAST
Works
nicely !
Try again with no in-vessel reflectors and no toroidal symmetry of laser beams !
STEP
How to survive ? - example #1
BEST
Will probably work
(horizontal channels only)

## Slide 25
Conclusions
Suggestions for development for power plants -class diagnostics based on facts and experience
Synthetic diagnostics accounts of substantial development time at first phase  feasible solutions for control
1:1 prototyping for part of the system is a must  risk mitigation
Turn-key laser sources and detectors(for laser-based diagnostics)  needs developments
Zero access for most of parts of the machine  No in-vessel optics (reflectors)/electromechanics
3D-symmetry of sensors(no poloidal symmetry)  very hard physics reconstruction  complicates real-time control
Control mechanisms components must be far away from plasma  mm accuracy beam tracking or cameras views via optical fibres

## Slide 26
UKAEA supports new diagnostics developments and share lessons learned
Plasma Physics and Controlled Fusion
Special Issue on the Physics and Engineering of Toroidal Fusion Plasma Operations
Plasma Science and Fusion Operations Division new initiative
<< DICE >>
Diagnostics Innovation Centre
of
Excellence
Operational since September 2024
Neutronics and Gamma Lab
Magnetics Lab
Optics Lab with five optical workstations areas
https://iopscience.iop.org/article/10.1088/1361-6587/ad5376
https://iopscience.iop.org/journal/0741-3335/page/Special%20Issue%20on%20the%20Physics%20and%20Engineering%20of%20Toroidal%20Fusion%20Plasma%20Operations

## Slide 27
Many people have worked with the FIR team, and it has been a tremendous collaborative effort, so I would like to thank to entire JET Team over the years.
I would like to remember here my predecessors and their deputies: Didier Veron, George Magyar, Jan de Hass, Perry Beaumont, Klaus Guenter, Paulo Puglia and the technicians Geoff Braithwaite, Dave Peach, Paul Trimble.
Interferometer/polarimeter operations in the JET DTE3 campaign would not have been possible without my colleagues James Macdonald, Mark J Brown, the electrical engineers Will Studholme and Simon Cramp.
And last, but not least, we would like to thank for our long-term collaborators F.Orsitto, C.Gil, B.Bieg, A.Murari, D.Mazon and M.Gelfusa.
JET, which was previously a European facility, is now a UK facility collectively used by all European fusion laboratories under the EUROfusion consortium. It is operated by the United Kingdom Atomic Energy Authority, supported by DESNZ and its European partners. This work, which has been carried out within the framework of the Contract for the Operation of the JET Facilities up to 31 October 2021, has been funded by the Euratom Research and Training Programme. Since 31 October 2021, UKAEA has continued to work with the EUROfusion Consortium as an Associated Partner of Max-Plank-Gesellschaft zur Förderung der Wissenschaft e.V represented by Max-Plank-Institut fur Plasmaphysik (“IPP”) pursuant to Article 9.1 of the EUROfusion Grant Agreement for Project No 101052200. The views and opinions expressed herein do not necessarily reflect those of the European Commission.
Acknowledgments

## Slide 28
Many thanks
Any questions ?
Email: Alexandru.Boboc@ukaea.uk
FUSIONICS