## Slide 1
Review on CoTS boards
Mario De Lucia
Università di Pisa, INFN – Sezione di Pisa

## Slide 2: Pisa
PISA LAB:
Low temperatures Group with over 10 members across all backgrounds and career stages
VIRGO

## Slide 3: A background as small as possible…

## Slide 4: Overview:
Why?
What for?
What do we need?
What’s available?

## Slide 5: Overview:
Who? + Why?
What for?
What do we need?
What’s available?

## Slide 6
Meandered
inductor
Interdigitated
capacitor
Microwave Kinetic
Inductance Detectors

## Slide 7: Microwave Kinetic Inductance Detectors
Main interesting features:
Single Photon Sensitivity
Energy resolving detectors
µs Time resolution
No Dark Count Rate
Inherent multiplexibility (FDM)
Easy to fabricate

## Slide 8: Microwave Kinetic Inductance Detectors
Applications for BULLKID DM
~160 pixels per wafer / Stack of 10 wafers
Energy resolution as low as possible (27±6 eV) *
~ 0.8GHz resonance frequency likely (0.5GHz – 1GHz)
µs Time resolution
No Dark Count Rate
* BULLKID: Monolithic array of particle absorbers sensed by Kinetic Inductance Detectors Cruciani A., et al. 2022

## Slide 9: Microwave Kinetic Inductance Detectors
Applications for the Cold Veto System
~1 pixel per each scintillator / number and geometry of scintillators TBD
Energy resolution as low as possible (61±6 eV) *
~ 2300MHz resonance frequency likely (2GHz – 4GHz)
µs Time resolution / long. Comparable to
No Dark Count Rate
* CALDER: cryogenic light detectors for background-free
searches., L. Cardani et al. 2015

## Slide 10: Readout Electronics:
Philosophy
What’s Available?
Pros and Cons - Table

## Slide 11: Readout Philosophy
On the board
Mezzanine
Up and down mixing goes here
if necessary

## Slide 12: Readout Philosophy
On the board
Mezzanine
Up and down mixing goes here
if necessary
Generate and
read tones

## Slide 13: Readout Philosophy
On the board
Mezzanine
Up and down mixing goes here
if necessary
Channelisation Coarse + Fine
: Polyphase Filter Bank :

## Slide 14: Readout Philosophy
On the board
Mezzanine
Up and down mixing goes here
if necessary
Digital Down Conversion Downsampling

## Slide 15: Readout Philosophy
On the board
Mezzanine
Up and down mixing goes here
if necessary
Maths, Pulse Triggering

## Slide 16: Overview:
Who? +Why?
What for?
What do we need?
What’s available?

## Slide 17
Board name | Logic Blocks | DSP Slices | ADC | DAC | Unit Price (€) | FPGA pixel count | ADCDAC Pixel count | Cost per pixel (max) | Comments
ROACH1 | 7,360 | 640 | 2x 12 bit 550 MSPS | 2x16 bit 1000 MSPS | 4000 | 256 | 512 | 16€ | Prod. Disc.
ROACH2 | 326,400 | 1,120 | 2x 12 bit 2.5 GSPS | 2x 16 bit 2.8 GSPS | 4000 | 1,000 | 1,400 | 4€ | Prod. Disc.
UniboardII | 4.6M | 1,518 | N/D | N/D | 20,000+ | 1,200 | N/D | 16.7€+ | +extra DAC/ADC
VCU118 | 2.586M | 6,840 | N/D | N/D | 5,800+ | 5,600 | N/D | 1.00€+ | +extra DAC/ADC
ZCU111 | 930,000 | 4,272 | 8 x 12 bit / 4.096 GSPS | 8x14 bit 6.554 GSPS | 8,995 | 3,500 | 8,000 | 2.25€ | Redundant +free firmware
4x2 | 930,000 | 4,272 | 4 x 12 bit / 4.096 GSPS | 2x14 bit 6.554 GSPS | 2,149* | 3,500 | 2,000 | 1.00€ | +free firmware *Xilinx edu only
ZCU216 | 930,000 | 4,272 | 16 x 14 bit/ 2.5 GSPS | 16 x 14 bit / 9.85 GSPS | 15,546 | 3,500 | 20,000 | 4.44€ | One board To rule them all
VP430 | 930,00 | 4,272 | 8x 12-bit, 4GSPS | 8x 14-bit, 6.4GSPS | ~15,000 | 3,500 | 8,000 | 4.30€ | 
AMC599 | 1.451M | 5,520 | 2 x 12 bit 6.4 GSPS | 2 x 16 bit 12 GSPS | 20,000 | 4,500 | 3,200 | 6.25€ | 
t0 | 930,000 | 4,272 | 8x 14 bit 5 GSPS | 8x 14 bit 9.85 GSPS | 10/60k | 8,000 | 8,000 | 1.3/7.5€ | Plug & play Support / Dear
USPRx300 | 326,080 | 840 | 2x 14 bit 200 MSPS | 2x 16 bit 800 MSPS | 10,000 | 700 | 200 | 500€ | Quite limited expensive

## Slide 18
Board name | Logic Blocks | DSP Slices | ADC | DAC | Unit Price (€) | FPGA pixel count | ADCDAC Pixel count | Cost per pixel (max) | Comments
ROACH1 | 7,360 | 640 | 2x 12 bit 550 MSPS | 2x16 bit 1000 MSPS | 4000 | 256 | 512 | 16€ | Prod. Disc.
ROACH2 | 326,400 | 1,120 | 2x 12 bit 2.5 GSPS | 2x 16 bit 2.8 GSPS | 4000 | 1,000 | 1,400 | 4€ | Prod. Disc.
UniboardII | 4.6M | 1,518 | N/D | N/D | 20,000+ | 1,200 | N/D | 16.7€+ | +extra DAC/ADC
VCU118 | 2.586M | 6,840 | N/D | N/D | 5,800+ | 5,600 | N/D | 1.00€+ | +extra DAC/ADC
ZCU111 | 930,000 | 4,272 | 8 x 12 bit / 4.096 GSPS | 8x14 bit 6.554 GSPS | 8,995 | 3,500 | 8,000 | 2.25€ | Redundant +free firmware
4x2 | 930,000 | 4,272 | 4 x 12 bit / 4.096 GSPS | 2x14 bit 6.554 GSPS | 2,149* | 3,500 | 2,000 | 1.00€ | +free firmware *Xilinx edu only
ZCU216 | 930,000 | 4,272 | 16 x 14 bit/ 2.5 GSPS | 16 x 14 bit / 9.85 GSPS | 15,546 | 3,500 | 20,000 | 4.44€ | One board To rule them all
VP430 | 930,00 | 4,272 | 8x 12-bit, 4GSPS | 8x 14-bit, 6.4GSPS | ~15,000 | 3,500 | 8,000 | 4.30€ | 
AMC599 | 1.451M | 5,520 | 2 x 12 bit 6.4 GSPS | 2 x 16 bit 12 GSPS | 20,000 | 4,500 | 3,200 | 6.25€ | 
t0 | 930,000 | 4,272 | 8x 14 bit 5 GSPS | 8x 14 bit 9.85 GSPS | 10/60k | 8,000 | 8,000 | 1.3/7.5€ | Plug & play Support / Dear
USPRx300 | 326,080 | 840 | 2x 14 bit 200 MSPS | 2x 16 bit 800 MSPS | 10,000 | 700 | 200 | 500€ | Quite limited expensive

## Slide 19
Board name | Logic Blocks | DSP Slices | ADC | DAC | Unit Price (€) | FPGA pixel count | ADCDAC Pixel count | Cost per pixel (max) | Comments
ROACH1 | 7,360 | 640 | 2x 12 bit 550 MSPS | 2x16 bit 1000 MSPS | 4000 | 256 | 512 | 16€ | Prod. Disc.
ROACH2 | 326,400 | 1,120 | 2x 12 bit 2.5 GSPS | 2x 16 bit 2.8 GSPS | 4000 | 1,000 | 1,400 | 4€ | Prod. Disc.
UniboardII | 4.6M | 1,518 | N/D | N/D | 20,000+ | 1,200 | N/D | 16.7€+ | +extra DAC/ADC
VCU118 | 2.586M | 6,840 | N/D | N/D | 5,800+ | 5,600 | N/D | 1.00€+ | +extra DAC/ADC
ZCU111 | 930,000 | 4,272 | 8 x 12 bit / 4.096 GSPS | 8x14 bit 6.554 GSPS | 8,995 | 3,500 | 8,000 | 2.25€ | Redundant +free firmware
4x2 | 930,000 | 4,272 | 4 x 12 bit / 4.096 GSPS | 2x14 bit 6.554 GSPS | 2,149* | 3,500 | 2,000 | 1.00€ | +free firmware *Xilinx edu only
ZCU216 | 930,000 | 4,272 | 16 x 14 bit/ 2.5 GSPS | 16 x 14 bit / 9.85 GSPS | 15,546 | 3,500 | 20,000 | 4.44€ | One board To rule them all
VP430 | 930,00 | 4,272 | 8x 12-bit, 4GSPS | 8x 14-bit, 6.4GSPS | ~15,000 | 3,500 | 8,000 | 4.30€ | 
AMC599 | 1.451M | 5,520 | 2 x 12 bit 6.4 GSPS | 2 x 16 bit 12 GSPS | 20,000 | 4,500 | 3,200 | 6.25€ | 
t0 | 930,000 | 4,272 | 8x 14 bit 5 GSPS | 8x 14 bit 9.85 GSPS | 10/60k | 8,000 | 8,000 | 1.3/7.5€ | Plug & play Support / Dear
USPRx300 | 326,080 | 840 | 2x 14 bit 200 MSPS | 2x 16 bit 800 MSPS | 10,000 | 700 | 200 | 500€ | Quite limited expensive
In-house custom developed readout board: full flexiblity probably more expensive

## Slide 20
Jenny Smith – UCSB - 2024
(Xilinx RFSoC ZCU111 – 2000 Channels FDM)