## Slide 1: 3D image reconstruction for Boron Dose distribution in BNCT with CZT-based Compton camera
February 26-28, 2024 @Trento Institute
Sesto Congresso TIFPA
nicola.ferrara@ba.infn.it
dayron.ramos@ba.infn,it

## Slide 2: INDEX
BNCT bases and dosimetry
Sensor for dose imaging
Image reconstruction method and simulations
Results using the iterative method
Alternative method of source reconstruction using image processing
First SPECT set-up for dose tomography

## Slide 3: Boron neutron capture therapy
Capture reaction: 10B(n,α)7Li
HADRON THERAPY INSIDE THE
SINGLE CELL
dD(x,y,z) ≈ nB10(x,y,z) Φ(x,y,z) dV

## Slide 4: Boron neutron capture therapy
SELECTIVITY
at single cell level
DISTRIBUTION of 10B

## Slide 5: Dosimetry on BNCT
Dtotal =
DB +
Dp +
Dn +
D𝛄
In BNCT we must correctly account for four separate radiation components
From 1H(n,𝛄)2H (E=2.2 MeV, 𝛔=0.33 b) & reactor background
Due to epithermal and fast neutron elastic scattering daily with H nuclei
From 14N(n,p)14C reactions (Ep=630 keV, 𝛔=1.9 b) due to thermal neutron
Therapeutic dose from 10B(n,⍺)7Li (𝛔=3837 b)

## Slide 6: CZT sensors for dose imaging on BNCT
Room-temperature gamma-ray spectroscopic
Sub-millimetre spatial resolution and excellent energy resolution (<1% FWHM at 661.7 keV)
drift strip detector
drift cell: a collecting anode strip + 2 drift strips on the right
side (RD) + 2 on the left side (LD) + 10 cathode strips
CZT single unit: 20x20x5 mm3,
planar transversal field (PTF),
orthogonal drift strip electrodes

## Slide 7: Compton imaging: single stage detector
front plane SCATTER detector
axes W from segment
(interaction-point 2 – interaction-point 1)
back-projection cone
Back plane ABSORBER detector

## Slide 8: Image reconstruction method
most probable source distribution
Iterative method
Maximum Likelihood Expectation Maximisation (MLEM)
UCT_MLEM – CUDA Compton Camera Image Reconstruction Code
https://github.com/mattcleigh/Compton_MLEM
λnj = calculated amplitude of pixel j at the nth iteration
sj = sensitivity, i.e. the probability that a gamma ray originating from pixel j is detected anywhere
Tij = imaging response matrix, i.e. the transition probabilities generated by the measured events(first estimation: based on back-projection, λ0)

## Slide 9: Simulation set-up
Detector
CZT crystal stack (5mm thickness each)
Phantom material
Air, soft tissue
Data
Hits positions, energies dep, replica number

## Slide 10: Algorithm commissioning
On Tissue phantom
I = 50
On Air phantom
Back-projection
I = 10
I = 50
I = 250

## Slide 11: Spheric source reconstruction
offset = 30 mm
offset = 60 mm
offset = 100 mm
Z-axis stretching
On Tissue phantom, 60 mm Offset, T = 5.0
I = 250
I = 50
Back-projection
I = 250
I = 50
Back-projection
On Tissue phantom, 60 mm Offset, T = 2.0

## Slide 12: Alternative source reconstruction method
Morphological filters
Erosion and dilation filters
to back-projection image
First limits approximation of the source
Dilation
Spur pixel deleted
Hole filling
Erosion and dilation cycles
Source reconstruction

## Slide 13: Alternative source reconstruction method
Morphological filters
Reconstructed sources examples
Spheric source on Air
Spheric source on
Tissue (T=5)
5 point-like source on Air
Spheric source on
Tissue (T=2)
Real source
Source reconstructed
Overlap

## Slide 14: Deep learning reconstruction
U-net architecture:
Directed Acyclic Graph
U-net Convolutional neural network: layer of convolution, pooling and upsampling
Residual: take difference between input and output

## Slide 15: Deep learning reconstruction
U-net architecture
Spheric source on
Tissue (T=5)
5 point-like source on Air
Spheric source on
Tissue (T=2)

## Slide 16: Outlook
Tomography reconstruction:
Set 4 module CZT
Collimator pin-hole
OSEM reconstruction

## Slide 17: Conclusions
                Sensor CZT good in reconstruction of boron dose
                Test of three reconstruction methods
                What’s next?
                Measure with detector at BNCT facility
                Build of a tomography reconstruction of dose

## Slide 18: THANKS!
nicola.ferrara@ba.infn.it
dayron.ramos@ba.infn.it