## Slide 1: QMA for monitoring and precise measurements of ...
Gaseous polarised nuclear targets consist of H (or D), the residual gas also in UHV
are mainly H2, and the injected atoms recombine in molecules, difficult to pump.
... atoms and molecules contents
in gaseous nuclear polarised target
accumulated in a storage cell,
... populations of quantum states of polarised atomic beams and nuclear targets.
Chopping beam techniques are needed, in order subtract
background (BG) gases.
TARGET GAS ANALYSER
TGA
BREIT-RABI POLARIMETER
BRP
Ciullo Giuseppe (work done with Statera Marco et al.)

## Slide 2: Some uses (experiments) of polarised nuclear target
Study of contribution of the spin of subnuclear
particle to the total spin of the nucleons.
Polarised positron collinding with
nuclear polarized targets
HERMES
HERA ring at DESY (Hamburg)
HERA Beam
Gases are injected in an accumulation open cell on the accelerated beam axis
We inject gases which are difficult to pump, H2, D2, He and N2, Kr, Xe too.
Target thickness
ttarget [at cm-2]
Intensity of projectiles
Ip [part s-1]
  R = s L
  s = cross section, L= luminosity [cm-2 s-1]
  For accelerated beam on fixed targets:
L=Ip ttarget
L [cm-2 s-1]

## Slide 3: The HERMES polarized nuclear target
Target/year | H¦¦(1997) | D¦¦(2000) | H(2003)
Pt | 0.851 ± 0.033 | 0.845 ± 0.028 | 0.795 ± 0.033
t(1014 nucl/cm2) | 0.7 | 2.1 | 1.1
FOM (P2t) | 0.5 | 1.5 | 0.7
The diagnostics of the nuclear polarized target are performed by the TGA (Target GAs Analyzer) and the Breit-Rabi Polarimeter, which are equipped both with two QMA (Quadrupole Mass Analyser)
Impressive performance of daily working system, parassitic target, just stopped for accelerator maintenance.
ABS
BRP
HERA Beam
HERA Beam

## Slide 4: Upgrading for filtering studies of antiprotons test on p 2005-2017 COSY (FZJ) – proposal for AD (CERN)
Tuning of the whole target,
monitoring of its perfomance
diagnostics and measurements
of atomic, molecular fraction and polarization
performed by two cross beam QMA equipped with baffles and
choppers for a good
BG subtraction and also monitoring vacuum conditions
and openable cell performances.
ABS
Target
chamber
Breit-Rabi
polarimeter
COSY/AD
beam axis
Electric power
Cooling water
Target Gas Analyser

## Slide 5: Easily moved on the beamline properly designed for the installation at COSY and at AD
Atomic Beam Source
Six additional quadrupoles:
4 from COSY
2 from CELSIUS
Target Gas Analyser
and
Breit-Rabi Polarimeter
Target chamber:
Detector system + storage cell
Momentum: <3.7 GeV/c
l Circumference: 183 m
l Polarized proton and deuteron

## Slide 6: Nuclear Polarised Target – PAX SCHEME
ABS
Target
chamber
Breit-Rabi
polarimeter
COSY/AD
beam axis
Electric power
Cooling water
QMA at the end of the Target Gas Analyzer
QMA at the end of the Target Gas Analyzer
TGA Chopper for background subtraction
The Breit-Rabi Polarimeter allows to monitor, to tune the RF transitions and finally to measure the population of quantum states and provides the polarisation of the sampled beam.
BRP chopper for background subtraction
The Target Gas Analyser (TGA) allows to
monitor and measure atomic and molecular
composition of sampled beam from the
accumulation cell, it allows also to
distinguish between the backgroung (BG) from the
TGA chamber itself, from the Cell Chamber,
and ballistic flow from ABS.
Production of
nuclear
polarized H or D
Accumulation
Diagnostics

## Slide 7: WHY an accumulation cell? Increasing ttarget
ingoing at/s=outgooing at/s
Free atomic beam (Target)
Accumulated atomic beam (Target)
accelerated beam
Thickness of the target of the free beam tb
1.5 1014 At cm-2
Integrated thickness (tc) of the target along the cell
2.6 1011 At cm-2
Elliptical cell
Free atomic beam
= cm s-1
d
Accumulated atomic beam

## Slide 8: From macroscopic to microscopic quantities
Throughput Q:
Stationary condition p cost: =p S (1)
Q12=(p1 – p2) C (2)
(1) Macro ; micro=n S
(2) Macro =(n1kT – n2kT)C; micro: =(n1 – n2) C
mbar l s-1
Transitory cost V (vessel) Q =- V = pS

## Slide 9: Microscopy derivation of Master equations
A1
A2
Stationary condition
Effective flow torwards less molecular density:
As a result we derive:

## Slide 10: If Q is the conserved quantity in vacuum systems
Qcal
{
Ccal
we need the measurement of CCal .
Pcal can be measured with high accuracy (0.15 %) with a baratron, but
Q calibration Unit is a required tool

## Slide 11: Measuring Pcal
Ccal measured by
    Evacuation of Vcal through Ccal
We need the Measurement of Vcal ?

## Slide 12: Measuring V cal
From the relation PV=constant
I need a reference volume Vref , measured (ex situ)
By filling it with liquid (ethanol) Vref .
In situ, by iterated expansion of gas (H2) in Vref,
Is possible to measure Vcal
We can measure the pressure in Vref, by the gauge in Vcal
When both volume are in comunication (valve open),
Then we close the valve between them, we pump down Vcal and
then we can expand again .

## Slide 13: But model indipendent starting from the definition
Then get the calibrated throughput at at fixed geometry of the connection between the
Vcal and the cell (Ccal depends on pressure pcal) in the viscous regime.
From the definition
which applied to the evacuation
of the Volume Vcal becomes:
from

## Slide 14: Test on a Q cal
Measurement precision ≈ 0.5 %
Reproducibility ≈ 0.5 %

## Slide 15: How you can measure IABS?
QABS=(pcv – pch)Ccv injecting ABS
Qcal=(pcv – pch)Ccv calibrated injection
IABS from corresponding Qcal I: IABS= Qcal/kT

## Slide 16: Unpolarized gas feeding system (HERMES)

## Slide 17: SMOG 2 injection upstream LHCb

## Slide 18: Cell thickness seen by the LHCb detector

## Slide 19: It is possible to measure the conductance of the cell
The measured pressure, is not actually on the center of the cell, but with a gauge,
connected with bellow, valves and capillary, and the model, which allow has to
estimated the cell, come out from the stationary system distributed throughput,
which I’ll remind in the next slide.

## Slide 20: Stationary systems: distributed throughput

## Slide 21: Checking if the cell close well
Measuring CCell
and

## Slide 22: Target polarisation
PT = total target polarization
a0 = atomic fraction in absence of recombination
ar =atomic fraction surviving recombination
Pa = polarization of atoms
Pm = polarization of recombined molecules
Relation to measured quantities, trough sampling corrections
    ar = ca arTGA
    Pa = cP PaBRP
    arTGA
    measured by cross beam QMA
    PaBRP
    measured by cross beam QMA
    with the substraction of BGs with
    chopping beam techiniques.
arTGA
    PaBRP

## Slide 23
The diagnostics system vacuum scheme
TGA: NEG GP500 MK5
BRP: NEG inside Ion-pump
Scroll Pumps
Initial scheme improved thanks to the characterization
of vacuum performances ad commissioning
of new pumping systems
the Ti-Ball in the TGA Chamber has been
exchanged with a NEG pump (CF150)
Also the Ion pump in the BRP has been equipped with a small NEG (CF40) inside it.
The two ecodry piston pumps, exchanged with two scroll
diode

## Slide 24: The diagnostics system inner components scheme

## Slide 25: The atoms and molecules counting devices
RF module
Amplifier and Electronics QME 100
only for calibration with Faraday cup
CEM
Channeltron
QMA CROSS BEAM
with 90° off axis CEM
(Channel Electron Multiplier)
the ion optics and signal processing have to be under control and continuous cross check, for
proper calibrations and quantitative estimation.

## Slide 26: Spectrometer QMA 400 Pfeiffers Balzers
Extractor
HV1
The whole filament powering,
the electrons and ions optics,
and the detector powering
and DAQ are controlled via
DAC for piloting any module PC Software
mainly with 0-10 VDC logic,
and some module also GPIB, RS-232, RS-485
The read-out of the piloting
parameter
are performed via ADC
via software.
U1
U2
U3
U4
or as Faraday Cup
(with U5 grounded)
U5
HV1
HV2
MOE
Extractor

## Slide 27: We need to check the stability of our tuning for the QMAs
HV1 ®
U5

## Slide 28: Tuning of ion optics parameters
TGA
BRP
Compromises between higher signals
and good stabilities.
For example in the tuning on TGA of m/q =1 and
and m/q=2, we choose the red one.
On BRP the setting for m/q =1 is on the max with a plateau.

## Slide 29: Target gas analyser
Scheme of the sampling of the target: atom and molecules diffuse towards the end of the extension tube and enter the TGA. A chopper to subtract the background. Baffles prevent atoms to hit the inner wall of ionisation volume of QMA, both tilted 7° from sampled beam axis, allowing the sampling of the beam toward the BRP
Picture of the inside of the TGA chamber,with QMA and chopper:
on the left the beam blocker, the chopper baffle and the fork chopper.
In the back the housing of the titanium sublimation pump.

## Slide 30: Schemes of data acquisition and signal processing
Two types of signals:
one system which acquires QMA pulses with fast counter (DAQ), and
a second system which records analog voltages with an Analog to Digital Converter (ADC) (SLOW control).
The DAQ and the SLOW control software is programmed with the programs of the package Labview running on a standard personal computer.
QMAs pulses are amplified by fast amplifier at 1 m from QMAs.
The amplifier pulses are discriminated and converted into NIM pulses, sent to a shielded connector block (NI SC-68) an acquired by a 32-Bit counter card (NI PCI-6602).
The setup enables pulses counting for both QMAs up to one MHz.

## Slide 31: Scheme of DAQ
On TGA
On M1 100 Cycle averaged,
On M2 Average on 300 Cycles

## Slide 32: The different uses of the QMA in TGA and BRP
The QMAs installed in the TGA chamber and the BRP chamber are used in different mode of operation, let’s distinguish in
CONTINOUS MODE
  monitor vacuum and procedure of its preparation and improvements, then confirm other kind of pumping, for istance adding NEG pumps in both chambers, observing the proper setting of pumping, in order to avoid backstreaming.
CHOPPED MODE peculiar for tuning and monitoring atomic beams
  Tune the atomic beam source,
  Tune the Breit-Rabi Polarimeter
and more sofisticated
  counts atoms and molecules coming from the cell in a sampled beam,
  determine the relative population of quantum states to provide the polarization of the sampled beam, which requires different calibrations and also montecarlo simulations.

## Slide 33: Continous mode: vacuum and bake-out effectiveness
Mass scan of the QMA in the TGA Chamber.
Just after the turbo-molecular pumps
reach the regime rotation speed.
m/q (countrate)
18 (570 kHz)
28 (100 kHz)
TGA
1 (100 kHz)
2 (20 kHz)

## Slide 34: I can have confidence on bake-out procedure
Looking in the base
partial pressure of
M2 and M1

## Slide 35: Continous mode: vacuum and bake-out effectiveness
Mass scan on m/q
for tuning of the QMA
and vacuum, partial pressures,
preparation,
test and monitoring
m/q (countrate)
18 (260 kHz)
80 (100 kHz)
Time scan on choosen m/q
In this case m/q=2 on TGA in order to follow the UHV
preparation of its Vacuum
chamber.
m/q (countrate)
2 (6 kHz)
1 (36 kHz)
TGA

## Slide 36: Bake out Elements and Temperatures in the real world
Two modules equipped with 16
channel of 1 kW, and interlocked with
pressure thresholds.
Each channel with PID T control, differential
and thermic switch,

## Slide 37: Continous mode: vacuum and bake-out effectiveness
Time sc on 18
Mass scan on m/q
to see the effectiveness
of the bake out in this case
on the BRP.
Sometime the available time
before running is limited.
BRP
Time scan on a fixed m/q
to see the effectiveness
of the Bake out in this case
on the BRP fixing m/q=18.
one day acquisition.

## Slide 38: Bake-out and conditioning, seasoning ...
m/q (countrate)
18 (180 kHz)
28 (38 kHz)
44 (20 kHz)
___ m/q =2 QMA signal
TGA
Tuning of the proper flahing mode for Ti-BALL, and monitoring of its effectiveness on the partial pressures.
Continous mode:

## Slide 39: Continous mode:
m/q (countrate)
TGA
2 (60 kHz)
18 (11 kHz)
28 (6 kHz)
m/q = 2
Flashing the
TGA Ti-Ball

## Slide 40: Difference after bake-out and help from Cryopumps

## Slide 41: Ti Ball flashing improve very little
BRP

## Slide 42: Back-streaming of He and Capability to pump down
He
TGA

## Slide 43: Solution: NEG pumps mainly for TGA
Our Main problem
is to protect the sextupole
magnets which cannot
overcome temperature
Higher than 70 oC.
These temperatures are in the range of what we get during the bake out of the TGA chamber. The conclusion is that we will need to cool down the flange of sextupole permanent magnets during the regeneration of the NEG pump, as usually is done for the bake-out.

## Slide 44: Measurements of pumping Speed o NEG
Before activating the NEG, preliminarily we measured the pumping speed of the turbo pump alone, which provided a value of Sturbo = 122 + 3 l/s for hydrogen.
After the activaation of the NEG we measured the pumping speed of the NEG and the turbo and we get as a result SNEG+turbo = 910 + 70 l/s.
We also measured the pumping speed of the NEG alone (with the valve on the turbo closed) and got only SNEG = 50 + 20 l/s, much different from the expected value from the simple subtraction;
SNEG= S NEG&turbo – Sturbo ~ 780 + 70 l/s.
NEG SAES Capacitorr D 2000 for TGA
Anyhow for us it was a good results, we
have not only a turbo pump on the TGA,
but also the two cyopumps.
For 10-5 mbar l s-1 sampled beam we expect 10-8 mbar in the TGA.

## Slide 45: Commissioning of NEG PUMPS on TGA and BRP
NEG SAES Capacitorr D 2000 for TGA
GP 50 2F per BRP.
After the installation and activation of the NEG pump in the TGA, opening the valve on the target chamber with the ABS in operation, the measured pressure in the TGA was
PTGA = 1.6 10-8 mbar, meanwhile the target chamber pressure was PTC = 3.5 10-7 mbar.
Comparing previous the value in the October 2010 run, when the TGA chamber was pumped by the TSP werecorded in the range of
PTGA = 2.5 10-7 mbar at the same PTC = 3.5 10-7 mbar.
Within the accuracy and the reproducibility of the gauge installed in the TGA,
it’s evident
that we improved the pumping speed of one order of magnitude.

## Slide 46: Stationary vacuum with cryopumps on
The ion gauge are below their minimum reading,
the cryo-pumps are running and the separation valve
is open.
The cryo-pumps are pumping also the BRP Chamber through a bellow of CF40 10 cm long and a gate valve CF40.
TGA
PTGA < 5 10-10 mbar
PBRP < 5 10-10 mbar

## Slide 47: Evidence of backgrounG from TGA to BRP, not on the ion gauge
Separation Valve
BRP
Continous mode:
Baked
Baked
¬ Not Baked®
Target Valve
PBRP = 3.5 10-9 mbar

## Slide 48: Checking BRP pumping
The BRP is pumped through
the bellow and the valve conductance
by the TGA pumping system
BRP
Separation Valve

## Slide 49: Opening the Target Valve and switching on ABS
2 (40 kHz)
2 (1.6 MHz)
TGA
TGA
Target Valve
Mass scan on m/q
to see the influence of BG
from Target Chamber:
PTGA & PBRP < 5 10-10 mbar
Mass scan on m/q
to see the influence of ABS on
from Target Chamber:
PTGA = 1.6 10-8 mbar
PBRP < 5 10-10 mbar
18 (156 kHz)
18 (108 kHz)
2 (35 kHz)

## Slide 50: Tuning on of mass ... our ROI (region of interest)
m/q
Injecting D2
1
1
2
3
4
18
19
20
17
16
15
14
TGA

## Slide 51: Chopped Mode: monitoring the ABS perfomance
The atomic flux Fa originate esclusively from ABS. From the dissociative ionization in the target can be neglected.
Measured just feeding molecular atoms in the cell
1 % has to be subtracted.
ABS Injection
extension tube
Sampled gas
To TGA
and BRP

## Slide 52
Chopped Mode: monitoring the ABS perfomance
BRP Signal
TGA Signal

## Slide 53: accurate measurements of Atomic and molecular fraction
The measure atomic (molecular) signal Sa (Sm) is proportional
to the corresponding flux Fa (Fm) into the TGA,
the ionization cross-section (),
the detection probability (
the mean inverse velocity of the particles <1/va> (<1/vm>),
will depend on the frienging magnetic field (B) of the target, and on the emission current (Iem) from the electron gun which we indicate as

## Slide 54: Our interest the number of nucleon in the target
For H and D the degree of dissociation measured by TGA using nucleon flow rates is defined as:
aTGA =
where flow rate of nucleons coming from atoms,
flow rate of nucleons coming from molecules.
Flow rate of nucleons are related to atomic and molecular flow rates as follow:
,
.

## Slide 55: From the ... :
for atoms and molecules thermalized on ext. tube

## Slide 56: k calibration
Assuming the ABS flux constant and no change in temperature and current emission (o using the normalized signal S/Iemiss and T correction:
+2
In HERMES changing (on purpose) the T of the cell, we change the recombination then we change for the same flow the Sa and Sm ratio.
Correction are needed with respect to the temperature of the extension tube and emission current.
But we are able to provide average calibration over the runs
for Hydrogen and Deuterium:
k measured during run for Hydrogen 0.5507 + 0.0023
k measured for Deuterium 0.861 + 0.015

## Slide 57: Just showing a tipical k calibration for Deuterium

## Slide 58: We have to distinguish and measure other contributions to molecules counting
Flow of molecules coming from the target chamber frg, through the cell, which can be measured by feeding gas in the chamber of the target at the same pressure read during atomic beam injection.
Ballistic flow from ABS fball, which can be determined by feeding in the cell different states of ABS, the ballist flow will be constant.
C. Baumgarten et al. A gas analyzer for the internal polarized target of the HERMES experiment
NIMA 508 (2003) 268-275

## Slide 59: Quantum state populations:
Breit-Rabi diagram for H and D
we can manage both H and D, for simplicity I’ll present more details for H
H
D
a scheme of hyperfinestates (hps) of H and D

## Slide 60: Quantum state selection (Stern-Gerlach and rot)
With proper RF transition units we can rotate the angular
momentum of the atoms and with a 2nd sextupole

## Slide 61: RF transitions MFT
H
D
Each transition requires proper Bhom, Bgrad, n, and they can be overimposed.
The requirements are different for H and D.
For the magnet we can use the same, for RF generator we have different modules for H and D.

## Slide 62: RF transitions MFT

## Slide 63: RF transitions (SFT)
H
D

## Slide 64: Type of transition depending on oscillating BHF
BHF
BHF

## Slide 65: Dual Cavity for H and D use SFT
Match existing geometry in the BRP, no change in volume occupied by the cavity
Oscillating field (BHF) at 45o respect to the static Field and to beam axis (previously along)
Push-Pull excitation of the two rods, we have
a BHF ortoghonal to the axis of the rods
BHF
BHF ┴
BHF║
Axis of B static and gradient
To excite both
s and
p HF transitions

## Slide 66: Dual Cavity
H. RF in
Site for D RF Pick-up
H tunable rod
H tunable rod
H fixed rod
Site for H RF Pick-Up
D. RF in
H rotable rods
D fixed rods
D rotable rods
lin = 44 mm

## Slide 67: First tuning the Atomic beam Source for the injection of state |1> in the cell.
CHOPPED MODE ABS MFT SCAN - BRP OFF
ABS

## Slide 68: Injecting state |1> and tuning the MFT 13 of BRP
CHOPPED MODE ABS MFT at |1> - BRP MFT 13
BRP
MFT 13
Sextupole Magnests defocus
|3> and |4> states

## Slide 69: Injecting state |1> and tuning the SFT 14 of BRP
CHOPPED MODE ABS inj. |1> - BRP SFT 14
BRP SFT 14
Sextupole Magnests defocus
|3> and |4> states

## Slide 70: Injecting state |2> and tuning the SFT 24 of BRP
CHOPPED MODE ABS |2> - BRP SFT 24
MFT 13
Sextupole Magnests defocus
|3> and |4> states

## Slide 71: BRP SFT - Behavior

## Slide 72: SFT dual Cavity for H and D cavity

## Slide 73: The tuning to be performed for ABS and BRP
this is the simplest for polarized H.

## Slide 74: BRP calibration requires measurements of
Si=S Mia Ia
Measured BRP signals Si are described by a set of linear functions of four hyperfine state intensities Ia
Mia a matrix measurements,which depends on transition efficiencies and transmission probabilities for the hyperfine states.
C. Baumgarten et al. An atomic beam polarimeter to measure the nuclear polarization in the HERMES gaseous polarized
Hydrogen and deuterium target NIMA 482 (2002) 606-618.
efficiencies and transmission probabilities
(11 unknowns for H, 40 for D))

## Slide 75: Polarization measurements injecting state \1>.

## Slide 76: The following slides are Spares, under construction, or still under brainstorming

## Slide 77: BRP vacuum check

## Slide 78: Measurements of the Cell Conductance

## Slide 79: Ionization cross section

## Slide 80: Relative Ionization Probability

## Slide 81: SEM or CEM