## Slide 1: Turbomolecular pump

## Slide 2: Pumps that operate in the Molecular Regime
Notes: 2

## Slide 3: Turbomolecular Pumps – General Info
Turbomolecular pumps are kinetic mechanics pumps where rotors at high rotational speed transfer momentum to particles present in vacuum system.
High and Ultra High Vacuum generation
Different models for high compression ration and high throughput
≤ 8 x 10-8 mbar (ISO F) or ≤ 5x 10-10 mbar (CF)
The always need forevacuum pumps
Coating of optical elements, instruments and display, architectonic glasses
Analytic technology, mass spectrometry, R&D
90 to 2100 l/s (325 to 7560 m3/h)
Principle of operation
Pros
Typical applications
Pumping Speed
Ultimate pressure

## Slide 4: Turbomolecular Pumps
Turbopumps are axial compressors designed for pumping gasses in the molecular flow regime
Operating range 10-2 to 10-10 Torr
Pumping speed 10 to 3,000 l/s
Infinite pumping capacity
Turbopumps are throughput pumps – meaning they have infinity capacity
Blade rotation peed ranges from 14,000 to 90,000 rpm – making them mechanical vulnerable
Notes: 4

## Slide 5: Principle of Operation
In molecular flow, particles are not «sucked» but they move freely in the available volume;
When a gas particle hits a rotors, an impulse is transferred to the particle in direction favorable for the pumping.
Therefore, rotors need to have speed near to the mean thermal speed of the gas particles;
The mean free path of the gas particles has to be larger than the distance rotor-stator.
Static Surface
Rotating Surface

## Slide 6: Theory of Turbo Pumping Effect
v
c
mean velocity
wall velocity

## Slide 7: Pompe turbomolecolari: principio di funzionamento/1
Impulse of the rotor blade
Stator
Impulse of the rotor blade
Thermal speed
Notes: 7

## Slide 8: Turbomolecular pump: principle of operation/1
Impulse of the rotor blade
Stator
Impulse of the rotor blade
Thermal speed
Adsorption, sojourn, desorption
Notes: 8

## Slide 9: Turbomolecular pump: principle of operation /3
Notes: 9

## Slide 10: Turbomolecular pump: principle of operation /4
The pumping effect (composition of speed) is significant only when the peripheral speed of the blades is of the
Balde speed is of the same order of magnitude as the average speed of the molecules that are to be pumped (at room temperature)
Notes: 10

## Slide 11: Turbomolecular pump: principle of operation/5
Axial compressor type pumps are very flexible design
  # stages can be varied
  Blade angles varied
  Hybrid pumps
Molecular flow exists through most of the a turbopump; however, transient and sometimes viscous flow occurs at the pump discharge
The key parameter of turbopumps is compression ratio, not Dp
Notes: 11

## Slide 12: Principle of operation

## Slide 13: Turbomolecular pump: key parameters
Important parameters to consider:
  Pumping speed for different gases;
  Maximum throughput for different gases;
  Compression ratio (capability to maintain a given pressure difference between inlet and outlet);
  Maximum counter pressure;
  Final pressure.
Notes: 13

## Slide 14: Turbomolecular Pumps
Notes: 14

## Slide 15: Compression ratio (1/3)
Compression ratio
Compression ratio with zero inlet flow
Probability of inlet into the turbo stage
K0
vb / cw
Using a Monte Carlo method, Kruger & Shapiro calculated K0 as function of different parameters
Throughput

## Slide 16: Compression ratio(2/3)
Exponential dependence on
K0
vb / cw
Chart: Mass Dependence
| 5.2915026221291814 | 2 | 1.4142135623730951
 | 200000000 | 600 | 200
Pompa TURBOVAC SL 80 C

## Slide 17: Compression ratio(3/3)
K0
vb / cw
Gas Inlet (Vacuum side)
Initial stages with large angle blades to increase pumping speed
Gas Outlet
Final stages with limited angle blades to increase compression ratio

## Slide 18: Pumping Speed
Velocità di pompaggio
Behaviour from geometrical consideration
Chart: Pumping Probability PHo
| 0.1 | 10 | 20 | 30 | 40 | 50 | 60
vb/vp = 0.5 | 0.0017423 | 0.14604 | 0.24322 | 0.30217 | 0.32994 | 0.32994 | 0.30217
vb/vp = 1 | 0.0034785 | 0.25485 | 0.39128 | 0.4641 | 0.49617 | 0.49617 | 0.4641
Pumping speed constant in a large range
Pumping speed decreases in the rasniion region where the flow changes from molecular to transient
Flusso attraverso un area

## Slide 19: Main components
Rotation speed in range 14’000 a 90’000 rpm
Precision bearings and axis balance.
Case resistant to rotor crash. Carefull in supporting the pump.
Holweck stage used also to separate bearing with labyrinths (see later slide).
Ball bearings may be:
  Oil lubricated
  Grease lubricated
Possibility for a magnetic suspended rotor
Hybrid configurations available
Water cooling only for large loads, otherwise air cooling is sufficient
High Vacuum Flange
Stator
Venting Valve
ForeVacuum Flange
Inlet net
Rotor
Case
Ball Bearings
Water Inlet
Ball Bearings
Motore a 3 fasi
TURBOVAC 151

## Slide 20: Ball Bearings and Magnetic Suspension
Good compatibility with particles due to flowing lubricants
Vertical installation only possibility
Low maintenance
Stainless ball bearing, grease lubricated
Installation with any orientation
Compatible with moving chart
Infinite lubrication
Not available for large pumps
Attention to heating up
Magnetic Control
1,2: Radial electromagnets
3: Radial sensors
4: Axial electromagnets
5: Axial sensor
6: Current amplifier
7: Control Unit
8: Rotor
No friction, no maintenance
Completely hydrocarbon free
Low vibrations
Installation with any orientation
Stainless ball bearing, oil lubricated

## Slide 21: Mechanical and Magnetic Bearings

## Slide 22: Further considerations
For magnetic suspended rotors
  In case of power supply deficiency, the pump acts as a generator and slowly stop the rotor on emergency bearings.
  In an abrupt venting event, the motor act as a dumper and the rotor stops on emergency bearings
  In case of an unforeseen cable removal, the motor suddenly stops, and the pump stops immediately.
To reach the ultimate pressure <5x10-10 mbar it is neccessary to bake the pump and use metal seals.
Corrosive gas pumping requires special attention:
  Injection of purging gases
  Protective coating
  Heating of the low vacuum region

## Slide 23: Purging gas usage

## Slide 24: Hybrid Turbomolecular pump magnetic levitated rotor
Holweck Stage
Magnetic Bearings
Notes: 24

## Slide 25: Holweck Stage
Gas particles are adsorbed on rotating surfaces. The moving surface push the gas particles in a long spiral channel.
The wall surface adds to the speed distribution of the gas particles, creating a net pumping effect.
Throughput depends on the inverse of the rotor to stator distance (gap 0.3-0.5 mm).
Stable material required for standing the centrifugal forces and temperature variations. Typical is a carbon fiber cylinder.
transition region near to molecular flow
cm mbar (air at 20°C)
mbar
After a x100 the stage is not more effective

## Slide 26: Holweck Stage

## Slide 27: Turbomolecular pump with magnetic suspension
Active 1-axis magnetic bearing system
Active 5-axis magnetic bearing system
Notes: 27

## Slide 28: Turbo pumps: technical specifications
Notes: 28

## Slide 29: Turbo pumps: compression ratio
Notes: 29

## Slide 30: Turbo pumps: wide range /1
Horizontal spyral
Notes: 30

## Slide 31: Turbo pumps: wide range/2
Vertical Spyral
Notes: 31

## Slide 32: Turbo pump wide range: technical specifications
Notes: 32

## Slide 33: Turbo pump wide range: technical specifications
Notes: 33

## Slide 34: Turbo pumps: technical specifications
<
Notes: 34

## Slide 35: Turbo pumps: technical specifications
Notes: 35

## Slide 36: Turbo pumps: typical system
High Vacuum Valve
Turbo Venting Valve
TMP (TurboMolecular Pump)
Primary Pump
Leak Detector Valve
Vacuum Chamber
Notes: 36

## Slide 37: High Vacuum Pumps (Diffusion and TMP) : Primary Pump Evaluation
Q = p A * Seff
Seff
p V
= Q [mbar l s-1]
p A
S V
Q
Q
Q
Q = p V * SV
S V =
p A * Seff
p V
Notes: 37

## Slide 38: High Vacuum Pumps (Diffusion and TMP) : Primary Pump Evaluation
Fore vacuum pumping speed is calculated with the following formula
S V =
p A * Seff
p V
= Q [mbar l s-1]
In a nutshell, we need to guarantee to the high vacuum pump a pressure on the outlet enough low also at the maximum throughput of the high vacuum pump.
Notes: 38

## Slide 39: High Vacuum Pumps (Diffusion and TMP) : Primary Pump Evaluation
Gases or vapors pumped by the high vacuum pump MUST be transported also by the forevacuum pump that operates at higher pressures.
It is MANDATORY NOT overcome the maximum pressure on the high vacuum pump.
For the turbo pump it is possible an overload of the motor or its overheating that induces shutdon of the pump.
If Q is the effective quantity of gas or vapor that the high vacuum pumps removes at the effective pumping speed Seff at pressure pa, the throughput of the primary pump, with effective pumping speed Sv at pressur pv is equal to Q (nothing is created, nothing is destroyed), Q is conserved! Then it holds
Q = pa × Seff = pv × Sv
Notes: 39

## Slide 40: Forevacuum pump dimensioning (1/2)
High Vacuum Valve
Turbo Venting Valve
TMP
Primary Pump
Leak Detector Valve
Vacuum Chamber

## Slide 41: Forevacuum pump dimensioning (1/2)
Maximum Throghput
QpV = 25 mbar lt/s @ PIn = 0.1 mbar
Reduced Throughput
QpV = 7.16x10-2 mbar lt/s @ PIn = 10-4 mbar
Chart: TURBOVAC 850i
Chart: TRIVAC
| 1000 | 630.95734448019311 | 398.10717055349699 | 251.18864315095823 | 158.48931924611142 | 100 | 63.095734448019307 | 39.810717055349691 | 25.11886431509582 | 15.84893192461114 | 10 | 6.3095734448019298 | 3.9810717055349696 | 2.5118864315095824 | 1.5848931924611143 | 1 | 0.63095734448019314 | 0.39810717055349737 | 0.25118864315095801 | 0.1584893192461114 | 0.1 | 6.3095734448019303E-2 | 3.9810717055349734E-2 | 2.5118864315095798E-2 | 1.5848931924611141E-2 | 0.01 | 6.3095734448019329E-3 | 4.04592007885996E-3
TRIVAC D 8 B | 2222.222 | 1288.82 | 808.3157 | 513.2993 | 326.1444 | 207.11 | 128.1351 | 80.634 | 51.374 | 32.0635 | 20.0085 | 12.4798 | 7.8014 | 4.8758 | 2.9983 | 1.846 | 1.1156 | 0.66143 | 0.38285 | 0.21557 | 0.11889 | 0.063909 | 0.033317 | 0.0166 | 0.0076846 | 0.0030886 | 0.00083822 | 0
TRIVAC D 40 B | 12500 | 7504.9443 | 4688.2947 | 2996.7641 | 1897.8728 | 1194.7099 | 755.7549 | 477.6488 | 300.0963 | 187.6418 | 119.4362 | 75.7714 | 47.7532 | 29.9126 | 18.6298 | 11.6282 | 7.2711 | 4.4066 | 2.6661 | 1.692 | 1.0239 | 0.60451 | 0.35475 | 0.20177 | 0.11339 | 0.060873 | 0.030978 | 0.014858
TRIVAC D 8 B PFore = 10-1 mbar
TRIVAC D 16 B PFore = 2 mbar

## Slide 42: Turbomolecular pump: characteristics
High cost
Mechanics critical
No risk for venting
It operates in any position
Need for a forevacuum pump
SHe high
Low Compression Ratio for low light gases
Infinite capacity
Between two service life> 10 000 hours
There is NO giant TMP (max 3000 l/s)
Water: not always necessary
Hydrocarbon contamination: no, if properly used
Final vacuum: 10-10 mbar
Vibration: low, really low for magnetic suspended
Notes: 42