## Slide 1: Introduction to Quadrupole Mass Spectrometry
2024-06-12
Valerie Derpmann, Norbert Müller

## Slide 2: Quadrupole Mass Spectrometry
Basics of Quadrupole Mass Spectrometry
  Ionization
  Mass separation
  Detection
Data acquisition modes
Applications
Exercises

## Slide 3
Total and Partial Pressure in vacuum chambers

## Slide 4: Total and Partial Pressure
Partial Pressure Analysis informs on
residual gases at base pressure incl. outgassing rates
air leaks
Helium leaks (leaktest )
composition of process gases
impurities in process gases

## Slide 5: MS basics
What does a quadrupole mass spectrometer do?
A mass spectrometer produces charged particles (ions) from the chemical substances that are to be analyzed.
The quadrupole mass spectrometer then uses electric fields to measure the mass (“weight” or more precisely the Mass/charge ratio, m/z or m/e) of the charged particles.
The charged particles that pass the mass spectrometer are collected, their charge is converted to a current and measured.

## Slide 6: MS basics
What does the mass tell us? Let us use water (H2O) as an example. A water molecule consists of two hydrogen's (H) and one oxygen (O). The total mass of a water molecule is the sum of the mass of two hydrogen's (approximately 1 atomic mass unit per hydrogen) and one oxygen (approximately 16 atomic mass units per oxygen): 2 H: 2 amu (atomic mass units) + O: 16 amu ___________ = H2O: 18 amu
1 amu = 1.66E-27kg
16O ~ 16 amu
= Electron ~ 0 AMU
= Neutron ~ 1 AMU
= Proton ~ 1AMU

## Slide 7: MS instrument set up

## Slide 8: Ion Source
and in general:
Faraday-Cup
Electron Multiplier
Ion-Source
Mass-Filter
Ion Counting

## Slide 9: MS basics
Ion source
A filament or cathode emits electrons which – on their path to the more positively charged formation chamber / Anode– collide with gas particles and knock out electrons. In this process the gas particles become positively charged, they are ionized.
An Extraction lens extracts the ions from the formation chamber, and focusing lenses ensure ideal conditions at the entrance to the mass filter.

## Slide 10: Introduction into Quadrupole Mass Spectrometry
I
Filament
+ -
+
Emission
e
-
U
Electron-Enenergy =
U
[eV]
M + e
-
---> M
+
+ 2e
-
The emission is a function of the filament temperature, the material (work-function), the applied voltage and the geometry (local field strength).
The filament is heating up the ion source by radiation, and can so increase the outgassing rate
Cathode
Anode
- Ion Source for Electron Impact Ionization

## Slide 11: MS basics – H2O ionization and fragmentation
Let's suppose that we put some water vapor into the ion source of the mass spectrometer. If we shoot a beam of electrons through the water vapor, some of the electrons will hit water molecules and knock off an electron. If we loose a (negatively charged) electron from the (neutral) water molecule, the water will be left with a net positive charge. In other words, we have produced charged particles, or "ions" from the water:
H2O + 1 (fast) electron --> [H2O]+ + 2 electrons
Some of the collisions between the water molecules and the electrons will be so hard that the water molecules will be broken into smaller pieces, or "fragments". For water, the only possible fragments are [OH]+, O+ and H+.

## Slide 12: MS basics - H2O
The mass spectrum of water will show peaks that can be assigned to masses of 1 (H+), 16 (O+), 17 (OH+), und 18 (H2O+) m/e.
The relative abundance and the fragmentation pattern are like a finger print of a substance.

## Slide 13: Introduction into Quadrupole Mass Spectrometry
Ionization produced by electron impact, as a function of the electron energy

## Slide 14: Introduction into Quadrupole Mass Spectrometry
Ionisation Probability compared to N2
The sensitivity of any mass-spectrometer
and ionization gauge is gas-specific!

## Slide 15: Introduction into Quadrupole Mass Spectrometry
Molecule + electron  Ionization of molecule  Fragmentation of molecule  Multiple charged ions  Isotopic effects are shown
Example: CO2 + e- (70eV)  CO2+ 44 m/e  C+ 12 m/e  O+ 16 m/e  CO+ 28 m/e  CO2++ 22 m/e  12C 16O 18O 46 m/e

## Slide 16: Introduction into Quadrupole Mass Spectrometry

## Slide 17: Introduction into Quadrupole Mass Spectrometry

## Slide 18: Introduction into Quadrupole Mass Spectrometry

## Slide 19: Introduction into Quadrupole Mass Spectrometry

## Slide 20: Introduction into Quadrupole Mass Spectrometry

## Slide 21: Measurement of pure Argon
Information from spectra
library

## Slide 22: Pure Argon - or not???
Analog spectra from MS, linear display

## Slide 23: Pure Argon - or not???
Analog spectrum from MS, logarithmic display

## Slide 24: Pure Argon - Quality 5.6
In the MS we see more then just a peak at mass 40. Why???
36, 38: Argon has stable Isotopes with Mass 36 and 38
20: Argon is partially double ionized in the ion source and appears as Ar++
The following peaks are system related (background):
18: Water vapor from the vacuum chamber, seals, etc.
14, 28: Nitrogen from the vacuum chamber, seals, etc.
16, 32: Oxygen from the vacuum chamber, seals, etc.
44: Carbon dioxide from the vacuum chamber, seals, etc.
1,2: Hydrogen from the vacuum chamber and from water vapor
All these are just traces in the level of <100ppm!

## Slide 25: What do we find in air???
1, 2: H2 from water vapor
4: He
7, 8: N++, O++
12: C von CO2, CXHY...
14, 15: 14N, 15N
16: O
17: HO
18: H2O
20: Ar
28: N2, CO
29: 14N15N
30: NO
32: O2
34: 16O18O
36, 38: 36Ar, 38Ar
40: Ar
44: CO2
82, 83, 84, 86: Kr
129, 131, 132: Xe

## Slide 26: Possible errors
The last slide showed the most simple interpretation.
In detail it can be much more difficult.
16: Could be CH4 instead of O
17: Could be NH3 instead of HO
20: Could be Ne instead of Ar++
29: Could be C2H5 (Ethanol) instead of 14N15N
36: Could be HCl instead of 36Ar
44: Could be N2O instead of CO2 Knowledge about the sample and possible reactions are necessary for an exact interpretation of the measurements. Mass spectra are reproducible, spectra libraries (e.g. Wiley, NIST) can help with the identification of unknown substances.
Nevertheless it should be noted that spectra vary from one mass spec to another and also by specific MS settings.

## Slide 27: List of fragments part 1
1 H+ H2 2
H+ H2O 18, 17, 16
H+ CxHy 12, 13, 14, 26, 27 etc.
2 H2+ H2 1
He++ He 4
4 He+ He 2
6 C++ CO 12, 28, 29
C++ CO2 12, 28, 44
C++ CxHy 12, 13, 14, 26, 27 etc.
7 N++ N2 14, 28, 29
8 O++ O2 16, 32, 34
O++ H2O 16, 17, 18
12 C+ CO 28, 29
C+ CO2 28, 29, 44
C+ CxHy 13, 14, 26, 27 etc.
13 CH+ CxHy 12, 14, 26, 27 etc.
14 N+ N2 28, 29
N+ NH3 15, 16, 17
CH2+ CxHy 12, 13, 26, 27 etc.
CO++ CO 28, 29
Add. masses
Possible origin
Fragment
Masse (m/e)
15 CH3+ CxHy 12, 13, 14, 26, 27 etc.
NH+ NH3 14, 16, 17
16 O+ O2 32, 34
O+ H2O 17, 18
CH4+ CH4 12, 13, 14, 15
NH2+ NH3 14, 15, 17
17 OH+ H2O 16, 18
NH3+ NH3 14, 15, 16
18 H2O+ H2O 16, 17
19 F+ F2 38
F+ HF 20
20 HF+ HF 19
20Ne+ 20Ne 22
Ar++ Ar 40
22 CO2++ CO2 6, 12, 28, 29, 44
22Ne+ 22Ne 20
24 C2+ CxHy 12, 13, 14, 26, 27
26 C2H2+ CxHy 12, 13, 14, 24, 27
27 C2H3+ CxHy 12, 13, 14, 24, 26
28 N2+ N2 7, 14, 29
C2H4+ CxHy 12, 13, 14, 24, 26
CO+ CO 6, 12, 29
CO+ CO2 6, 12, 29, 44
29 C2H5+ CxHy 12, 13, 14, 24, 26, 27, 28
14N15N+ N2 7, 14, 28

## Slide 28: List of fragments part 2
30 C2H6+ C2H6 12, 13, 14, 24, 26, 27, 28, 29
NO+ NO 14, 16
31 C2H2OH+C2H3OH 26, 28, 29
32 O2+ O2 8, 16, 34
S+ H2S 33, 34
S+ SO2 34, 64
34 H232S+ H2S 32, 33
34S + H2S 32, 33, 34
16O18O+ O2 8, 16, 32
35 35Cl+ Cl2 37, 70, 72,74
35Cl+ HCL 36, 37, 38
36 H35Cl+ HCl 35, 37, 38
36Ar+ Ar 20, 38, 40
37 37Cl+ Cl2 35, 70, 72, 74
37Cl+ HCl 35, 36, 38
38 H37Cl+ HCl 35,36,37
C3H2+ CxHy 12, 13, 14, 24, 26, 27, 28
38Ar+ Ar 20, 36, 40
39 C3H3+ CxHy 12, 13, 14, 24, 26, 27, 28, 38
39K+ K 41
40 Ar+ Ar 36, 38
C3H4+ CxHy 12, 13, 14, 24, 26, 27, 28, 38, 39
41 C3H5+ CxHy 12, 13, 14, 24, 26, 27, 28, 38, 39
41K+ K 39
42 C3H6+ CxHy 12, 13, 14, 24, 26, 27, 28, 38, 39
43 C3H7+ CxHy 12, 13, 14, 24, 26, 27, 28, 38, 39
C2H3O+ C2H5OH 31, 44, 45, 46
44 C3H8+ C3H8 41, 42, 43
CO2+ CO2 6, 12, 28, 29
C2H4OH+ C2H5OH (alcohol) 31, 43, 45, 46
N2O+ N2O 14, 16, 28
45 C2H5O+ C2H5OH (alcohol) 31, 43, 44, 46
13CO2+ CO2 6, 12, 28, 29, 44
46 NO2+ NO2 14, 16
C2H5OH+ C2H5OH (alcohol) 31, 43, 44, 45
48 SO+ SO2 32, 64
55 C4H7+ CxHy 12, 13, 14, 24, 26, 27, 28, 38, 39
57 C4H9+ CxHy 12, 13, 14, 24, 26, 27, 28, 38, 39
58 (CH3)2CO+,C3H6O (acetone) 43
64 SO2+ SO2 32, 48
77 C6H5+ Phenyl 50, 51, 52
78 C6H6+ C6H6 (benzene) 50, 51, 52
92.5 185Re++ Rhenium (filament) 93.5, 185, 187
93.5 187Re++ Rhenium (filament) 92.5, 185, 187
149 Phthalic ester (softening agent)
182 182W+ Tungsten (filament) 183, 184, 186
183 183W+ Tungsten (filament) 182, 184, 186
184 184W+ Tungsten (filament) 182, 183, 186
185 185Re+ Rhenium (filament) 187
187 187Re+ Rhenium (filament) 185

## Slide 29: Spectra library
Most RGA software packages contain a spectra library
Additional spectra are available from the NIST library https://webbook.nist.gov/chemistry/

## Slide 30: Introduction into Quadrupole Mass Spectrometry
Axial Ion-Source
Mass-Filter
Filament
e
-
Ion
+
-
+
-
Extraction Lense
Cross Beam Ion-Source
Mass-Filter
Filament
e
-
Ion
+
Ion
+
e
-
Molecular Beam
Residual Gas Analysis
beam analysis
Extraction Lense

## Slide 31: Introduction into Quadrupole Mass Spectrometry
Mass-Filter
Ion
+
Filament
Extractionlens
Grid-
Ion-Source:
Filament
Filament
Elektrons
Ions
10
-2
mbar
Closed
Ion-Source:
(SPM)
Process Gas Analysis, p>1E-4mbar
Ultra-High-Vacuum
Mass-Filter,differentially pumped

## Slide 32: Residual Gas Analyzer
Transpector MPH

## Slide 33: Differentially pumped analyzer
Approx. 3E-6mbar
DN40CF
Transpector CPM3
SPM 800

## Slide 34: SPM 700 Ion Source
Ion Formation Chamber (interal volume open to process) Filaments (hidden)
Extraction Lens
Focus Lens
Base plate, connection to Quadrupole

## Slide 35: SPM 700 Ion Source
Grid =
Formation Chamber
Filament baseplate =
Filament
Wehnelt =
Extraction
Focus
Housing =
Quadrupole rod
Faraday =
Pressure = 1E-2mbar
Pressure = 5E-6mbar

## Slide 36: SPM 700 Ion Source Potentials
SEM

## Slide 37: SPM 700 Ion Source Potentials
All voltages are measured with «Ion Reference» as reference
«Cathode» defines the energy of the electrons
«Field axis» is the potential on the axis of the Quadrupole
«Field Axis» defines energy / speed of ions in the Quadrupole e.g. IonRef= 150V and Field Axis = -6V means Field Axis =144V from ground

## Slide 38: SPM 700 Ion Source
Flight paths of electrons
Quadrupole rod system
Entrance aperture
Focus
Extraction
Filament
Formation chamber
Faraday

## Slide 39: SPM 700 Ion Source
Flight paths of electrons

## Slide 40
SPM 700 Ion Source
Flight paths of Ions
Quadrupole rod system
Entrance aperture
Focus
Extraction
Formation chamber
Filament
Faraday

## Slide 41
SPM 700 Ion Source
Flight paths of Ions

## Slide 42: Axial ion source, open
Characteristics
  High sensitivity
  Robust, simple design
  Easy to clean
  Good ion injection into the QMA
  Single filament Re, W, Y2O3-Ir
Typical applications
  General gas analysis
  Residual gas and outgassing analysis at system with base pressure >1E-10mbar

## Slide 43: Axial ion source, open
Wehnelt
Filament
Anode / Ionref
Entrance to quadrupole
Focus

## Slide 44: Axial ion source, electrodes and potentials

## Slide 45: Axial ion source
Flight paths of electrons
Quadrupole rod system
Entrance aperture
Focus
Formation chamber
Filament
Wehnelt

## Slide 46: Axial ion source
Flight paths of electrons

## Slide 47: Axial ion source
Flight paths of positive ions
Quadrupole rod system
Entrance aperture
Focus
Formation chamber
Filament
Wehnelt

## Slide 48: Axial ion source
Flight paths of positive ions

## Slide 49: High Sensitivity (HS) ion source, open
Characteristics
  open design
  Fast reaction on gas composition changes
  High sensitivity
  Long lifetime
  Good ion injection into the QMA
  Double filament W, Y2O3-Ir
  Closed version available
Typical applications
  Residual gas analysis
  Quantitative and qualitative gas analysis
0000-00-00
Presentation Title Goes Here
49

## Slide 50: High Sensitivity (HS) ion source, open
0000-00-00
Presentation Title Goes Here
50

## Slide 51: Ion Source Potentials

## Slide 52: High Sensitivity ion source
Flight paths of electrons
Quadrupole rod system
Entrance aperture
Focus
Grounded electrode
Extraction
Formation chamber
Filament
0000-00-00
Presentation Title Goes Here
52

## Slide 53: High Sensitivity ion source
Flight paths of electrons
0000-00-00
Presentation Title Goes Here
53

## Slide 54: High Sensitivity ion source
Flight paths of positive ions
Quadrupole rod system
Entrance aperture
Focus
Grounded electrode
Extraction
Formation chamber
Filament
0000-00-00
Presentation Title Goes Here
54

## Slide 55: High Sensitivity ion source
Flight paths of positive ions
0000-00-00
Presentation Title Goes Here
55

## Slide 56: HS ion source, variable electron energy
Characteristcis
  variable electron energy from 10 eV to 150eV
Typical applications
  „low price“ appearance potential determination
  analysis of radicals
0000-00-00
Presentation Title Goes Here
56

## Slide 57: Grid ion source
Characteristics
  open design
  high sensitivity
  low degassing rate
  easy to degas
  double W- filament
Typical applications
  residual gas analysis in UHV / XHV
  desorption analysis (TDMS)

## Slide 58: Grid ion source, electrodes and potentials

## Slide 59: Grid ion source
Formation chamber
Quadrupole rod system
Entrance aperture
Filament
Flight paths of electrons

## Slide 60: Grid ion source
Flight paths of electrons

## Slide 61: Grid ion source
Flight paths of positive ions
Quadrupole rod system
Entrance aperture
Filament
Formation chamber

## Slide 62: Grid ion source
Flight paths of positive ions

## Slide 63
Introduction into Quadrupole Mass Spectrometry
Filament materials
Rhenium: rather high vapor pressure used mainly in High Vacuum
Tungsten used in the UHV, has a longer lifetime than Rhenium
Tungsten gets brittle at high Hydrocarbon concentrations
Yt-Oxide coating on Iridium wire rather resistant against air, temperature is much lower compared to pure metals

## Slide 64
Quadrupole Mass Filter
and in general:
Faraday-Cup
Electron Multiplier
Ion-Source
Mass-Filter
Ion Counting

## Slide 65
The quadrupole
Mass filter Four precisely aligned rods form an electric quadrupole field.
RF and DC voltages applied on the two rod pairs allow for separation of the charged particles according to their mass over charge ratio.
Introduction into Quadrupole Mass Spectrometry

## Slide 66
Introduction into Quadrupole Mass Spectrometry
A superposition of an AC and a DC electrical field is applied onto the quadrupole rods.
Ions with an appropriate ratio of mass to charge achieve stable trajectories in the
high frequency field and pass the filter.
E 1
E 2
E 3
E 4
-1.2
0
1.2
-1.2
0
1.2
2 r
o
+(U + V cos ω t) -(U + V cos ω t)
Mass Filter - The Quadrupole-Field

## Slide 67
Introduction into Quadrupole Mass Spectrometry
Electrical potentials at one point in time

## Slide 68
Mass Filter - The Quadrupole-Field, variations over time
Electrical field at t=0
Electrical field at t=0+1/2
Introduction into Quadrupole Mass Spectrometry

## Slide 69
Introduction into Quadrupole Mass Spectrometry
Use of the transformations
gives the following differential equations (Mathieu‘s equations)
Equations of motion

## Slide 70
Introduction into Quadrupole Mass Spectrometry
q =
4 * e * V
0.2
0.1
0.2 0.4 0.6 0.8 1.0
q
a
0.23699
0.706
a =
8 * e * U
m * r
0
2
* ω
2
m * r
0
2
* ω
2
stable region
instable region
instable region
Mass Filter – Stability-Diagram

## Slide 71
Introduction into Quadrupole Mass Spectrometry
The ratio a/q (which is proportional to U/V) is kept
constant and the absolute values of U and V are
increased.
For a/q< 0.1678 (which is the top of the stability-diagram) the stability region is entered periodically for the different masses.
The slope a/q determines the mass resolution.
Mass Filter – Mass-Tuning

## Slide 72
Introduction into Quadrupole Mass Spectrometry
The ratio a/q (which is proportional to U/V) is kept
constant and the absolute values of U and V are
increased --> A mass spectrum will result

## Slide 73
Mass Filter - The Quadrupole-Field, force on an ion
Electrical field at t=0
Electrical field at t=0+1/2
Introduction into Quadrupole Mass Spectrometry

## Slide 74
Quadrupole Mass Spectrometer
Flight path of ions through the mass analyzer
 Ion mass identical to set mass

## Slide 75
Quadrupole Mass Spectrometer
Flight path of ions through the mass analyzer
 Ion mass higher than set mass

## Slide 76
Quadrupole Mass Spectrometer
Flight path of ions through the mass analyzer
 Ion mass lower than set mass

## Slide 77
Introduction into Quadrupole Mass Spectrometry
16 mm rod system for highest resolution, stability and transmission (e.g. He/D2 separation)
8 mm rod system for High-End Applications
6 mm rod system for common analytical purpose (e.g. standard RGA)
Mass Filters

## Slide 78
Introduction into Quadrupole Mass Spectrometry
Mass-Filter, dynamic range, air analysis
≙7.8ppb

## Slide 79
Introduction into Quadrupole Mass Spectrometry
Unit Resolution: All peaks are equally broad over the whole mass range (one amu) and
10% valley
between two neighbouring peaks of equal heigth
Resolution := M/ M
With unit resolution the
resolution increases with
increasing mass number
mass
intensity

## Slide 80
Introduction into Quadrupole Mass Spectrometry
Mass-Filter, peak form and contribution to neigbouring mass

## Slide 81
Detector
and in general:
Ion-Source
Mass-Filter
Electron Multiplier
Faraday

## Slide 82
Introduction into Quadrupole Mass Spectrometry
Faraday Cup for „robust every day operation“, detection of high ion currents, and determination of the basic sensitivity (A/mbar)
SEM for high amplification and detection of trace components
Detectors

## Slide 83
Detector basics
Ion detection / Faraday and SEM
The ions that have passed the filter impinge on the Faraday collector where they release their charge. The very small resulting current is the input to the electrometer preamplifier. If a secondary electron multiplier (SEM) is configured and switched on, the Deflection Unit directs the ions towards the SEM. The negative high voltage on the input of the SEM attracts and accelerates the ions. On impact they release electrons and these in turn release additional electrons which leads to an amplification of several orders of magnitude.
Electrometer preamplifier This instrument converts very small currents (10 -15 ...10 -5 A) from the Faraday or SEM to a voltage which is subsequently processed in the quadrupole controller.
Introduction into Quadrupole Mass Spectrometry

## Slide 84: Introduction into Quadrupole Mass Spectrometry
Ions leaving the quadrupole mass spectrometer fly into a Faraday-Cup and give off their charge
Charge is measured as a current by an electrometer amplifier
Measurable currents are from 1E-15 to 1E-8 A

## Slide 85: Introduction into Quadrupole Mass Spectrometry
Secondary Electron Multipliers
SEM 217, 17 discrete dynodes
ChanneltronTM, continuous dynode (incl. Faraday Detector)

## Slide 86
Introduction into Quadrupole Mass Spectrometry
Secondary Electron Multiplier
SEV 217, 17 discrete dynodes

## Slide 87
Introduction into Quadrupole Mass Spectrometry
Particles (ions, neutrals, electrons, photons) hitting a surface with high energy release several „secondary electrons“
Use of several dynodes allows for an amplification of up to 108
Low ion currents can be detected easily
Measuring range 1E-15 to 1E-5A
„Counting“ of individual electron bursts allows for the detection of single ions (1E-19A)
Positive Ion
Electrons
-1 to - 3.5 kV
Secondary Electron Multiplier, discrete

## Slide 88: Introduction into Quadrupole Mass Spectrometry
Multiplier consists of glass tube with semi-conducting layer
High voltage ( -1 to - 3kV) is applied to the front end of the C-SEM
Electrons fly within the tube and release secondary electrons when hitting a surface
Bended tubes produce more hits of the electrons onto the surface than straight tubes
Ion Electrons
Secondary Electron Multiplier, continuous

## Slide 89: Introduction into Quadrupole Mass Spectrometry
If no voltage is applied to the C-SEM, the ions fly straight onto the Faraday-Cup  analysis of high ion currents
A high voltage applied to the C-SEM deflects the ions into its entrance  analysis of low ion currents
One electrometer can measure currents both from the Faraday-Cup and from the C-SEM
Combined Faraday-Cup, C-SEM detector

## Slide 90: MS basics
General To ensure filament longevity, ion mobility and detector operation the analyzer has to be kept under high vacuum conditions.
  At high pressure filaments burn quickly
  At high pressure Secondary Electron Multipliers degrade quickly

## Slide 91: MS basics
mean free path (log. scale)
Pressure at the place of the Mass-Spectrometer
10
-5
10
-4
10
-3
10
-2
5 meter
50 cm
5 cm
5 mm
Ion-Molecule-Reactions
mean free path has to be
much larger than the
dimensions of the mass-
spectrometer
Mass-Filter
Why is it necessary to work under vacuum?

## Slide 92
Introduction into Quadrupole Mass Spectrometry
Scan Analog
Scan Bargraph
Multiple Ion Determination (MID)
Data acquisition modes

## Slide 93
Data acquisition modes - Scan Analog
Here a range of masses are scanned and the data appears as
intensity vs. mass. This mode is mainly used when all available information from the mass spectrometer is needed, for diagnostics,
and checking mass spec performance, peak shape etc.

## Slide 94
Introduction into Quadrupole Mass Spectrometry
0
20
40
60
80
100
1
10
100
1000
10000
100000
1000000
Residual Gas Spectrum
pressure in the process chamber : 5*10 E-8 mbar
(turbo molecular pump)
H2
,water vapour,CO and CO2
are the dominat gases
"hydrocarbons" are clearly detected as traces
Intensity [Arb.Units]
mass number
Vacuum-System

## Slide 95
Introduction into Quadrupole Mass Spectrometry
Classical residual gas spectrum in the UHV:
Lowest level of interference between process gas and residual gas
UHV Vacuum- System

## Slide 96
Data acquisition modes - Scan Bargraph
This mode is similar to Analog but it only reports data when a mass
peak is found or in steps of 1m/e - this reduces the data for simplicity.
This mode is mainly used for scanning for unknown components.

## Slide 97
Data acquisition modes - MID
MID - here the masses to be monitored are known and the
system will only report back these points. It is the fastest mode and
very useful for trending data over time
MID:
Multiple Ion Determination

## Slide 98
Industrial, modular systems for outgassing measurements
D & M Vacuumsystemen B.V.
https://www.dm-vacuumsystems.com/

## Slide 99
Industrial, modular systems for outgassing measurements
Protocol of a calibration run

## Slide 100
Protocol of a sample run over several hours with background data for comparison
Industrial, modular systems for outgassing measurements

## Slide 101
Industrial, modular systems for outgassing measurements
Pfeiffer Vacuum
https://www.pfeiffer-vacuum.com/global/de/products/systems/residual-gas-analysis-systems

## Slide 102
Introduction into Quadrupole Mass Spectrometry
Filament
e
-
-
+
Ionisation
e
-
e
-
e
-
e
-
Ions
-
Filament
e
-
Ionisation
e
-
Ions
+90V +150V +145V
e
-
Electrons can impact onto the chamber
walls --> electron impact desorption
Emissioncurrent
about 1 mA
Ioncurrent < 1e-5 mA
Electrons only can impact onto the anode,
if there is electron impact desorption, then only from
well defined surfaces
commonly used technique advantage of a biased system
Voltages
method for background reduction

## Slide 103
Introduction into Quadrupole Mass Spectrometry
Mass-Filter
Filament
e
-
Ionisation
e
-
Ions
e
-
+90V +150V +140V
150eV
Energy
Ions
(10eV Energy)
Chassis at ground potential
electrically isolated RF-generator
Ions generated in the ionisation room are accelerated by a high electrical field, they move with an energy of 150 eV (in this example)
Thus the ions can cross fringing fields very easily, excellent injection conditions into the mass filter are achieved, a high transmission and excellent mass seperation is achieved.
Inside the filter the ions travel with a well defined energy, given by the so called „field axis voltage“ (here 10 eV)
Voltages
“field-axis“-technique

## Slide 104
Spectra interpretation
Exercises

## Slide 105
Spectra interpretation - exercises
Basic steps:
Look for major peaks
Check the library/key fragment ion listing
If you suspect a certain substance, look for additional typical peaks
Remember:
Molecular peaks, represent the molecular weight of the substance
Fragment peaks, represent functional groups of a sample
Atomic peaks, represent the atomic weight of the element
Isotope peaks, occur in the same abundance as they occur in nature (if not intentionally changed)

## Slide 106
Air residual gas in system

## Slide 108
Helium used for leak detection
4

## Slide 110
Acetone in air e.g. cleaning residue
43
58

## Slide 112
Ethanol in air e.g. cleaning residue
45
31

## Slide 114
Carbon tetrafluoride (CF4) in air Etch gas for Semi processes
69
50

## Slide 116
Hydrogen Chloride (HCl) in air e.g. Product from PVC
70
36
38
72
74

## Slide 118
Sulfur hexafluoride (SF6) in air e.g. Tracer gas, isolation gas
127
89
108

## Slide 120
Solvent in air (Thinner for Rubber cement)

## Slide 122
Solvent in air Filling of Zippo-Lighter

## Slide 124
Solvent in air, Solvent of Tippex (whiteout) Linear Alkanes are characterized by CH2 steps
55,56,57 +14
41,42,43 +14
69,70,71 +14
83,84,85....