## Slide 1: X-ray spectrometers with curved crystal optics Konstantin Klementiev MAX IV Laboratory

## Slide 2
Software library xrt
xrt can calculate reflectivity and transmittivity of elastically deformed bent crystals (added in 2023)

## Slide 3
pyTTE — a package to calculate XRD of deformed crystals
https://github.com/aripekka/pyTTE
Honkanen & Huotari, General method to calculate the elastic deformation and X-ray diffraction properties of bent crystal wafers. IUCrJ (2021) 8, 102-115, https://doi.org/10.1107/S2052252520014165
pyTTE is (was) the state-of-the-art bent crystal calculator
Notes: Components of the strain tensor in terms of transverse stretching stress. This is in addition to the pure bending strain.

## Slide 4
Bent crystal reflectivity calculator of xrt
xraylib → xrt.materials (effortless installation)
scipy.integrate.BDF → RKDP 45 (SIMD parallelization)
multiprocessing → OpenCL (performance gain up to ×25000)
Conversion
from pyTTE to xrt
xrt accelerates pyTTE by up to 4 orders of magnitude
Notes: backward-differentiation formulas, Dormand–Prince, Single Instruction Multiple Data, direct methods vs adaptive

## Slide 5
Bent crystal reflectivity calculator of xrt: xrtBentXtal
xrtBentXtal is a GUI for comparative bent crystal calculations
“magic” angles Si111: 64.900°
Si220: 62.865°
Si200: 58.214°

## Slide 6
See more about spectrometer studies by ray tracing here: xrt.readthedocs.io/gallery2.html
von Hámos
Johansson geometry provides (almost) exact point-to-point imaging and is the best for efficiency and energy resolution. Balder’s spectrometer SCANIA-2D is of this type.
Comparison of XES spectrometer types

## Slide 7
Why the efficiency is different?
Johansson is best for fixed emission energy applications, von Hamos is convenient for (wide band) XES
von Hámos
Johansson
Johann
Only a very narrow crystal band is effective, but the crystal is usable in dispersive detection
The complete crystal length is effective
Only the central part is effective but easy to make
Especially attractive for fixed excitation energy XES
Especially efficient for fixed emission energy detection (HERFD XANES)

## Slide 8
XES spectrometer at Balder beamline
Even a mild 1D bending force affects energy resolution of crystal analysers at high index reflexes
Without crystal deformation:
With crystal deformation by means TT eqs:

## Slide 9
Main design ideas
Backscattering relative to the sample
Detector under the table
Large Bragg angle range (85° – 35°)
Dectris EigerX 1M Hybrid-pixel detector 3 kHz frame rate
2 independent crystal branches for “2-color” detection
Ground, polished and thinned (300 µm) test crystal
Notice: it is not bent, it is ground!
Made by Iulian Preda
Designed and built in-house. Named as SCANIA-2D — Segmented Crystal ANalyzer with Image Acquisition in 2D.
XES spectrometer at Balder beamline

## Slide 10: Balder experimental hutch
beam conditioning
XES spectrometer SCANIA-2D
2 fluorescence detectors
7-element SDD, 2-20 keV – Canberra
7-element Ge detector, 5-40 keV – Canberra
LHe cryostat
gas system
(not shown)
2D XRD detector
In-situ sample cells
  1000oC furnace and controller
  MAX IV gas-flow cell for catalysis research
  Microfluidic liquid flow cell
up to 50 bar with fast mixing, vapor and MS
Normal sample holder
SCANIA-2D takes a significant part of the hutch. Compatible with all our sample environments and detection techniques except XRD that in practice requires a small slitted-down beam.

## Slide 11
Assembly technology and optical quality measurements were established in Nov 2020. Main limitations seem to come from dust particles between crystal and reference cylinder.
Crystal assembly technology

## Slide 12
SCANIA-2D in metal
SCANIA-2D has been built with one branch so far. The chamber is a plastic helium box.

## Slide 13
On-the-fly data reduction pipeline for SCANIA-2D
Note: The animation timeline is not real. The actual scans have a few hundred scan points with a typical total scan duration of ~10 min.
Scanning axis
Tangential detector axis
Scanning axis or projected tangential-to-scanning
XES
Tangential detector axis
Sagittal detector axis
Scanning axis
Press Enter to see animations
Original 3D data
Reduced 2D data
XES or HERFD XAS
Sheer correction, sagittal integration
θ-2θ band integration
The reduction pipeline was implemented in 2024, running automatically and asynchronously

## Slide 14
XES energy calibration by elastic lines
θ-θ band (rotates together with the crystal) due to the Borrmann effect and the scattering from the back of the crystal
With a 2D detector we see a zoo of effects: Borrmann, powder diffraction from the back surface, multiple diffraction. The θ-θ band dictates the alignment strategy: all crystal are used together, not individually.

## Slide 15
Energy resolution as given by elastic lines
At the magic Bragg angle (here, 64.9°), energy resolution is much better than the band width of Si111 DCM

## Slide 16
Conclusions
xrt can calculate elastically bend crystals and do ray tracing with these calculations applied to each ray.
There exists a magic angle at which the reflectivity curve shrinks.
Johansson spectrometer offers both high efficiency and high resolution.
The SCANIA-2D spectrometer at Balder beamline has begun its user operation.
Thank you for your attention!

## Slide 17
The calculated DCM bands match the experimental ones
Johansson analyzer with elastically deformed crystal
flat source | line source | 7 lines
perfect crystal
deformed crystal
(Takagi- Taupin)

## Slide 18
ParSeq-XES data pipeline

## Slide 19
The Co Ka1 HERFD (High Energy Resolution Fluorescence Detected) XANES spectra were collected together with expert commissioning users (A. Jentys group from TU Munich) in July 2022. A full spectrometer branch (3 crystals) was used for the first time.
Here, the scanning range had 650 points (rather excessive) at 0.5 sec per point. Note the noise level in the inset plot that would remain much smaller than signal with a sample dilution down to 1wt%.
HERFD XANES is one of the main applications mode of SCANIA-2D
HERFD XANES

## Slide 20
The calculated DCM bands match the experimental ones
DCM bands