## Slide 1
Davide Sisana, PhD student
on behalf of the BEaTriX (INAF-OAB) and in collaboration with IMEM-CNR and ESRF
High Precision X-Ray Measurements 2025, 16-20 June, LNF INFN
BEaTriX and beyond: an innovative compact X-ray calibration facility for astronomical optics enabling wide, collimated, and monochromatic beams

## Slide 2
Outline
  Why BEaTriX
  Opto-mechanical design
  4.51 keV BEaTriX beamline
  1.49 keV BEaTriX beamline
  BEaTriX+
  Application beyond astronomy
June 16, 2025
2
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 3
June 16, 2025
3
Context
AIT facilities for NewAthena X-ray space observatory
The largest X-ray mirror for astronomy ever built - 2.4 m diameter primary mirror
0.3 – 12 keV energy bandwidth
Requirement
X-ray calibration facility to test (PSF and Effective Area) about 500 SPO MMs of NewAthena at a rate compatible with MM production (~ 3 MM/day)
How to re-create astronomical X-ray source at an infinite distance?
Silicon Pore Optics technology
Why BEaTriX
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 4
June 16, 2025
4
State of the art for conventional X-ray facilities
Source at large distance
(PANTER, MSFC, …)
Synchrotron source
(ESRF, ELETTRA, …)
Large volumes in vacuum
Beam divergence
Pencil-beam
PSF reconstruction
Critical aspects
BEaTriX
Innovative optical design
Compact system
Collimated beam
Wide beam
Monochromatic beam
9 x 18 m2
divergence < 2 arcsec
17 x 6 cm2
10-5 eV
Beam Expander Testing X-ray facility
Why BEaTriX
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB
100 meters

## Slide 5
Where BEaTriX
June 16, 2025
5
INAF-Brera Astronomical Observatory (Merate)
BEaTriX
Compact system
Collimated beam
Wide beam
Monochromatic beam
9 x 18 m2
divergence < 2 arcsec
17 x 6 cm2
10-5 eV
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 6
Optical layout at 4.51 keV
June 16, 2025
6
Asymmetrically cut
Silicon crystal (220)
High dispersive element, a tight monochromation is needed
Low photoelectric absorption
Symmetrically cut Silicon crystal (220)
An X-ray source with a high brilliance is needed
Tight monochromation
X-ray source with Titanium anode
+
Paraboloidal mirror
(fused Quartz HOQ 310)
Coating: Cr + Pt
Fluorescence at 4.51 keV
High brilliance: > 5 x 1011 ph/s/sr
Spot size: 35 µm FWHM
Mirror surface quality: 2 arcsec
EXPANSION STAGE
(by diffraction)
MONOCHROMATION STAGE
(by diffraction)
X-RAY SOURCE and COLLIMATING MIRROR
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 7
Optical layout at 4.51 keV
June 16, 2025
7
High flux
High collimation
4.51 keV configuration
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 8
Optical layout at 1.49 keV
June 16, 2025
8
Asymmetrically cut ADP crystal (101)
High dispersive element, a tight monochromation is needed
fused quartz HOQ 310
EXPANSION STAGE
(by diffraction)
MONOCHROMATION STAGE
(by diffraction)
X-RAY SOURCE and COLLIMATING MIRROR
Low photoelectric absorption
Large d-spacing
Asymmetrically cut Quartz crystal (100)
An X-ray source with a high brilliance is needed
Rocking curve narrower than ADP one
Tight monochromation
X-ray source with Aluminum anode
+
Paraboloidal mirror
(fused Quartz HOQ 310)
Fluorescence at 1.49 keV
High brilliance: > 5 x 1011 ph/s/sr
Spot size: 35 µm FWHM
Mirror surface quality: 2 arcsec
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 9
Optical layout
June 16, 2025
9
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 10
Mechanical design
4.51 keV
X-ray source
Optical chamber
Experimental chamber
CCD detector
Long arm
(< 12 m)
Short arms
(< 5 m)
June 16, 2025
10
1.49 keV
X-ray source
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 11
From design to reality
4.51 keV
X-ray source
Experimental chamber with hexapod and thermal box
CCD and cold finger
Long arm
Short arm
June 16, 2025
11
Paraboloidal mirror
Coating: Cr + Pt
Monochromators CCC
  Si (220) symmetrically-cut
Beam expander
  Si (220) asymmetrically-cut
Chiller
  Focal spot = 35 µm
  Flux = 10 11 ph/s/sr on Ti-Kα
Optical chamber
  Size: 27.6 mm
  Pixel: 13.5 µm
[Salmaso et al., Proc. SPIE 12181, 2022]
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 12
4.51 keV beamline – Performance
170 mm
60 mm
June 16, 2025
12
After dithering of the beam expander (low-frequency)
SDD detector for the optical element's alignment
170 mm
60 mm
 | High flux | High collimation
2nd CCC rotation | 0 arcsec | 10 arcsec
Flux | 60 ph/s/cm2 | 10 ph/s/cm2
HEWvertical | < 2 arcsec | 
HEWhorizontal | 4 arcsec | 2.5 arcsec
Uniformity | Up to 95% | 
[ Basso et al, ‘’First light of BEaTriX, the new testing facility for the modular X-ray optics of the ATHENA mission’’, A&A 664, A173, 2022]
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 13
4.51 keV beamline – Experiments
June 16, 2025
13
Astronomical measurements: SPO MMs for NewAthena
Point Spread Function
Effective Area
Focal length
Thermal cycles
Straylight analysis
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 14
4.51 keV beamline – Experiments
June 16, 2025
14
Deconvolution of the image with the BEaTriX X-ray beam
High collimation conf. for PSF
High flux conf. for Effective Area
To avoid double measurements:
high flux conf. + deconvolution
Requirement: 3 MM/day
Deconvolution
[Chauhan et al., Proc. SPIE 2025, in preparation]
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 15
1.49 keV beamline – Crystals
June 16, 2025
15
Under test at PANTER facility
Tested at PANTER facility, ready for the installation
Characterized at IMEM-CNR (Parma)
Characterized at IMEM-CNR and ESRF synchrotron
80 x 80 x 20 mm3
80 x 30 x 15 mm3
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 16
1.49 keV beamline – Crystals
June 16, 2025
16
Characterization of the crystals
Requirement on crystal planes RoC: 22 km
ADP is more challenging with respect to Si
Crystal planes with RoC of 22 km are difficult to be measured
RoCs out of specifications
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 17
1.49 keV beamline – Simulations
June 16, 2025
17
How to improve the performance and to address the challenging aspects?
Simulations to evaluate the impact of the crystal planes RoC out of spec, both with Reference IDL code and OASYS-SHADOW4 software
[in collaboration with Sanchez del Rio, ESRF]
[Spiga et al., Proc. SPIE ICSO 12777, 2023]
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 18
1.49 keV beamline – Simulations
June 16, 2025
18
Parabola | Quartz#1 | Quartz#2 | ADP | OASYS-SHADOW4 |  | Reference IDL code | 
 |  |  |  | HEW_hor | HEW_vert | HEW_hor | HEW_vert
Ideal | Ideal | Ideal | Ideal | 2.81 arcsec | 0.87 arcsec | 3.21 arcsec | 0.87 arcsec
Ideal | Ideal | Ideal | RoC_h: 3.13 km RoC_v: 6.43 km | 5.98 arcsec | 1.16 arcsec | 7.73 arcsec | 1.70 arcsec
How to improve the performance and to address the challenging aspects?
The main contribution on the HEW values is due to the crystal planes curvature
Bending system for ADP crystals
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 19
1.49 keV beamline – Bending system
June 16, 2025
19
Design of Concept: bending system for ADP crystal
Mechanical design
Finite Element Analysis
Surface reconstruction
Pseudo-inverse interaction matrix method
Matlab code
OASYS-SHADOW4 simulations
HEW values
IMEM-CNR and ESRF measurements
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 20
1.49 keV beamline – Bending system
June 16, 2025
20
Proof of Concept: bending system for ADP crystal
The bending system design is frozen; the pieces are under production at INAF-OAB workshop
The system will be calibrated in optical wavelength at ZYGO interferometer at INAF-OAB promises
Final test: X-ray measurement at ESRF synchrotron
Installation in BEaTriX facility
TRL 4
Validation in laboratory
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 21
1.49 keV beamline – Bending system
June 16, 2025
21
Bending system for ADP crystal – Mechanical design
Piezo motor
System to compensate the horizontal RoC
System to compensate the vertical RoC
ADP crystal
155 mm
187 mm
164 mm
[Sisana et al., Proc. SPIE 2025, in preparation]
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 22
12 m
10 m
8 m
Detector tower moving
BEaTriX +
BEaTriX+
funded by INAF within “Potenziamento Laboratori Spaziali”
June 16, 2025
22
State of the art
3 possible focal length configurations
Flexible focal length from 1.5 m to 13.7 m
The long arm is replaced with shorter sections of 300 to 600 mm to enable the positioning of the detector, hosted in a stand-alone module, at any position in the range 1.5 – 13.7 m
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 23
Application beyond astronomy
June 16, 2025
23
BEaTriX: X-ray lab-based imaging system enabling wide, collimated, and monochromatic beams
Experimental chamber size:
  750 mm diameter x 718 mm height
Thermal control during experiment:
  -15 °C to 55 °C
Flexibility on the CCD detector position:
  from 1.5 m to 12 m
Beam properties | 4.51 keV | 1.49 keV
Size | > 150 mm x 50 mm | 
Collimation | < 2 arcsec | < 3.5 arcsec *
Monochromation | 35 meV | 23 meV
Polarization in the vertical plane | Linear | Partially linear (95%)
* With the crystal planes RoCs in spec
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB

## Slide 24
Application beyond astronomy
June 16, 2025
24
BEaTriX: X-ray lab-based imaging system enabling wide, collimated, and monochromatic beams
What suggestions do you have for making BEaTriX a facility also for your community?
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB
Biological sample
(insects, fabrics, …)
Astronomy
Biomedical purpose
Add/built a high energy beamline using the same optical design of BEaTriX
Mammography (~ 25 keV)
…

## Slide 25
High Precision X-Ray Measurements 2025, 16-20 June – BEaTriX, Sisana Davide – PhD student Politecnico di Milano/INAF-OAB
Thank you for your attention
About BEaTriX
Contacts:
davide.sisana@inaf.it (speaker)
bianca.salmaso@inaf.it (PM of BEaTriX)