## Slide 1: A medical and multi-disciplinary in-air transport beam line for the 70 MeV protons from the INFN-LNL cyclotron
Dr Roberto Catalano and Dr Giada Petringa
Istituto Nazionale di Fisica Nucleare – Laboratori Nazionali del Sud (Italy)
G.A.P. Cirrone, G. Cuttone,
Istituto Nazionale di Fisica Nucleare - Laboratori Nazionali del Sud (Italy)
G. Russo, F. Cammarata
CNR – Istituto di Bioimmagini e Sistemi Biologici Complessi

## Slide 2
2002
The INFN-LNS activity on beamline realization for clinical proton beams
2018
2017
2025
I-LUCE facility
First proton therapy treatment @CATANA
2019
The LNS activity in the framework of beam facility
First improvement of Ion irradiation facility @LNS (light ions modulation)
Second improvement of Ion irradiation facility @LNS (high intensity)
Collaboration with Best Cyclotron
ELIMED installation
Double ring implementation @TIFPA
2020
How can LNS contribute to the future research program with the high power cyclotron of SPES-LNL?

## Slide 3: Activity @LNS with conventional proton beams

## Slide 4
CATANA @LNS-INFN
Centro di AdroTerapia ed Applicazioni Nucelari Avanzate
Cuttone et al. EPJ 126 (2011)
Mainly dedicated to proton irradiation (eye melanoma treatments)
First Italian proton therapy facility
First Patient treated on March 2002 (about 400 treated)
Passive proton beam line
Modulated 62 MeV protons
Dosimetry and radiobiology experiments
Collimated beams up to 35 mm (lateral homogeneity ≈ 3%)

## Slide 5
CATANA @LNS-INFN
A passive proton beamline
Scattering System
Modulation System
Collimator
1
2
3
4

## Slide 6
CATANA @LNS-INFN
Scattering system
Double Scattering system with one occluding central stopper
20um Ta
15um Ta
brass
Gaussian beam → Clinical beam
SEM

## Slide 7
CATANA @LNS-INFN
Modulation system
The dose in the SOBP can be expressed as the superimposition of many Bragg curves expressed as:
where Bi(x,y,z) is a Bragg curve that can be determined experimentally and analytically; wi is the optimized weight determining the modulator shape, and accounting for the contribution of each peak i to the final SOBP. The weights can be obtained by the minimization of a cost function, Q, defined as:
Range shifter + modulator (homogeneity along Z direction)

## Slide 8
CATANA @LNS-INFN
Online Monitor system
collection efficiency
beam flux
mean energy needed to produce an ion pair for a charged particle of energy E
mean free path in the gas
gas density
mass electronic stopping power
In-trasmission ionization chambers
Multistrip ionization chambers detector to monitor online the beam geometry (position and uniformity)

## Slide 9
Collaboration with the
Best-Cyclotron company
to realize a beamline to treat ocular melanoma with commercial cyclotron (both low and high intensity beams)
Feasibility study of a new beamline
A new monitor chamber for both low and high intensity beams

## Slide 10
Irradiation point
In-transmission
ionization chamber
Modulation system for ions (Ridge and Ripple fliters)
Secondary Emission Monitoring (SEM) for scattering and beam fluence
Beamline dedicated to dosimetric tests and radiobiological experiments with ion beams of 62 MeV/n:
- 4He
- 12C
- 16O
Ion irradiation facility @LNS
0 degree beamline

## Slide 11
Ion irradiation facility @LNS
Modulation system

## Slide 12
Faraday cup
Irradiation setup for high intensity proton beams (absolute dosimetry)
Ion irradiation facility @LNS
The cylindrical symmetry of the electric field provided by the external electrode is broken due to the presence of the internal one.
The resulting effect is a strongly asymmetric electric field, characterized by a significant transversal component able to maximize the deflection of the secondary electrons genera-ted by both the entrance window and the cup.

## Slide 13
Ion irradiation facility @LNS
Irradiation setup for high intensity proton beams (relative dosimetry)
15um Tantalum foil
Insulator PMMA frame
Charge integration for normalisation purposes
Scattering foil for beam diffusion
f2/f1 = efficiency ratio
f1 = Efficiency of the chamber 1
Determined by the measured ratio of the currents
Efficiency of chamber 1 can be determined
Two adjacent IC, gaps of 5 mm and 10 mm, independently biased (maximum applied voltage ± 1000 V and ± 2000 V, respectively)
Anode: thin layers of 5 µm of copper and 2 µm of nickel, deposited on a 25 µm layer of kapton
Cathode:12 µm-thick layer of aluminized mylar
Secondary Emission Monitoring
Dual-gap ionization chamber
Specifically designed for high
intensity beams

## Slide 14
Instantaneous current [nA] | 5 - 50
Pulse duration [ms] | 10 - 200
Instantaneous dose rate [Gy/s] | 20 - 230
Beam spot shape and size [mm2] | Squared, 9.8 x 9.8
Nominal energy [MeV] | 60 ± 0.5
SEM
DGIC
FC
Dosimetric test with flash beams

## Slide 15
Biological sample irradiation @LNS
Analysis pre and post irradiation
In-vitro positioning system
Types of equipment into the Bio-Lab:
Laboratory Hood
Inverted microscopy
Centrifuge
Incubator
-80°C for storage of biological samples
Dewar for long term storage of different cellular batch
In-vivo positioning system
Small animals storage

## Slide 16: Collaboration with TIFPA’s group and ELI-BEAMLINES

## Slide 17
Ridge
filter
pre-collimator
collimator
Monitor
Chamber
Scattering
System
F. Tommasino et al.,A new facility for proton radiobiology at the Trento Proton Therapy centre: design and implementation”,Physica Medica, 10.1016/j.ejmp.2019.02.001
Irradiation facility @TIFPA

## Slide 18
Dosimetric system for high intensity proton beams
ELIMAIA Beamline | L3 HAPLS | L4 ATON
Peak power | ³1 PW | 10 PW
Energy in pulse | ³30 J | ³1.5 kJ
Pulse duration | £30 fs | £150 fs
Rep rate | 10 Hz | 1 per min
Supplier | LLNL | National Energetics
ELI-Beamlines | Compressor, short pulse diagnostics, controls & timing systems | Compressor design, OPCPA design, short pulse diagnostics, timing system
- A set of a 5 PMQ allows to focus the particle beam and provide a rough energy selection
The ESS consists of four electromagnetic dipoles and a slit; it was designed to select protons with energies up to 300 MeV and heavier ions with energies up to 60 MeV per nucleon
Downstream the ESS, the ion beam is transported into the final beam shaping section, consisting of two resistive quadrupoles and two steerers for a fine focusing and alignment of the ion beam

## Slide 19
Monte Carlo Simulation for beamline design
Monte Carlo simulations for beamline design
A dedicated Geant4 application was develo-ped to optimize and realize the beamlines:
CATANA (LNS)
Ion irradiation facility (LNS)
Best Cyclotron beam line
ELIMED (ELI)
I-LUCE (LNS)

## Slide 20
Monte Carlo Simulations for dosimetric and microdosimetric purposes: LET and RBE
Open source G4 Monte Carlo codes
Hadrontherapy: Geant4-based application specifically developed to address typical needs related to the proton and ion therapy
Radiobiology: Geant4-based application realized for dosimetric and radiobiological applications of proton and ion beams
Microdosimetry (INFN, Section of Catania)
MicroPlus probe at the isocenter position of the CATANA protontherapy facility
Mini-TEPC at the isocenter position of the CATANA protontherapy facility
Silicon Telescope at the isocenter position of the CATANA protontherapy facility
20
LET-dose
LET-track

## Slide 21
Monte Carlo Simulations for dosimetric and microdosimetric purposes: LET and RBE
Experimental survival data of the U87 cells irradiated in six different positions along the 62 MeV proton pristine Bragg peak

## Slide 22: Current activity in the framework of I-LUCE

## Slide 23
Cyclotron
Tandem
Exp beamlines
Offices
Laser-area

## Slide 24
Two laser systems (L1 and L2)
Three laser outputs
Three experimental stations (ES1, ES2 and ES3)
I-LUCE (INFN-Laser Induced radiation production)
1J+/23fs/45TW
7J+/23fs/320TW
9mJ/100fs/1 kHz

## Slide 25
Transport beamline and diagnostic for ion beams
ES2
Up to 900 targets
Max rep rate: 3 Hz
1 um resolution in Z
EMP compatible
Target System
Transport up to 30 MeV with an energy resolution
of 5 %
Permanent Magnet Quadrupoles
Energy Selection System
Transport System
SEM
ICT
Faraday cup
Scintillator
SiC & Diamond
Diagnostic System
Dosimetric System
Dosimetric System
Transport System
Max energy | 50 MeV
Particle per pulse | 1011 MeV-1 Sr-1
Beam divergence | ± 20°

## Slide 26
How can we contribute?
By providing our expertise in the design and implementation of proton transport lines in air.
By performing Monte Carlo simulations to optimize proton transport in air.
By supplying monitoring and dosimetric systems for both low- and high-intensity proton beams, in continuous and pulsed modes.
By offering the readout electronics developed for all dosimetric and beam monitoring systems.
By supporting the development of a dedicated irradiation system for biological samples.

## Slide 27: Thanks to everyone
F. Abubaker, C. Altana, A. Amato, S. Arjmand, D. Bandieramonte, D. Bonanno, F. Cammarata, R. Catalano, G.A.P. Cirrone,
G. Cuttone, F. Farokhi, S. Fattori, M. Guarrera, A. Hassan, A. Kurmanova, C. Manna, N. Macaluso, A. Miraglia, M. Musumeci,
D. Oliva, A.D. Pappalardo, G. Petringa, L. Raffaele, D. Rizzo, G. Russo, A. Sciuto, J. Suarez, S. Tudisco, M. Tringale