## Slide 1
A. Bondar, A. Buzulutskov , V. Parkhomchuk,
A. Petroghitsky, T. Shakirova, A. Sokolov
Budker Institute of Nuclear Physics,
Novosibirsk State University
February 25, 2025
2ND International Workshop on
Detection Systems and Techniques
For Fundamental and Applied Physics - DeSyt-2025
Low-pressure TPC with GEM readout for experiments with low energy ion beams

## Slide 2
Outline
Accelerator mass spectrometry
New concept of ion identification. Proof of Concept.
Experimental setup
Measurements of energy spectra and track ranges
Recent results with AMS

## Slide 3
Isotopes used for dating
Analyzed isotopes | Half life | Stable isotopes | Stable isobars
10Ве | 1,39 million years | 9Be | 10B
14C | 5730 years | 12,13С | 14N*
26Al | 717 thousand years | 27Al | 26Mg*
36Cl | 301 thousand years | 35,37Cl | 36Ar*, 36S
41Ca | 102 thousand years | 40,42,43,44Ca | 41K
129I | 15,7 million years | 127I | 129Xe*
In the current AMS BINP setup the time-of-flight technique is used for the carbon isotopes separation. But that technique has a serious problem of separating the isobars - different chemical elements having the same atomic mass. The typical example are radioactive isotopes 10Be and 10B.
* - isobars that do not form stable negative ions.

## Slide 4
Accelerator mass spectrometry
Accelerator mass spectrometry (AMS) is an ultra-sensitive method of counting individual atoms. Usually it is the rare radioactive atoms with a long half-life. The archetypal example is 14C which has a half-life of 5730 years and an abundance in living organisms of 10-12 relative to stable 12C isotope.
AMS facilities operate in more than 100 physical laboratories worldwide, two of them are located in Novosibirsk at Geochronology of the Cenozoic Era Center for Collective Use and Novosibirsk State University.

## Slide 5
1. Formation of an ion beam from atoms of the test sample
2. Ion selection at low energy
3. Ion acceleration
4. Recharging of atomic ions and destruction of molecular ions in magnesium target
5. Ion selection in a high voltage terminal
6. Ion acceleration
7. Ion selection at high energy
8. Identification and counting of ions
BINP AMS provides reliable separation of a pure beam of radiocarbon ions from the accompanying ion background.
BINP AMS
Mg stripper

## Slide 6
∆Е-Е detectors
A cross section through the multi-element ionization chamber. The upper panel shows a plan view of the anode electrode.
Detectors used to count the AMS isotope:
silicon detectors;
time-of-flight systems (BINP AMS);
ionization chambers.
Traditional concept of ion identification at AMS
<number>
Ions track lengths at 50 Torr of i-C4H10

## Slide 7
Proof of concept: measuring track ranges
SRIM (The Stopping and Range of Ions in Matter Software) –
is a collection of software packages which calculate many features of the transport of ions in matter.
Ion ranges distributions in the low-pressure TPC for alpha particles with different energies for 200 nm silicon nitride window and 50 torr isobutane, obtained using SRIM simulation.
Track ranges distributions in the low-pressure TPC for 4.025MeV 10B and 10Be ions for 200 nm silicon nitride window and 50 torr isobutane, obtained using SRIM simulation.

## Slide 8
New concept of ion identification: low-pressure TPC
Schematic layout
Ion beam from AMS
or alpha particles
Si3N4 window (200 nm)
Ion track
Ionization
GEM
Anode
Charge-sensitive
Preamplifiers CAEN1523
Cathode
Shaping amplifiers
Naicam 226
Oscilloscope
LeCroy 4054AR
Isobutane
24 mm
Field cage

## Slide 9
Effective gain of GEM
GEM effective gain as a function of the voltage in isobutane at pressures varying from 40 to 70 torr in the low-pressure TPC.
<number>
Electric field inside the holes.
50 µm

## Slide 10
pulse width -> track range
pulse area -> energy
Typical waveform shape of the signal from the alpha particle in low-pressure TPC
TPC Prototype: Principle of operation
Diameter – 178 mm
Length – 300 mm
178
300
226Ra
source

## Slide 11
Test installation
Test installation
Test installation
Test installation

## Slide 12
The measurement of track ranges
The alpha-particle source
226Ra
Pressure = 50 Torr
Gain = 225
Shaping time = 200 ns
2D plot of pulse width versus pulse area and their axis projection spectra for alpha particles from 226Ra (4.8, 5.3, 5.6, 6.0 and 7.7 MeV) source, measured in low-pressure TPC in Isobutane at 50 Torr and GEM gain of 320. The peaks of 5.3 and 5.6 MeV had the same track length’s difference as the 10B and 10Be ions. FWHM of this peaks is 6 mm, which corresponds sigma of 2.5 mm.
11
12 mm

## Slide 13
10B
10Be
Results
Source | Shaping time | Gain | Pressure | Sigma/Range, % | Separation in sigma between two peaks
226Ra | 200 ns | 320 | 50 Torr | 7 | 5.5
Using these results and SRIM code simulations, it is shown that isobaric boron and beryllium ions can be effectively separated at AMS, providing efficient dating at a 10 million years scale. This technique will be applied in the AMS facility in Novosibirsk.
Spectra of ion ranges for 10B and 10Be with energy 4.025 MeV in Isobutane at 50 Torr simulated with SRIM.
11

## Slide 14
11
TPC for BINP AMS
100 mm
70 mm

## Slide 15: TPC for BINP AMS
Test of the low-pressure TPC in laboratory, before installation.

## Slide 16: Installation of TPC on BINP AMS
TPC
TOF detectors

## Slide 17: Signal from 14C and 4He
Signal from 4MeV 14C ion
Signal from 5.1 MeV alpha particle

## Slide 18: Results on carbon beam
Sample with “alive” carbon sample
Sample with “dead” carbon sample
14C

## Slide 19: Conclusion
A new concept of the detector for the accelerator-based mass spectrometer was proposed for identifying ions by their stopping range in gas;
A low-pressure TPC prototype, based on this concept, have been made and successfully tested;
The TPC have been installed on AMS;
First results on 14C beam are promising;
We are preparing to the tests with 10Be samples.

## Slide 20
Thank you for your attention!
19

## Slide 21: Installation of TPC on MICADAS

## Slide 22
Backup slides
19

## Slide 23: Results on carbon beam

## Slide 24
Formation and application 10Be
Application in-situ and
meteoric 10Be:
exposure dating to identified the growths and decays of the Antarctic ice sheet;
understanding ice shelf collapse history;
paleomagnetic excursions history reconstructions using ice cores;
understanding the erosion rates using depth profiles of mid latitudes outcrops;
identifying the timing of formation of the impact crater and so forth.
Time intervals of dating:
14C from 50 years to 40-60 thousand years
10Be from 1 thousand years to 10 million years

## Slide 25
Measurements of energy spectra using semiconductor detector
Alpha particle source– 233U, 239Pu , 238Pu
Energy spectrum of alpha particles from 233U (4.8 MeV), 239Pu (5.2 MeV) and 238Pu (5.5MeV) sources, measured using semiconductor detector
Si Charged Particle Radiation Detectors for Alpha Spectroscopy
19

## Slide 26
The measurement of track ranges
The alpha-particle source
233U, 239Pu, 238Pu
Pressure = 50 Torr
THGEM Gain = 67
Shaping time = 200 ns
19
2D plot of pulse width versus pulse area and their axis projection spectra for alpha particles from 233U (4.8 MeV), 239Pu (5.2 MeV) and 238Pu (5.5 MeV) source, measured in low-pressure TPC in Isobutane at 50 Torr and THGEM gain of 67.
The pulse width and pulse area spectra reflect those of the track range and energy.

## Slide 27
Silicon nitride membrane windows
M. Dobeli et al., Nucl. Instr. and Meth. B, 219-220 (2004) 415-419 doi:10.1016/j.nimb.2004.01.093
Figure of silicon nitride membrane windows
19

## Slide 28
BINP AMS scheme:
1 – accelerator tube, 2 – accelerator body, 3 – electrostatic turn on 180°, 4 – magnesium vapor target , 5 – corrector, 6 – electrostatic lens, 7 – vacuum pump, 8 – ion detector, 9 – high energy dipole magnetic spectrometer, 10 – low energy dipole magnetic spectrometer, 11 – ion source.
BINP AMS characteristics:
magnesium vapor target for molecular destruction localized site of molecular breakdown;
sorting of ions by energy immediately after the destruction of molecular ions; effective screening of fragments of molecules; the energy of the pieces of the molecule is always less than the energy of the molecule;
time-of-flight detector for registering the moment of ions arrival.
BINP AMS
19