## Slide 1
Advancements in Silicon Carbide-Based Detectors for High-Performance Radiation Dosimetry and Beam Monitoring: from simulations to test phase
Giuseppe Ferdinando D’Agata
Università degli Studi di Catania & INFN – Sezione di Catania
DesyT-2025, February 24th, 2025

## Slide 2
The SAMOTHRACE ecosystem aims to realize the vision of a global collaboration environment among major actors in microelectronics, microsystems, materials and micro technologies that have been operating with base in Sicily but with a global perspective.
SAMOTHRACE focuses on the European Commission global challenge “Digital, Industry & Space”, but it also addresses other challenges as “Health”, “Energy & Mobility”, “Agriculture & Environment”.
A collaboration between UniCT, INFN – LNS and INFN – Sez. Catania has recently started with the aim to develop and characterize new innovative SiC detector.

## Slide 3
To develop such a device, the first step is to simulate the detector (GEANT4) keeping in mind:
The detector response
  Expected energies detected
  Necessary thickness for the various purposes
  Solid angle covered
The effects coming from the interaction between the different parts of the detector
  Cross-talk
  Interaction between electric fields
  edge effects

## Slide 4
The chosen geometry in this phase has been a padded one (10 or 100 thickness) suitable for different purposes (medical and nuclear applications)
What we need now is a suitable simulation, able to reproduce the features of the environment, of the «beam» (source for now), and the detector geometry.
The simulation must also be able to easily reproduce the detector resolution.

## Slide 5
1
2
4
3
Given the various purposes and geometries of the padded detector, the simulation must be versatile enough to easily modify
The matrix structure of the detector (how many pads and in which configuration).
The relative position between the beam and the detector.
The characteristics of the beam itself (rate, spatial/energetic distribution)
It is also necessary to implement some feature able to provide information on lost events (almost simple), edge events (harder) and electric field effects (very hard).

## Slide 6
Different α sources (and standard 3-peak)
Different electronics used
AMETEK ORTEK 570 preamplifier - MAESTRO Multichannel Analyzer Emulation Software
Mesytec 16 channel fast preamplifier (MPR-16), incoming signals are processed by means of a
CAEN 571 8 channel digitizer.

## Slide 7
Given the structure and purpose of the detectors, it is crucial to understand how and why impinging particles can also lead to improper, incomplete or totally missing signal reconstruction.
For example, if a particle does not end up in the active zone of the SiC detector, hitting the detector in the inter-pad sections
In such cases either the particle is lost, or the incoming can be detected at the same time. The signal will then be recorded by both pads.
The possibility of missing particles cannot be ruled out a priori

## Slide 8
Edge effect proves to be a non- negligible hindrance to the correct reconstruction of the events. Such events, although in the energy tail of good one, on the other hand show a risetime with a broad distribution. Such an effect is quite spread in the detectors

## Slide 9
The same effect can be detected in the case, with a different, less energetic source (). In this case the effect is even more problematic to treat, given that it all goes in to a «parasite» structure sitting on the low-energy tail of the energy distribution and evenly spread on a long rise time tail.

## Slide 10
The first idea to thoroughly investigate the effect, can be to strongly collimate the source of one of the four pads.
As it is clear, the effect is due to the construction of the detector, given that the «edge» events are taken out if the source is collimated directly on the center of the detector.

## Slide 11
We know that those are not-correctly reconstructed events, and that this effect comes from the sides of the detectors
What we still do not know is how many particles are slipping to the inter-pad regions without being detected
To do so we can compare two simulation, one considering the padded detector with another one that has the same area but is not segmented.
Also, this simulation must be compared with the experimental data

## Slide 12
Experimental data has been coupled with simulations
source (for the )
We know the activity of the source, and can compare the number of event expected with the simulation (solid angle must be considered)
Less then 1% of events are lost in the cracks of the detector
Presence of a small satellite peak at lower channels in the experimental data

## Slide 13
Standard mixed source (for the )
We know the activity of the source, and can compare the number of event expected with the simulation (solid angle must be considered)
Less then 1% of events are lost in the cracks of the detector
Probably due to the higher (and different) energies involved, low energy peaks are broader and less visible

## Slide 14
From the inter-pad zone most the events are not lost, but poorly reconstructed. This could be mainly due to electric field effects in the detector.
Thoroughly simulate this occurrence can be very tricky, due to the difficulties to properly mimic the electric field produced by the different regions of the detectors,
In first approximation we assume that the charge collection in the two adjacent pads generated by an impinging particle is directly proportional to the distance from the hit point and the centre of the pad. In this way we can obtain a first (rough) approximation of the number and energy of the charged particles detected in the inter-pad.

## Slide 15
Simulations and experimental data are reasonably in agreement (here source for reference)
There is a low energy region (in the simulation) that is hard to understand, but could be due to the dead layer thickness or to the glue position considered in the simulation
The low energy region is probably cut due to the dynamic range of the preamplifier

## Slide 16
Given the sturdiness of the detector, we decided to increase the recommended voltage
The result is that the main peak does not move much in both position and charge collected, while the «parasite» one (that now we know comes form the sides of the detector)
What moves, on the other hand, is the unwanted contribution from the sides of the detectors

## Slide 17
As it can be seen, while increasing the bias, the satellite peak moves to lower channels, disappearing at one point (between and ).
This leads to the conclusion that this effect is probably due to field effects in the inter-pad region
Understanding the effects of the electric field in the inter-pad regions is not an easy task.

## Slide 18
CONCLUSIONS:
Inter- pad effects have been studied from simulations
The events from the inter-pad have been simulated, and experimental data are almost in agreement with simulations
Uncertainties rose up from the edge effects
Probably this last point can not be treated via GEANT4 simulations and, different simulation codes must be used (Synopsys,…)