## Slide 1
Paolo Finocchiaro
Istituto Nazionale di Fisica Nucleare
Laboratori Nazionali del Sud, Catania, Italy
Sensor systems for radwaste monitoring and nuclear decommissioning:
experience from the Euratom projects MICADO, PREDIS, CLEANDEM
DeSyT-2025
2nd International Workshop on Detection Systems and Techniques for Fundamental and Applied Physics, Catania (Italy), Feb 24-26, 2025
MICADO, PREDIS, CLEANDEM projects have received funding from the European Union’s Horizon 2020 research and innovation programme under grant agreements No 847641, No 945335, No 945098.

## Slide 2
INFN-Energy committee
supports energy-related initiatives
EU
EU projects
SME

## Slide 3
2008
our R&D activity in the radwaste
monitoring field started in 2008
Ansaldo Nucleare funded
5 fellowships on the subject
2025
contract between INFN and Sogin 2013-2014
2019-2023: Euratom H2020 MICADO
2020-2024: Euratom H2020 PREDIS
2021-2024: Euratom H2020 CLEANDEM
INFN-E DMNR
INFN-E HELNEM
INFN-E RADSIEVE
INFN-E PI3SO

## Slide 4
Type | Name | Takes part in | Period | #Partners | Note
 |  | WP7 (leader) | 2019-2023 concluded | 9 | 
 | SmarTent |  |  |  | Detectors derived from MICADO
 | Portal |  |  |  | Detectors derived from MICADO
 |  | WP7 | 2020-2024 concluded | 47 | Pending new spinoff activities
 |  | WP3, WP4 | 2021-2024 concluded | 11 | 
 |  |  |  |  | Detectors derived from CLEANDEM
EU
EU
EU
triggered participation to H2020 Euratom projects
and the development of new applications
PI3SO

## Slide 6
Euratom call NFRP-2018-10
Measurement and Instrumentation for Cleaning And Decommissioning Operations
EU
This project has received funding from the European Union’s Horizon 2020 innovation action programme under grant agreement No 847641.
This text reflects only the author’s views and the Commission is not liable for any use that may be made of the information contained therein.
Partner | Country
CAEN | IT
ORANO | FR
CEA | FR
XIE | GE
CTU | CZ
ENEA | IT
INFN | IT
SCK•CEN | BE
SOGIN | IT
Radwaste characterization, classification, monitoring and data archiving
Flexibility, Transportability, Digitization, Quantification, Traceability

## Slide 7
alpha particles are immediately stopped inside the material
neutrons, from fission or (a,n) reactions, come out very easily typically at low dose-rates; in some cases there can be many
gamma rays are ubiquitous, penetrating, and come out easily and abundantly
fission fragments are immediately stopped inside the material
beta particles are mostly stopped inside the material and in air
Radwaste radiological monitoring
basic radioactivity coming out is gamma rays and neutrons
they carry information from inside and are worth monitoring
easier external monitoring
what comes out?

## Slide 8
Sensors developed by INFN
SciFi SiLiF
gamma and neutron detectors for long term radwaste monitoring
SiLiF (neutron)
SciFi (gamma)
g/n rejection
≈ 10-9
same threshold
no V bias fine-tuning
Details later in Gaetano Elio Poma's talk

## Slide 9
tested with low level waste (LLW)
Nucleco (Casaccia, Rome) 2023
1-month data taking
10-5 to 10-4 n/s
1.5 to 2 𝛄/s
bg ≈ 0.5 𝛄/s

## Slide 10
tested with low level waste (LLW)
Nucleco (Casaccia, Rome) 2023
neutrons
gammas

## Slide 11
tested with intermediate level waste (ILW with Plutonium)
Orano (La Hague, France) 2022
2-day data taking
10-3 to 10-2 n/s
5 to 45 𝛄/s
bg ≈ 0.5 𝛄/s

## Slide 12
tested with intermediate level waste (ILW with Plutonium)
Orano (La Hague, France) 2022
neutrons
gammas

## Slide 13
SiLiF & SciFi tested with spent fuel (2 days)
Zwilag (Würenlingen, Switzerland) 2018
neutron spectra
(5 casks)
!!!
!!!
cask 181
vitrified waste
non-moderated

## Slide 14
SciFi tested with high intensity X-ray machine
SciFi tested with 1 TBq gamma sources
further developments & collaborations in progress (NDA's)
huge X-ray flux & dose monitoring

## Slide 15
Side development from MICADO gamma counters:
SmarTent
Safeguards inspections:
equip a tent or a room with g counters
SmarTent
further developments & collaborations in progress (ENEA)

## Slide 16
demonstrator with 26 SciFi detectors
signals
directly accepts the SiPM output
handles 64 SiPMs i.e. up to 32 SciFi
provides the bias to SiPMs (low voltage)
2m
3m
2m
SmarTent

## Slide 17
test with 137Cs gamma source - 1.4 MBq
30s collection time
SmarTent

## Slide 18
Portal
Side development from MICADO gamma counters:
gamma portal
natural background: ≈ 1 cps
with 1MBq 137Cs source: ≈ 30 cps
further developments & collaborations in progress (NDA's)

## Slide 20
development and implementation of innovative technologies for
pre-disposal of radwaste streams other than nuclear fuel and HLW
Euratom call NFRP-2019-10
EU
PRE-DISPOSAL MANAGEMENT OF RADIOACTIVE WASTE
radwaste treatment, safety, radionuclides, nuclear science and technology, package monitoring, material science
This project has received funding from the European Union’s Horizon 2020 innovation action programme under grant agreement No 945098.
This text reflects only the author’s views and the Commission is not liable for any use that may be made of the information contained therein.
VTT | FI
NNL | UK
JRC | BE
IMTA | FR
CEA | FR
SCK•CEN | BE
BAM | DE
MAGICS | BE
TUS | BG
CTU | CZ
CVRez | CZ
SÚRO | CZ
UJV | CZ
UniHel | FI
CNRS | FR
ECL | FR
IRSN | FR
ORANO | FR
DMT | DE
FZJ | DE
KIT | DE
NCSRD | EL
ISOT | HU
SORC | HU
TSE | HU
ANN | IT
ENEA | IT
INFN | IT
NUC | IT
POLIMI | IT
SOGIN | IT
UNIPI | IT
FTMC | LT
NRG | NL
IFE | NO
RATEN | RO
A21 | ES
CIEMAT | ES
CSIC | ES
ENRESA | ES
UAM | ES
PSI | CH
GSL | UK
MCM | UK
UoM | UK
USFD | UK
KIPT | UA
SIIEG | UA
metallic materials (WP4)
liquid organic waste (WP5)
solid organic waste (WP6)
innovations in cemented waste handling and pre-disposal storage (WP7)
47 partners
17 countries
25 End User Group members
7 Work Packages

## Slide 21
+
+
Same sensors as
in MICADO...
...with newly developed Electronics &
Wireless Data Acquisition
pre-disposal storage: monitoring of cemented waste (WP7.3)
SciFi gamma sensor
SiLiF neutron sensor
SciFi front-end electronics, DAQ
& wireless data transmission
SiLiF front-end electronics, DAQ
& wireless data transmission

## Slide 22
simulation of expected dose rate
cemented mockup drum
4 SciFi sensors
SciFi sensors
167 MBq 137Cs
final demonstration UJV-REZ (Prague, Czech Republic) 2024
pre-disposal storage: monitoring of ILW cemented waste (WP7.3)

## Slide 23
pre-disposal storage: monitoring of ILW cemented waste (WP7.3)
final demonstration UJV-REZ (Prague, Czech Republic) 2024
2-month data taking

## Slide 24
Database
installed around single drums
hooked on single drums
installed on stillages

## Slide 26
Euratom call NFRP-2019/20-09
EU
Cyber physicaL Equipment for unmAnned Nuclear DEcommissioning Measurements
development of a small Unmanned Ground Vehicle (UGV) equipped with sensors
Partner | Country
CEA | FR
ORANO | FR
CAEN | IT
ENEA | IT
INFN | IT
ARTTIC | FR
SOGIN | IT
ANSALDO | IT
RINA-CSM | IT
AINT | DE
TECNALIA | ES
Dismantling and Decommissioning of nuclear facilities and installations
10 Work Packages
This project has received funding from the European Union’s Horizon 2020 innovation action programme under grant agreement No 945335.
This text reflects only the author’s views and the Commission is not liable for any use that may be made of the information contained therein.

## Slide 27
AiNT (Aachen, Germany) 2024
n-generator
demo mission with 4 sources and neutron generator
Unmanned Ground Vehicle for decommissioning & accident remediation

## Slide 28
Side development from CLEANDEM miniradmeter:
PI3SO - Proximity Imaging System for Sort and Segregate Operations
64 × 1cm3 scintillators
64 × 1cm3 scintillators
PI3SO
test spectra with 137Cs gamma source
Details in Chiara Rita Failla's talk

## Slide 29
Top view
Bottom view
60Co
56 kBq
22Na
15.4 kBq
137Cs
1.4 MBq
PI3SO
Proximity Imaging for Spectroscopic
Sort & Segregate Operations
Details in Chiara Rita Failla's talk

## Slide 30
Summary and prospects
flexible: radwaste, ambient monitoring, scanners, dosimeters, ...
successfully tested from very low to very high dose rates
discussions under way with companies
reliable gamma and neutron detectors have been developed at INFN
solutions in search of problems?
😀
safety & security: airrports, buildings, nuclear material transports,...

## Slide 31
M. Ripani et al., Field Test of the MiniRadMeter Gamma and Neutron Detector for the EU Project CLEANDEM, Sensors 2024, 24, 5905. https://doi.org/10.3390/s24185905
G.E. Poma et al., PI3SO: A Spectroscopic g-Ray Scanner Table for Sort and Segregate Radwaste Analysis, Inventions 2024, 9, 85. https://doi.org/10.3390/inventions9040085
M. Romoli et al., A Wireless Gamma-Ray Monitoring System for Cemented Radwaste Drums, Sensors 2024, 24, 2332. https://doi.org/10.3390/s24072332
L. Cosentino et al., SiLiF Neutron Counters to Monitor Nuclear Materials in the MICADO Project, Sensors 2021, 21, 2630. https://doi.org/10.3390/s21082630
L. Cosentino et al., Gamma-Ray Counters to Monitor Radioactive Waste Packages in the MICADO Project, Instruments 2021, 5, 19. https://doi.org/10.3390/instruments5020019
P. Finocchiaro, From nuclear physics to applications: detectors for beam handling, medical diagnostics and radioactive waste monitoring, Eur. Phys. J. Plus (2020) 135:345. https://doi.org/10.1140/epjp/s13360-020-00352-5
Fabio Longhitano et al., A Scintillator Array Table with Spectroscopic Features, Sensors 2022, 22, 4754. https://doi.org/10.3390/s22134754
A.P. Caricato et al., Thermal neutron conversion by high purity 10B-enriched layers: PLD-growth, thickness-dependence and neutron-detection performances, Eur. Phys. J. Plus (2022) 137:431. https://doi.org/10.1140/epjp/s13360-022-02523-y
A. Massara et al., 6LiF Converters for Neutron Detection: Production Procedures and Detector Tests, Instruments 2023, 7, 1. https://doi.org/10.3390/instruments7010001
F. Rossi et al., The Gamma and Neutron Sensor System for Rapid Dose Rate Mapping in the CLEANDEM Project, Sensors 2023, 23, 4210. https://doi.org/10.3390/s23094210
P. Finocchiaro et al., New technologies for radioactive waste monitoring: Results and perspectives from recent experience, IL NUOVO CIMENTO 46 C (2023) 34. https://doi.org/10.1393/ncc/i2023-23034-9
C. Provenzano et al., Development of a High-Efficiency Device for Thermal Neutron Detection Using a Sandwich of Two High-Purity 10B Enriched Layers, Sensors 2023, 23, 9831. https://doi.org/10.3390/s23249831
M. Morichi et al., The MICADO project and its possible upgrades, IL NUOVO CIMENTO 46 C (2023) 46. https://doi.org/10.1393/ncc/i2023-23046-5
last 4 years:
13 articles in peer reviewed journals
overall scientific production since 2008:
> 100 contributions
≈ 30 students (diploma, master, PhD)

## Slide 32
Thank you for your attention
Gianni, Ricco, Emilio Migneco, Marcello Lattuada, Marco Ripani, Eugenio Nappi, Luigi Cosentino, Alfio Pappalardo, Gianfranco Vecchio, Sergio Scirè, Massimo Barbagallo, Giuseppe Greco, Carlotta Scirè, Carmela Greco, Stefania Grillo, Martina Giuffrida, Carmelo Marchetta, Antonio Massara, Pietro Litrico, Giuseppe Passaro, Fabio Longhitano, Gaetano Elio Poma, Simone Amaducci, Gabriel Frei, Stefano Caruso, Federica Pancotti, Monica Sisti, Nadia Cherubini, Giuseppe Marzo, Luigi Lepore, Giada Gandolfo, Nicolò Ratto, Sergio Lo Meo, Antonietta Rizzo, Matias Simonetto, Quentin Ducasse, Mauro Romoli, Michele Di Giovanni, Paolo Di Meo, Antonio Pandalone, Claudio Principe, Ernst Niederleithinger, Sabah Ben Lagha, Karel Prchal, Jakub Zaruba, Coralie Chapuzet, Thibaut Hemedinger, Lionel Tondut, Christopher Helbert, Eric Breuil, Maugan Michel, Camille Theroine, Anna Paola Caricato, Gianluca Quarta, Marcella Marra, Chiara Provenzano, Chiara Failla
...and thanks to all these people for their contributions and support
...apologies to those I could have (surely!) forgotten