## Slide 1
Radwaste reconditioning:
a new approach through a gamma
scanner/imaging system
PI3SO
Chiara Rita Failla
Dipartimento di Fisica e Astronomia Università di Catania
INFN Laboratori Nazionali del Sud

## Slide 2
Radioactive waste
22 concrete silos
~11.000 m³ of radwaste.
Sellafield
Nuclear waste silo opened (BBC)
‘’At various steps in the predisposal management of radioactive waste, the radioactive waste shall be characterized and classified in accordance with requirements established or approved by the regulatory body ‘’.

## Slide 3
Radioactive waste
>2kW/m3
<2kW/m3
<4 GBq/t for α
<12GBq/t for β, γ
Not harmful
LLW Repository site (NDA)
‘’At various steps in the predisposal management of radioactive waste, the radioactive waste shall be characterized and classified in accordance with requirements established or approved by the regulatory body ‘’.

## Slide 4
Nuclear waste decommissioning
Sort & Segregate
g-Rays
PI3SO
How do we integrate it together?
ILW
LLW
VLLW
how to sort and segregate the waste?

## Slide 5
P = Proximity
I = Imaging
3S = System for
Sort and
Segregate
O = Operations
g counter
Spectroscopic detectors
Proximity imaging system
PI3SO
Funded by INFN-Energy

## Slide 6
• 120 × 80 cm² surface • 108 × 64 cm² scanning area
• top and bottom g-detectors:
- 2×64 linear detector arrays
- 1 cm³ crystals & 36 mm² SiPMs
• anticollision system (LED)
• two motion engines:
- horizontal (both detectors)
- vertical (top)
• 2 mm Carbon fiber tabletop
• software: motion & DAQ
The Mechanical Scanning system

## Slide 7
Focus on the detectors setup

## Slide 8
1 cm³ CsI(Tl) crystal
• light yield 6x104 photons/MeV
• spectral emission around 550 nm • density: 4.51 g/cm3 
good intrinsic detection efficiency
6x6 mm² SiPM
• MPPC Hamamatsu
S14161-6050HS
2 CAEN VX2745 digitizers
• 2x64 channels, 16 bit
125 MS/s (8 ns/sample)
Detectors and Data Acquisition

## Slide 9
The Mechanical Scanning system
Motion software

## Slide 10
The Mechanical Scanning system

## Slide 11
PI3SO
Scanner Tests
• complete table with detectors
• 8 modules at once
• unique bias supply
• 8 flat cables from boards to digitizers
2 digitizers (one serves 4 modules)
2 linear scanner arrays:
Top: (4 modules – 4x16 detectors)
Bottom(4 modules – 4x16 detectors)
• hot spots finding
• Spectroscopic measurement
• 3 scan tests:
Simulation of waste with radioactive sources.
137Cs source in a beamline cross.
Investigation of decommissioned parts of a cyclotron electrostatic deflector.
SiPM power supply

## Slide 12
Chart: Calibration sources spectra
Characterization of the 128 detectors
Sample spectrum
Det. n. 79 – Tab n. 5
Calibration
PI3SO
• peaks finding
• energy calibration
• FWHM Energy Resolution
Source | Peak Energy [keV] | Activity [MBq]
137Cs | 662 | 1.4000
22Na | 511, 1274 | 0.0154
60Co | 1173, 1331 | 0.0560
Chart: Linear - fit result
| -55.9833 | 511.1431 | 662.07190000000003 | 1117.1451 | 1273.7909 | 1332.1043 | 2123.3370999999997
calibration line | 48 | 544 | 676 | 1074 | 1211 | 1262 | 1954
Data Peaks | 544 | 676 | 1074 | 1211 | 1262

## Slide 13
Bakcground scan
NO sources on the table
PI3SO
Background scan

## Slide 14
Top
Bottom
PI3SO
Scan with 3 sources:
60Co, 22Na, 137Cs
60Co
22Na
137Cs

## Slide 15
= background count rate assumed 5 counts/s;
= sample count time assumed 1 s;
background count rate assumed 10 counts/s;
= detector efficiency in counts/decay, i.e., 0.3 times the geometric efficiency that varies with the distance;
C = conversion factor to the other units (1).

## Slide 16
PI3SO
Simulation of waste containing radioactive sources
1.4MBq 137Cs source on top of shielding objects in the “drum”
12,6kBq 22Na source in the “drum” – nothing in between
50,0kBq 60Co source in the “drum” – nothing in between

## Slide 17
PI3SO
Bottom - Log Z
Top - Linear Z

## Slide 18
PI3SO
137Cs source in steel cross pipe.
1.4MBq 137Cs source on top of the table inside a steel cross pipe
12,6kBq 22Na source on top of the table
50,0kBq 60Co source on top of the table

## Slide 19
PI3SO
Top - Linear Z
Top - Log Z

## Slide 20
PI3SO
Bottom - Log Z
Bottom - Linear Z

## Slide 21
Investigation on decommissioned (2012) parts of a cyclotron deflector
PI3SO
?

## Slide 22
PI3SO
Lamina E1 - 15
Electrode E2 - 15
Electrode E2 - 14
hot-spots coordinates

## Slide 23
PI3SO
Foil E1 - 15
spectroscopy example

## Slide 24
PI3SO
(30, 17) germanium detector (lin, 30 min)
(30, 17) PI3SO (lin, 15 min)
foil E1 - 15
¹³³Ba 10.5 y
(302, 356, 383 keV)
60Co 5.27 y
(1173, 1330 keV)
spectroscopy example

## Slide 25
PI3SO
Electrode E2 - 15
spectroscopy example

## Slide 26
PI3SO
(78, 49) PI3SO (lin, 15 min)
(53, 52) PI3SO (lin, 15 min)
[(78, 49); (53, 52)] germanium detector (log, 30 min)
Electrode E2 - 15
44Ti 59.1 y
(67.9, 78.3 kEv)
44Sc, 22Na β+
(511 keV)
44Sc 3.96 h
(1157 keV)
22Na 2.6 y
(1274 keV)
spectroscopy example

## Slide 27
PI3SO
Conclusions
Low cost, compact, robust, commercial components
Fast scan, results in line with expectations
Reliable, repeatable, safe operations
Digital data storage, full scan history
Fast and accurate scan: optimal search for hot–spots
Gamma-ray images
Spectroscopic investigation with each of the 128 detectors
The system
Measurement results

## Slide 28
Thank you
for
your attention

## Slide 29
backup

## Slide 30
PI3SO
Final operational tests:
hot spots findings
Top
Bottom
Background scan:
- 15 mins scan;
- 8.3 s/cm speed;
- 5 counts/sec;
- Background subcracted.

## Slide 31
Final operational tests:
hot spots findings
Top
Bottom
Sources free scan:
- 15 mins scan;
- 8.3 s/cm speed;
- 2.5 cm min distance from top detectors
- 2.1 cm min distance from bottom detectors
- Background subcracted.

## Slide 32
Final operational tests:
hot spots findings
Top
Bottom
1 min data acquisition on the three detected hot-spots:
137Cs, 60Co, 22Na

## Slide 33
Final operational tests:
hot spots findings
Top
Bottom
Sources lead scan:
- 15 mins scan;
- 8.3 s/cm speed;
- 2.5 cm min distance from top detectors
- 2.1 cm min distance from bottom detectors
- 1 cm thick lead plate on top of 137Cs source;
- Background subcracted.

## Slide 34
PI3SO
Cesio in crociera
Final operational tests:
hot spots findings
Top
Bottom

## Slide 35
PI3SO
Final operational tests:
hot spots finding
Lead
presence
on top of the 137Cs source

## Slide 36
PI3SO
Final operational tests:
hot spots findings
Source | X ± s X [cm] | Y ± s Y [cm]
137Cs top | 15.1 ± 3.2 | 9.6 ± 2.7
137Cs bottom | 15.4 ± 3.8 | 11.6 ± 3.2
60Co top | 30.8 ± 3.2 | 54.2 ± 2.8
60Co bottom | 30.7 ± 3.6 | 54.3 ± 3.0
22Na top | 71.7 ± 3.2 | 31.8 ± 2.8
22Na bottom | 71.7 ± 3.7 | 32.8 ± 3.3
137Cs lead top | 15.0 ± 2.5 | 9.6 ± 2.1
137Cs lead bottom | 15.4 ± 3.7 | 11.6 ± 3.1
60Co lead top | 30.8 ± 2.9 | 54.2 ± 2.8
60Co lead bottom | 30.8 ± 3.7 | 54.3 ± 3.2
22Na lead top | 71.9 ± 3.1 | 31.8 ± 2.7
22Na lead bottom | 71.9 ± 3.6 | 32.7 ± 3.2

## Slide 37
PI3SO
Final operational tests:
ROI drawing & (S – B)/B
Background scan
Sources free scan
137Cs lead shielded

## Slide 38
Background scan
NO sources on the table
Sources position on table
Sources free scan
137Cs lead shielded
X 108 cm
X 108 cm
Y (64 cm)
Y (64 cm)
PI3SO
Preliminary results
Background scan.
Sources on top of the table
Shielded 137Cs source.

## Slide 39
PI3SO
Final operational tests:
Spectroscopic measurements
Top
Bottom
Top
Bottom
Top
Bottom

## Slide 42
VIRERO (VIrtual REmote RObotics for Radiometric Sorting), is currently being developed in Germany by AiNT GmbH, Friedrich-Alexander-Universität Erlangen-Nürnberg, and Framatome GmbH, funded by the Bundes Ministerium für Bildung und Forsch-ung (BMBF)

## Slide 43
CLEANDEM The Gamma and Neutron Sensor System for Rapid Dose Rate Mapping in the CLEANDEM Project

## Slide 46
As in real applications the speed can be a relevant requirement, we also tested the data quality at a ten times faster scanning speed, i.e., with a total duration of 1.5 min. In this case, in order to improve the statistics, the plots were also binned at 1 cm along the X (sliding direction). We obtained basically the same images with worse statistics, but the hot-spots are still recognized.
Top
Bottom
Lead
Lead

## Slide 48
PI3SO
Conclusions
Fast scan, results in line with expectations
Precision measurements; no need of operator
Detectors: low cost, compact, robust, commercial components
Reliable, repeatable, safe operations
Digital data storage, full scan history available