## Slide 1
18/03/2025
Characterization of IXPE Gas Pixel Detectors with the X-ray Calibration Facility
Stefano Tugliani
Simone Maldera, Andrea Frassà, Raffaella Bonino, Nicolò Cibrario, Luca Latronico, Alessio Gorgi

## Slide 2
APPENDIX A
2
Characterization of IXPE Gas Pixel Detectors with the X-ray Calibration Facility
18/03/2025
The Gas Pixel Detector GPD
V drift = -2800 V
V top = -840 V
V bottom = -400 V
V ASIC = 0 V
ADC
Photoelectron
emission direction
: angle of emission of the photoelectron wrt the polarization direction
10
mm
0.7
mm
ASIC
GEM
DRIFT PLANE
Dymethyl-ether
DME
Interaction point
Final absorption
point
Final absorption
point
Interaction point
Interaction point
Final absorption
point
Barycenter
Emission direction
Ref. [2]
Detector
Unit
IXPE

## Slide 3
APPENDIX A
3
Characterization of IXPE Gas Pixel Detectors with the X-ray Calibration Facility
18/03/2025
The X-ray Calibration Facility
X-ray tube: radiation source
GPD
CMOS camera
SDD
Handling system
Beryllium windows transparency measurements for future X-ray detectors
Gas Pixel Detectors:
study of systematic effects that take place in the detector
GPD gain map
Unpolarized beam line
w1/air
w2/air
Ref. [2],[5],[6]

## Slide 4
APPENDIX A
4
Characterization of IXPE Gas Pixel Detectors with the X-ray Calibration Facility
18/03/2025
The X-ray Calibration Facility
Polarized beam line
GPD
CMOS camera
SDD
X-ray tube: radiation source
Plane of incidence
Diffracted polarized radiation
Unpolarized incident radiation
Polarizing crystal
Handling system
GPD response to the polarization:
Modulation factor (response to 100% linearly polarized radiation) vs Energy.
B
A
Ref. [2],[3],[5],[6]

## Slide 5
5
Astrophysical X-ray Polarimetry with IXPE
THANK YOU FOR YOUR ATTENTION
18/03/2025

## Slide 6
APPENDIX A
6
Astrophysical X-ray Polarimetry with IXPE
Bibliography
[1] Weisskopf et al. The Imaging X-Ray Polarimetry Explorer (IXPE): Pre-Launch. 2021. doi: 10.48550/ARXIV.2112.01269. url: https://arxiv.org/abs/2112.01269.
[2] L. Baldini et al. “Design, construction, and test of the Gas Pixel Detectors for the IXPE mission”. In: Astroparticle Physics 133 (2021), p. 102628. issn: 0927-6505. doi: https:// doi.org/10.1016/j.astropartphys.2021.102628. url: https://www.sciencedirect. com/science/article/pii/S0927650521000670.
[3] Fabio Muleri Sergio Fabiani. Astronomical X-ray Polarimetry. Aracne, 2014.
[4] Fabio Muleri. “ON THE OPERATION OF X-RAY POLARIMETERS WITH A LARGE FIELD OF VIEW”. In: The Astrophysical Journal 782.1 (Jan. 2014), p. 28. doi: 10.1088/ 0004-637x/782/1/28. url: https://doi.org/10.1088%2F0004-637x%2F782%2F1%2F28.
[5] S. Tugliani et al. “IXPE Gas Pixel Detectors characterization using the X-ray Calibration Facility”. In: Nuovo Cim. C 47.3 (2024), p. 141. doi: 10.1393/ncc/i2024-24141-9.
[6] S. Tugliani et al. “IXPE Gas Pixel Detector characterization with the X-ray calibration facility”. In: Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spec- trometers, Detectors and Associated Equipment 1069 (2024), p. 169800. issn: 0168-9002. doi: https://doi.org/10.1016/j.nima.2024.169800.
[7] F. Kislat et al. “Analyzing the data from X-ray polarimeters with Stokes parameters”. In: Astroparticle Physics 68 (Aug. 2015), pp. 45–51. doi: 10.1016/j.astropartphys.2015. 02.007. url: https://doi.org/10.1016%2Fj.astropartphys.2015.02.007.
[8] C. Tomaiuolo et al. " Time-dependent instrumental effects in IXPE: Pressure variation and GEM charging inside GPDs ". In: Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment, 1069 (2024), p. 169881. issn: 0168-9002. doi: https://doi.org/10.1016/j.nima.2024.169881.
16/10/24

## Slide 7
APPENDIX A
7
Astrophysical X-ray Polarimetry with IXPE
IXPE: Imaging X-ray Polarimetry Explorer
Solar array
Boom w/ thermal Sock deployed
Mirror Module Assembly (x3) (Wolter I)
X-ray Shields (x3) deployed
Forward star tracker
Detector Unit (x3)
18/03/2025
Ref. [1]

## Slide 8
APPENDIX A
8
Astrophysical X-ray Polarimetry with IXPE
IXPE data processing 1
Level 1 data:
Event reconstruction
in detector coordinates
Level 2 data:
Data in sky coordinates
Celestial Coordinates
3 Detector
Units
Star Tracker:
Celestial
Coordinates system
DU1
DU2
DU3
Detector Coordinates
12/10/24

## Slide 9
APPENDIX A
9
Astrophysical X-ray Polarimetry with IXPE
Event in Detector Coordinates
Event in Celestial Coordinates
Polarization information: recovered on a statistical basis from the azimuthal distribution of the photoelectron directions of emission.
Photoelectric differential cross section:
The total counts are modulated as:
IXPE data processing 2
Azimuthal angle : photelectron emission direction.
Polarization angle:
Polarization Degree: derived from the Stokes parameters and from
Ref. [3], [4], [7]
12/10/24

## Slide 10
10
Astrophysical X-ray Polarimetry with IXPE
GPD long term studies
GPD long term studies: secular pressure decrease
Secular decrease of the pressure in the sealed gas cell, due to internal adsorption of DME with a time scale of months.
The pressure decrease can be retrieved analyzing:
Increase of track
length
Increase of gain
Decrease of rate
Ref. [8]
16/10/24