## Slide 2
ID03, the ESRF Dark-Field X-ray microscopy beamline
C. Detlefs (ESRF), on behalf of the ID03 team:
Helena Isern
Raquel Rodriguez Lamas
Thierry Brochard
Thomas Dufrane
Can Yildirim
and many collaborators

## Slide 3
T. Brochard, T. Dufrane, H. Isern. C. Detlefs
ESRF Support Services

## Slide 4: How to build a beamline
Figure out what the beamline should do
Scientific case
Technical concept  Conceptual Design Report (CDR)
Figure out how to do that
Technical Design Report (TDR)
Source
Optics
End station

## Slide 5: Dark field x-ray microscopy
Magnified diffraction topography
using a lens between sample and detector
You take a sample and align a Bragg reflection
You put an imaging detector in the diffracted beam
You insert an objective lens between the sample and the detector
Diffraction topography
Rocking curve imaging
X-ray microscopy
H. Simons et al. (2014). Dark-field X-ray microscopy for multiscale structural characterization. Nature Communications 6, 6098. DOI: 10.1038/ncomms7098

## Slide 6: Dark field x-ray microscopy
Magnified diffraction topography
using a lens between sample and detector
Detector resolution 1 μm
Magnification 10x – 30x
Resolution at the sample 100 nm – 30 nm.
Advantages:
Full field method,
High spatial and angular resolution,
Diffraction
H. Simons et al. (2014). Dark-field X-ray microscopy for multiscale structural characterization. Nature Communications 6, 6098. DOI: 10.1038/ncomms7098

## Slide 7: ID03: Scientific case
Metallurgy
Pattern formation
Materials fatigue
Recovery and recrystallization
Biominerals
Microstructure
Image
Ceramics
Dislocation- toughening
Nano- twinning
Functional materials
Strain at grain boundaries
Formation of domain patterns
Dynamics of domain switching
3D strain maps with 150 nm spatial resolution
C. Yildirim et al. (2020). Probing nanoscale structure and strain by dark-field x-ray microscopy. MRS Bulletin, 45, 277. DOI:10.1557/mrs.2020.89

## Slide 8
Technical concept

## Slide 9: Conceptual design of ID03/EBSL2
Boundary conditions:
Evolution of the prototype beamline ID06-HXM
Re-use the end station from ID06-HXM Key parameters:
  Typical working energy 17-33 keV, up to 60 keV for complementary techniques (3DXRD, DCT)
  Resolution ~150 nm, diffraction limits ~60 nm
Design new optics fully optimized for Dark-Field x-ray microscopy Pink beam and monochromatic beam
Rebuild beamline infrastructure from ground up
Budget envelope ~4 MEuros, excluding in-house man power
M. Kutsal et al. (2019). The ESRF dark-field x-ray microscope at ID06. IOP Conference Series: Materials Science and Engineering, 580, 012007. DOI: 10.1088/1757-899X/580/1/012007.

## Slide 10: Technical design of ID03/EBSL2: priorities (1)
Dark-field x-ray microscopy is a notoriously photon-starved technique
Optimize photon flux
Use the best available source device
  New generation of cryogenically-cooled permanent magnet undulators (CPMU)
  Machine project for minimum gap 4 mm (currently 5 mm)
Allow “pink beam” operation
  Band width ΔE/E ~ 10-2 given by peak width of the undulator harmonic
  gain of ~ x100 compared to Si(111) crystal monochromator

## Slide 11: Technical design of ID03/EBSL2: Priorities (2)
Spatial resolution at present ~150 nm, future target ~60 nm
 Optimize stability
Optical elements with horizontal deflection
  Ground vibrations are mainly vertical  horizontal optics less sensitive
  Vertical source size is still smaller than horizontal source size
Thermal stability of end station
  High performance HVAC
  Use of granite for support frames

## Slide 12: Conceptual design of ID03/EBSL2
CDR was approved
“Gate B letter” 11/03/2021
Marching order to develop the Technical Design Report (TDR)
Total budget envelope of 4 Meuros
Implementation on sector ID03
Target date September 2023 for start of user operations.
Implicit: Allocation of internal resources for design studies
Project team
Helena Isern (Beamline operations manager)
Thierry Brochard (Mechanical engineer)
Carsten Detlefs (Beard)
Thomas Dufrane (Beamline technician)

## Slide 13: How to build a beamline
Figure out what the beamline should do
Scientific case
Technical concept  Conceptual Design Report (CDR)
Figure out how to do that
Technical Design Report (TDR)
Source
Optics
End station

## Slide 14
Technical concept
The design process was iterative, so difficult to present in a linear fashion.
Starting points were the ID06 optics and the new ID27, which has similar specifications in terms of photon energy, source device, etc.
We optimized and modified from there.

## Slide 15: How to build a beamline
Figure out what the beamline should do
Scientific case
Technical concept  Conceptual Design Report (CDR)
Figure out how to do that
Technical Design Report (TDR)
Source
Optics
End station

## Slide 16: Source device
Reducing the gap from 5 mm to 4 mm closes the “dark” energy range between the fundamental and 3rd harmonic.
Credit: Ph. Brumund, R. Versteegen
Disadvantage: At 4 mm gap, the heat load becomes …. interesting.
We had to do a lot of heat load studies and re-design some optical components  Front end apertures
 Primary slits

## Slide 17: How to handle the heat load
Design choices:
Front end apertures 2mm (h) x 1 mm(v) Leave most of the heat load in the Front End 17.4 kW  3.4 kW
Credit: Ph. Brumund
The complete technical design is made for operation with 4 mm gap
We are still waiting for the ASD to develop the minigap straight sections
 Until further notice, the minimum operation gap is 5 mm

## Slide 18: How to build a beamline
Figure out what the beamline should do
Scientific case
Technical concept  Conceptual Design Report (CDR)
Figure out how to do that
Technical Design Report (TDR)
Source
Optics
End station

## Slide 19: Optics overview
Primary slits
High power attenuators
Double multilayer monochromator
Diagnostics
Crystal monochromator
Transfocator
Diagnostics + beam shutter
Modules, each on a granite support

## Slide 20
Double multilayer monochromator
Design principle:
Organized in “blocks”
Diagnostics elements after each optical element

## Slide 21
Double multilayer monochromator: Modified ESRF standard design
(Carole Clavel, Philipp Brumund, Christian Morawe, Ray Barrett, PAMU)
R. Baker et al (2013) New generation mirror systems for the ESRF upgrade beamlines, J. Phys. Conf. Series 425, 052015 (2013), DOI: 10.1088/1742-6596/425/5/052015

## Slide 22
Multilayer specifications:
300 mm long
Angle of incidence: 3.5 – 7.4 mrad (0.2 – 0.4 deg)  projected size 1 – 2 mm
3 multilayer stripes
Stripe | #1 | #2 | #3
Multilayer system | [Pd(3.50)/C(5.00)]20 | [W(2.00)/B4C(2.80)]30 | [W(1.45)/B4C(1.60)]70
Energy range (keV) | 12 – 24 | 20 – 40 | 30 – 60
Total Transmission | 82% – 93% | 66% – 87% | 72% - 88%
dE/E FWHM | 12% | 6.0% – 6.9% | 2.5% – 2.6%
Credit: Christian Morawe, ESRF

## Slide 23
Credit: Ph. Brumund, ESRF
Ph. Brumund et al: Thermal optimization of a high-heat-load double-multilayer monochromator, J. Synchrotron Rad. 28, 1423 (2021) DOI: 10.1107/S160057752100758X

## Slide 24
Diagnostics block

## Slide 25
Diagnostics block 3:
(Thierry Brochard, Thomas Dufrane)
Emergency white beam stop
Beam limiting aperture
Secondary slits Water cooled, new design
Intensity monitor back-scattering diamond with metal coatings 4-quadrant photodiode
“Beam viewer” Diamond with video camera
Beam

## Slide 26
Crystal monochromator

## Slide 27
Crystal monochromator:
Subcontracted to Cinel (Italy)
Channel cut Si(111) crystal
Liquid Nitrogen cooled
Horizontally deflecting
8 mm offset
12 – 60 keV
In-vacuum:
  In-out translation (y)
  Main Bragg rotation
  Second surface pitch fine-tuning
Channel cut crystal produced in-house

## Slide 28
Trade-off: Large offset vs. small offset
Large offset:
Allows larger angles/lower energies
For channel cut: Easier access to optical surfaces
Small offset:
Smaller crystal size/less travel when changing Bragg angle
For fixed offset: Smaller variation of beam offset
Ph. Brumund et al, Design Simulations Of A Horizontally Deflecting High-Heat-Load Monochromator, J. Synchrotron Rad 28, 91 (2021) DOI: 10.1107/S1600577520014009

## Slide 29
Channel cut
Advantages:
Mechanical stability
Good thermal contact to second “crystal”
Simpler mechanics
Disadvantages:
Crystal surfaces more difficult to polish
No “roll” motion possible (probably not needed)
Fixed exit not possible (but change in offset negligible for high energies)

## Slide 30
Transfocator

## Slide 31
Transfocator: In-house development (now ESRF standard)
9 Actuators
Diamond lenses (Palm Scientific) R=320 µm, Aperture 1mm R=640 µm and R=1280 µm for fine tuning
Alignment motions ex-vacuum
New flexor design to align lens cartridges with reference V rail
Al lenses would melt in the pink beam!

## Slide 32
Transfocator:
CRLs are highly chromatic (f α E2) Need to vary number of lenses when changing working energy
Either focus onto sample position
Or Collimation
Actuator | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
Radius | 1280 | 640 | 320 | 320 | 320 | 320 | 320 | 320 | 320
Number | 1 | 1 | 1 | 2 | 4 | 8 | 16 | 16 | 16

## Slide 33: How to build a beamline
Figure out what the beamline should do
Scientific case
Technical concept  Conceptual Design Report (CDR)
Figure out how to do that
Technical Design Report (TDR)
Source
Optics
End station

## Slide 35
End station: Dismantled from ID06-HXM, re-installed on ID03
Upgrades:
EH1 is ~10 m further downstream  a bit more space around the bench
High-stability HVAC +/- 0.1 deg C Large HVAC installation on roof of CC4 Thermal insulation on walls, false floor
New goniometer (LAB) Optimized for topo-tomo scans in vertical plane More space and better xyz positioning of the sample with Hexapod
Upgrade of far field detector Back-illuminated sCMOS camera with ~2x better detection efficiency Move the previous sCMOS camera to the near field camera

## Slide 36
New goniometer
LAB motion systems
2 air bearing rotations
Hexapod for sample xyz and sample tilts
Slip rings for user electrical signals
Hexapod positioning not quite as precise and reproducible as we would like
Software coordinate system to be worked out
One of the air bearings seized  back at manufacturer