## Slide 1: CoCo – Cockroft Walton Feedback Control Circuit
Deepak G, Paul T, Vladimir G
On the behalf of

## Slide 2: Motivation for HV and its regulation…
3” Photo Multiplier Tubes (PMT) operate with high voltage (HV) in the order of kV.
The gain of the PMT varies linearly with HV.
The HV is generated using a Cockroft Walton (CW) multiplier circuit.
This HV should be regulated in order to have stable operation of the PMT.
No second source for the present driver

## Slide 3: Present Solution using COTS component
The input 7 μs DC pulses to the CW multiplier were supplied with a load dependent frequency between 1 and 5 kHz.
The circuit had stability problems at lower HV outputs.
The output load current rating was too high.
House keeping power high because its generic functions.
The COTS is also costly at low volume (~0.5M).
AD 1111

## Slide 4
COTS..we intend to replace
Cockroft-Walton Multiplier
PROMiS

## Slide 5: Goals for CoCo
CoCo should regulate the frequency of the 6.5 μs pulses.
For example, during start-up of the HV circuit, CoCo should supply the pulses at a maximum of 50kHz and when the desired HV is reached, the pulses are slowed down to a few Hz based on the feedback-> Low power operation .
The width of these pulses should also be controlled by sensing the current, to avoid saturation of the transformer.
The regulation should be stable, even at lower voltages.
Should have smaller footprint, cheaper and consume less power.

## Slide 6: CoCo sample simulation
Across Capacitor
Driver out
200ns pulse
6.5us pulse
CLK

## Slide 7: Are we in line with the goal?
CoCo is designed to control and regulate the feedback of the CW HV generator.
CoCo is designed in 350nm CMOS technology, consumes ~ 1.7 mW power.
CoCo is stable even at lower CW voltages.

## Slide 8: Layout of CoCo
GND
FB_IN
OPAMP_OUT
VDD
GND
CF
CSNS
SW
GND
PSC
NSC
BG
CSNS_C
200n
6u5
CLK
1 mm x 1 mm in size
16 pads

## Slide 9: Dec 5 Tapeout, Dies received 3rd Feb.

## Slide 10: Test Board schematic
Designed in away that it can be tested on the lab table, with a few DC supplies and also with the HV PMT Base!

## Slide 11: Test board Layout

## Slide 12
1M5
100k
10u
100n
560
1p
1p
1p
1p
1p
1p
1p
220p
3.3
100p
10M
10M
1u
10M
1M5
1M5
100n
3G
GND
3.3V
HV from
PMT Base
Internal DAC test V 0->2.4V
HV DAC setting 1.9->2.7V finally from PROMIS_V2
Jumper while testing on the CoCo testboard
Jumper while testing from PMT Base
To transformer on PMT Base
CRO Monitor Points
Now3.3k
J20
J2
Now300k

## Slide 14
Bonded Dies

## Slide 15
CLK
CLK
Neg edge comp
Pos edge comp
Across Capacitor
CLK
HV Ripple
Current in the Transformer
Current Sense
6.5 us output
CLK
Bandgap
Switch output
6.5 us output
200 ns output
HV Ripple
Current in the Transformer
Current Sense
Current Sense
Current Sense
Current Sense
6.5 us output

## Slide 16: Measurements on the testboard
Chart: Testboard 3
| 2.4 | 2.3899999999999997 | 2.38 | 2.3699999999999997 | 2.36 | 2.3499999999999988 | 2.34 | 2.3299999999999987 | 2.3199999999999981 | 2.3099999999999987 | 2.2999999999999998 | 2.2000000000000002 | 2.1 | 2 | 1.9 | 1.8 | 1.7000000000000008 | 1.6 | 1.5 | 1.4 | 1.3 | 1.29 | 1.28 | 1.27 | 1.26 | 1.25 | 1.24 | 1.23 | 1.22 | 1.21
Vdd = 3.3 V | 51.3 | 51.3 | 51 | 50.8 | 50.5 | 50 | 49.8 | 49.3 | 48.8 | 48.5 | 47.8 | 43.9 | 39.8 | 35.5 | 31.3 | 26.9 | 22.5 | 18 | 13.6 | 9.1 | 4.5 | 4 | 3.45 | 3 | 2.44 | 2 | 1.5 | 1.05 | 0.6 | 0.2
Vdd = 3.0 V | 51.3 | 50.8 | 50.8 | 50.3 | 50.3 | 49.8 | 49.3 | 49 | 48.5 | 48.1 | 47.6 | 43.9 | 39.4 | 35.2 | 31.1 | 26.6 | 22.2 | 17.7 | 13.3 | 8.8 | 4.5 | 4 | 3.6 | 3.09 | 2.45 | 2.2 | 1.7 | 1.1 | 0.8 | 0.2
Vdd = 3.6 V | 51.3 | 51.2 | 51.2 | 51 | 50.8 | 50.3 | 50 | 49.6 | 49.3 | 49 | 48.3 | 44.2 | 40.2 | 35.6 | 31.4 | 27 | 22.5 | 18 | 13.6 | 9.05 | 4.29 | 4 | 3.5 | 2.9 | 2.65 | 2.07 | 1.6 | 1.25 | 0.8 | 0.2
Chart: Testboard 1
| 2.4 | 2.3899999999999997 | 2.38 | 2.3699999999999997 | 2.36 | 2.3499999999999988 | 2.34 | 2.3299999999999987 | 2.3199999999999981 | 2.3099999999999987 | 2.2999999999999998 | 2.2000000000000002 | 2.1 | 2 | 1.9000000000000001 | 1.8 | 1.7 | 1.6 | 1.5 | 1.4 | 1.3 | 1.29 | 1.28 | 1.27 | 1.26 | 1.25 | 1.24 | 1.23 | 1.22 | 1.21
Vdd = 3.3 V | 51.4 | 51.4 | 51.3 | 51 | 50.8 | 50.5 | 50.3 | 50 | 49.8 | 49.3 | 48.5 | 44.6 | 40.3 | 36.1 | 32.1 | 27.7 | 23.1 | 18.6 | 14.3 | 9.71 | 4.72 | 4.12 | 3.7 | 3.4 | 2.91 | 2.33 | 1.6 | 1.23 | 0.9 | 0.2
Vdd = 3.0 V | 51.2 | 51.2 | 51 | 50.6 | 50.4 | 50.1 | 49.6 | 49.4 | 49 | 48.5 | 48.1 | 44.1 | 40 | 36 | 31.7 | 27.5 | 23 | 18.5 | 14 | 9.39 | 4.74 | 4.3 | 3.89 | 3.3 | 2.67 | 2.3 | 1.92 | 1.4 | 0.8 | 0.3
Vdd = 3.6 V | 51.4 | 51.4 | 51.3 | 51.3 | 51.3 | 51 | 50.6 | 50.3 | 49.8 | 49.4 | 49 | 44.8 | 40.7 | 36.5 | 32.2 | 27.9 | 23.5 | 18.9 | 14.2 | 9.43 | 4.65 | 4.2 | 3.8 | 3.4 | 2.9 | 2.3 | 1.7 | 1.16 | 0.7 | 0.2

## Slide 17
Chart: Stable operation frequency
| 1.901 | 2.0009999999999999 | 2.101 | 2.2010000000000001 | 2.3009999999999997 | 2.4009999999999998 | 2.5009999999999999 | 2.601 | 2.7010000000000001
Discrete_HV_Freq. (kHz) | 0.675 | 1.03 | 1.49 | 2.05 | 2.73 | 3.6 | 4.67 | 5.98 | 7.46
CoCo_Frequency(kHz) | 0.9 | 1.2 | 1.4 | 1.6 | 2.6 | 3 | 3.7 | 4.8 | 6.3
Chart: High Voltage generation
| 1.901 | 2.0009999999999999 | 2.101 | 2.2010000000000001 | 2.3009999999999997 | 2.4009999999999998 | 2.5009999999999999 | 2.601 | 2.7010000000000001
Discrete_HV | -699 | -798 | -896.7 | -994 | -1092 | -1190 | -1287 | -1384 | -1480
CoCo_HV (1Gohm probe) | -698 | -797 | -896 | -994 | -1092 | -1190 | -1287 | -1384 | -1480
Chart: CoCo current consumption
| 1.901 | 2.0009999999999999 | 2.101 | 2.2010000000000001 | 2.3009999999999997 | 2.4009999999999998 | 2.5009999999999999 | 2.601 | 2.7010000000000001
Current (without 1Gohm) | 0.574 | 0.653 | 0.713 | 0.816 | 0.935 | 1.076 | 1.25 | 1.45 | 1.7
Measurements with the HV PMT Base!

## Slide 18
COTS..we intend to replace
Cockroft-Walton Multiplier
PROMiS
Cockroft-Walton Multiplier
PROMiS_V2
CoCo

## Slide 19: Thank you for your attention

## Slide 20: Chip Schematics
Bandgap
Bias Block
Current sense comparator
200ns Monostable
6.5us Monostable
Driver
CCO
Feedback opamp
SR Latch

## Slide 21: Current controlled Oscillator
SR Latch
FB control Opamp
Charging Current Mirror
Discharging Current Mirror
Positive Swing Comparator
Negative Swing Comparator
2V
1V
1.2 V Bandgap

## Slide 22: Monostable multivibrator
Comparator
DFF
RC = 200ns or 6.5us
1.2 V Bandgap

## Slide 23: PCB Bonding
Chip goes here! Direct bonding on the PCB
ORIENT THE LOGOS 