## Slide 1
Safety in Underground Cryogenic Systems
P. Rapagnani

## Slide 2
We deal with the main issues regarding the installation of Cryostat and use of cryogenic liquids in a confined environment.
On this topic, a vast expertise, also inside INFN, is available, e.g:
CERN with LHC
INFN with the Cryogenic Dark Matter Detectors at LNGS, ARIA facility in Sardinia
LAGUNA Underground Laboratory Studies

## Slide 13
It is more difficult to avoid these risks in an underground lab.
We can follow for instance the guidelines for Liquid Argon Dark Matter Detectors:
PSAM 12: Probabilistic Safety Assessment and Management,
Hawaii, USA 2014

## Slide 14
Following Markoulaki & Papazoglou 2014:
Main risk in Underground Cryogenics: Loss of Containment (LOC)
LN2 is released from its containment, evaporate, and:
reach concentrations in a particular confined space that can expel oxygen from the air
result to extremely low temperatures in the confined space
ODH (Oxygen Deficiency Hazard)
All containment failures may be divided in:
Structural Failures
They occur if the stress employed on the containment by the various operating conditions is larger than the stress of the containment.
Containment Bypassing
It occurs whenever an engineered opening in the containment (like a valve) opens inadvertently when it is supposed to be closed. Such failures are mainly caused by human actions either during normal operation or during test and maintenance activities (not treated here)

## Slide 15
Overpressure and Related Safety Measures
Increase of the internal pressure can cause an increase of the stress on the containment exceeding its design limits and resulting in LOC (Loss of Containment).
Accident sequences are sequences of events (either hardware failures and/or human actions) that lead to LOC and hence to LN2 release.
Example: Loss of Offsite Power
Mean Time To Repair
Mission Duration =
= “grace period”:
Time before overpressure becomes dangerous

## Slide 16
Loss of Service Water Cooling System
Initiating events:
(a) loss of a pump
(b) loss of a Heat exchanger (HX).
Logic: Success if the failed component is repaired within 2.2 hours (if instrumentation is turned off) and in 2 hours (if instrumentation is not turned off).
If the failed component is not repaired within the available grace period the sequences result in failure due to overpressure.
e.g.: Sequences #2 & #5 are accident sequences, as they involve failure of the SWCS pump with failure to recover it within the available time. Sequences #3 & #6 involve Ar release in the tank cavern since PS6 has failed. Similar accident sequences #8 & #11 and #9 & #12 are calculated for shorter grace period since the Instrumentation is left ON.
LN2 released in cavern

## Slide 17
These are just examples of the kind of analysis to be done...
Of course, for Archimedes Cryostat, the worst accidents are due to a loss of insulation vacuum.
Insulation Vacuum Pump stops
pressure builds up in
~ 5 hrs: “grace period”
Pump Repaired in time
Success
Failure to repair in time
LN2 Venting
For instance…
Experiment Switched Off
pressure builds up in
~ 10 hrs: “grace period”
Pump Repaired in time
Success
Failure to repair in time
LN2 Venting

## Slide 18
Of course, the grace period can vary from days (e.g. pumps stop with solid N2 inside) to minutes, if a gasket fails.
These cases, with the relevant weak points, must be considered, and the related fault trees identified
In this way, we can introduce redundancies and additional safety systems, where required.
In any case, release of LN2 in the cavern must be avoided as much as possible:
the venting tunnel is necessary
Even in case of release of LN2 in the closed environment, an emergency status of the air conditioning system can be implemented, to mitigate its effects.
For instance, in the ICARUS T600 experiment at LNGS, air conditioning was designed to pump eventual LAr spills through the floor (2015 JINST 10 P12004)

## Slide 19
Explorer was in operation from 1983 to 2010: 27 years
In that period, we had only some minor emergencies, due to:
Pumping system failures
Electrical power cuts
o’rings freezing during refill
which were contained with only minor damages
Some experience:
It contained ~ 600 l of LN2, refilled manually once every 14 days

## Slide 20
Conclusions
Underground cryogenics safety is challenging, but there is a wide expertise available, e.g., both at CERN and in INFN, that can be accessed.
A careful analysis of several accident scenarios must be performed, in order to increase reliability and minimize damage and down time.
Excluding really catastrophic (and unlikely) events, the time to reach overpressure could be of minutes, or even hours (to be determined with symulations and a thermodynamical analysis)
A venting to the surface is necessary.
Venting must be forced for heavier than air cryogenic fluids

## Slide 21
END