## Slide 1: CH4rLiE
Updates
9/4/25

## Slide 2: New setup in Pavia lab
CH4
100 ppm
Ar
CARTRIDGE
vent
  Simple setup to test the methane sensor
  Empty cartridge
Sensor

## Slide 3: New setup in Pavia lab
CH4
100 ppm
Ar
CARTRIDGE
vent
CARTRIDGE
  Setup similar to CERN one
  First cartridge for the moisture
Sensor

## Slide 4: New setup
18 cm
2 cm
2x smaller cartridge:
  Possibility to perfrom test with the methane sensors
  Less material needed
  Possibility to perform the high temperatue regeneration

## Slide 5: First test with the sensor
Step | CH4 [l/min] | Ar [l/min] | Plateaux [ppm]
1 | 1.3 | 0.0 | 100.03
2 | 1.3 | 1.0 | 61.58
3 | 0.9 | 1.0 | 51.31
4 | 0.9 | 0.5 | 66.50
5 | 0.9 | 0.2 | 81.17
6 | 0.6 | 0.2 | 59.36
7 | 0.6 | 0.5 | 40.60
8 | 0.6 | 1.0 | 27.40
1
2
3
4
5
6
7
8

## Slide 6: CO2 measure
  Unexpected high values of CO2
  Test with pure Ar:
    Still around 600ppm
    Probably an offset

## Slide 7: CO2 – CH4 correlation
  Unexpected correlation between CO2 and CH4
  Not a leak
  CO2 in the bottle ?

## Slide 8: To do next
  Confirm the presence of CO2 in the bottle
  Test with different integration time
  Test to see if the sensor loose precision if the flux is lower than the recommended one
  Prove the feasibility of adsorbtion tests in the lab with this setup
  Test the new material