## Slide 1: LNL Accelerator Status E. Fagotti

## Slide 2: Outline
SPES
PTA operation and status
PIAVE-TANDEM-ALPI availability

## Slide 3: Outline
SPES
PTA operation and status
PIAVE-TANDEM-ALPI availability

## Slide 4: SPES phased approach
SPES Phase 1:
Cyclotron and BL2
SPES Phase 3:
Charge breeder + MRMS + RFQ
SPES Phase 2B:
LERIB
SPES Phase 5:
RIB selection
SPES Phase 4: Radioisotopes production
SPES
Phase 2A:
RIB production
SPES Phase 1
SPES Phase 2A
SPES Phase 2B
SPES Phase 3
SPES Phase 4
SPES Phase 5
Other

## Slide 5: Phase 1 at completion: beam delivered May 2024
SPES-driver
SPES ISOL2
Phase 1
Phase 1 completed: cyclotron delivered to LNL
Target station tested.

## Slide 6: Three irradiation runs with 35, 50 and 70 MeV proton beams
Beamline characterization process using IAEA monitor reactions https://www-nds.iaea.org/medical/monitor_reactions.html
natAl(p,x)22/24Na
natNi(p,x)57Ni
natTi(p,x)46Sc
Beam
Remeasuring xs for 67Cu production route
68Zn(p,x)67/64Cu
70Zn(p,x)67/64Cu
70 MeV max
Remeasuring xs for 47Sc production route
natV(p,x)47Sc
70 MeV max
Production cross section value of the natZn(p,x)67Cu nuclear reaction

## Slide 7: Phase 2: RIB commissioning of the ISOL machine
TIS
Phase 2a
TIS
Phase 2b
Completed 11/2024
Completed 04/2025
RIB
RIB

## Slide 8: Phase 2A: RIB commissioning of the ISOL machine
TIS
Phase 2a
Completed 11/2024
RIB
TIS
Phase 2b
Completed 04/2025
RIB

## Slide 9: Phase 2A: RIB commissioning of the ISOL machine

## Slide 10: Phase 2A: First RIB production
28Si: 2+→0+
Beam ON
HPGe spectrum in coincidence with beta signal

## Slide 11: Phase 2B: RIB commissioning of the ISOL machine
TIS
Phase 2a
Completed 11/2024
RIB
TIS
Phase 2b
Completed 04/2025
RIB

## Slide 12: Phase 2B: LERIB (Low Energy Radioactive Ion Beam) transport line
June 2024
July 2024
May 2024

## Slide 13: Phase 2B: First RIB analysis at tape station
Beta plastic scintillator
(temporary arrangment)
HPGe
Moving tape
Mass 28 - 28Al
t1/2=2.25 min

## Slide 14: Phase 3: 1+ injector + Charge Breeder + Medium Resolution Mass Spectrometer + RFQ

## Slide 15: Phase 3A: 1+ injector + Charge Breeder + Medium Resolution Mass Spectrometer

## Slide 16: Phase 3B: Radio-Frequency Quadrupole
High power test results:
201 kW @ 0 dBm
Efficiency = 66.1%
1 dB BW >= 1 MHz
Linearity of ± 0.943 dB in CW in the range [20 kW, 180 kW]
Long run test for 117 hrs successfully completed
RFQ tuning results:
Q0 mode is at 78.94 MHz
Differences with simulated values less than 0.4% for all modes
Dipole-free zone around Q0 [-2.45 MHz, 1.97 MHz]
Q0 = 14000 (70% of 2D value) will improve with final tuners and plates

## Slide 17: Phase 5: Beam Cooler + HRMS
+
-
+
-
-
+
1 eV
Beam cooler

## Slide 18: Outline
SPES
PTA operation and status
PIAVE-TANDEM-ALPI availability

## Slide 19: Global Beam on Target @ LNL
Chart: YEAR
| 2009 | 2010 | 2011 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017 | 2018 | 2019 | 2020 | 2021 | 2022 | 2023 | 2024 | 2025
LNL BoT | 6032 | 5364 | 6774 | 5593.5 | 6154.5 | 5401 | 5593.5 | 5453 | 3016 | 4409.5 | 2344 | 2127 | 3701 | 4856 | 6436 | 5358.74 | 1827

## Slide 20: AGATA Beam on Target @ LNL
Chart: YEAR
| 15-21 (GANIL) | 2022 | 2023 | 2024 | 2025 | 22-23-24-25
Beam on Target | 2560 | 3476 | 3115 | 1424 | 10575
AGATA Beam on Target | 6568 | 1712.5 | 3073.5 | 2282.07 | 1011.4 | 8079.47
Beam Preparation | 541 | 1131.5 | 832.33 | 83.7 | 2588.53
AGATA Beam Preparation | 368.5 | 869 | 624.5 | 58.4 | 1920.4
Accelerator Division Tests | 240 | 214 | 480 | 0 | 934
Unscheduled Maintenance | 461 | 716.5 | 468.93 | 513 | 2159.43

## Slide 21: TANDEM & Injector status
Laddertron chain reached 6750 hrs with reduced degradation
7300 hrs operation will be reached by the end of July
Till now, 1011/1424 hrs operation reserved for AGATA
513 hrs unscheduled maintenance, all related to the injector
The injector recovered performance after transformer maintenance and vacuum pump replacement
New injector ionizers arrived at the end of June. The bake-out process will give us the first information on the feasibility of Ca and Mg shifts. In any case, a Ca beam test will be needed before the green light Ca and Mg beams

## Slide 22: PIAVE-ALPI status: ECR upgrade
URANIUM COULD BE PRODUCED ONLY BY SPUTTERING BUT…
IT CANNOT BE TESTED AS A CONVENTIONAL BEAM SO….
HOW CAN WE DEDUCE ITS FEASIBILITY?
0.7 µA
32+
Uranium by sputtering @ GANIL
5.5 µA
181Ta
Another beam produced with the same technique
181Ta @ LNL
9 µA
THE LNL ECR SOURCE WAS ABLE TO PRODUCE Ta INTENSITIES COMPARABLE (IF NOT HIGHER) THAN THOSE PRODUCED AT GANIL. IT IS THEN REASONABLE TO EXPECT THAT THE PRODUCTION OF 500 nA OF U32+ WITH THE SAME TECHNIQUE IS FEASABLE.

## Slide 23: ECR needed upgrades
Sputtering system connected to LEGIS
To dismount safely the sputtering system with U a further piece with a valve must be installed
IN THIS CONDITION THE SPUTTERING SYSTEM CANNOT BE CONTAINED IN THE SOURCE BOX
THE SOURCE BOX THEN THE HV PLATFORM MUST BE ENLARGED BY 50 cm
THE EXECUTIVE PROJECT FOR THE PLATFORM MODIFICATION IS READY. RESEARCH OF QUALIFIED COMPANIES TO PUT IT IN PRACTICE IS ONGOING

## Slide 24: PIAVE - ALPI maintenance
PIAVE-ALPI maintenance will start after summer:
Vacuum upgrade of the low-beta section
Tuners’ vacuum connection special maintenance
Valve boxed upgrade
PIAVE and ALPI conventional maintenance

## Slide 25: Outline
SPES
PTA operation and status
PIAVE-TANDEM-ALPI availability

## Slide 26: PTA availability
TANDEM operation will go on, as scheduled, till August 1st
A fast maintenance is foreseen in September, but there will be a delay related to the stripper foils replacement. Users’ shifts will start again on October 6th
Next December, Laddertron will overcome 9500 hours of operation. If and only if the January 2026 maintenance confirms Laddertron’s acceptable status, we will proceed with the same configuration till April 3rd.
On April 13th, the PIAVE-ALPI operation will be resumed, and new Uranium beams will be available. All shifts requiring Uranium beams will be grouped. At the end of the Uranium campaign, some days will be needed to replace the ECR plasma chamber. It is highly recommended to schedule the uranium campaign at the end of the experimental campaign to minimize risks related to the above-mentioned operation
No PTA beam is foreseen between the end of July 2026 and the end of December 2026

## Slide 27
Thank you