## Slide 1: IRIS – WP4 Superconducting Magnet Lab
Carlo Santini – Enrico Beneduce
14/07/2025

## Slide 2: New cryostats
Thermal design of the new service flanges for inner SOLEMI and Hilumi cryostats
  Final design of the gas-solid heat exchangers
  Final dimensions of the optimized current leads
Redefinition of technical requests for the new 20K / 55K GHe cooling system
  Calculations to support official communications with VS&S/CryoPride for supply details
SOLEMI 1 revamping:
  Cooldown in liquid nitrogen
  Power supply test
  Electrical insulation test
  Vacuum issues

## Slide 4: New service flanges for VT Test
HiLumi cryostat
Solemi insert cryostat
Re-use of exhisting tanks with updated top flanges

## Slide 5: New service flanges for VT Test
HiLumi (GHe)
SOLEMI GHe
SOLEMI cryocooler
GHe circulation system

## Slide 6: 20K GHe cooling system
Losses [W] | 1 Bize | 2 Bohmwind | 1 Bohmwind
From cryofan inefficiency | 33 | 14,3 | 8,7
From friction losses in the application | 3.8 | 6,3 | 2,7
Conduction inside the cryofan | 17 | 16 | 8
From transfer lines | 19 | 19 | 19
Application | 20 | 20 | 20
Total | 83,8 | 75,6 | 58,4
 | 1 Bize | 2 Bohmwind | 1 Bohmwind
[K] | 24.3 | 21.9 | 20.3
[K] | 26.0 | 23.4 | 21.9
 | 1 Bize | 2 Bohmwind | 1 Bohmwind
 | 24.3 | 21.9 | 20.3
 | 26.0 | 23.4 | 21.9
VS
Accepted solution

## Slide 7: Gas temperature
With 52W /C.L. + 10 W load on the thermal shield, 4.3g/s flow:
T cryocooler [K] | T gas inlet [K] | T gas outlet [K]
30 | 31 | 36

## Slide 8: Heat exchange plates for CL
Outer side
Inner side
GHe in
GHe out
Drill
Plug
GHe in
GHe in
Swagelock 90° mitre bend fitting
Current leads are cooled in series
Heat load for one optimized current lead 300K-60K: 46W

## Slide 9
Pressure drop
Hydraulic correlations (Miller, Internal flow systems), fixed-properties GHe at 40K, 4bara
Main loss from 90° mitre bends inside the exchangers (“plates”) and corrugated ½ inch pipes
Resulting for a Ø16 mm channel
Friction losses VS cryofan capacity
Resulting heat exchange coefficient

## Slide 10: Current leads insulation
Kapton MT ®
Current lead
Exchanger
Pressure applied with 8 M8 bolts
Differential contractions: brass bolts
Resulting applied average pressure: 4MPa
Kapton MT ® total heat conductivity at this pressure, 40K: 390 (Cryogenics 39 (1999) 803-809)
Conductivity at zero applied pressure: 320

## Slide 11: Simulations - GHe
OLD design! Illustration purpose only
Optimized section
Electrical feedthrough
Section for heat exchange
Exchangers channels
Contact with HTS current lead (must be T<64K, target 60K)
=1000 A
Kapton interfaces
Heat conduction through channels walls: ; (second lead)
Neglecting all electrical interface resistivities
Main results (similar for SOLEMI and HiLumi):
 | [K] | [K] | Heat to gas [W]
 | 60.8 | 75-80 | 49
 | 63 | 75-80 | 49
 |  |  | Heat to gas [W]
 | 60.8 | 75-80 | 49
 | 63 | 75-80 | 49

## Slide 12: Simulations - Cryocooler
Kapton interface
Cu-In-Cu
=500 A
60W?
55K
By reducing the current to 300 A (thermal load≈40W) temperatures may decrease by 10K!

## Slide 13: Last updates from suppliers’ requests
Swagelock pipes
Screw-on heat exchanger
Thermal contact on the donward surface of the cryocooler
Commercial copper braids (not shown)

## Slide 14: Rescheduling
Final drawings INFN → SRV: officially on 31/01/2025
Original due delivery date: start production + 200gg = end of July 2025
Due to last changes, significant delay → end of semptember 2025, not sure.
Mitigations (tender objectives splitting) will be applied if necessary, similar to other tenders

## Slide 15: Solemi 1 - cooldown
~ 400h cooldown

## Slide 16: Temperature rise after stop cooling and vacuum leak
date time | frac day | hours | mbara | mbar/h | T [K]
30/06/2025 16:21 | 0 | 0 | 4,10E-03 |  | 120
01/07/2025 12:49 | 0,85 | 20 | 5,90E-03 | 8,79E-05 | 
03/07/2025 10:32 | 2,75 | 66 | 1,20E-02 | 1,33E-04 | 
04/07/2025 11:56 | 3,82 | 92 | 1,80E-02 | 2,36E-04 | 170

## Slide 17: Power supply tests with ITECH 32V – 480A

## Slide 18: Power supply tests with ITECH 32V – 480A

## Slide 19: Electrical insulation test
Duration [s] | Nom. Voltage [V] | Peak voltage [V] | Avg. Resistance [GΩ] | Avg. Current [nA]
30 | 25 | 34 | 2.426 | 14.1
30 | 50 | 58 | 2.58 | 22.8
30 | 75 | 83 | 2.66 | 31.4
Test with MEGGER. Insulation coils-to-ground.

## Slide 20
Thank you for attention!

## Slide 21: Final thermal design of the new service flanges for inner Solemi and Hilumi cryostats
MDC precision electrical feedthrough
200mm(GHe)
150mm (Cryocooler)
?

## Slide 22: 2-materials optimization
(brass /steel)
(Cu)
Universal optimized solution