## Slide 1
Ph.D. Second Year Progress report
Surface Impedance of Superconductors Under Conditions of
Interest for Fundamental Physics
Measurements and Methods
Curriculum: Elettrotecnica ed Elettrotecnica per Acceleratori
Student: Gaurav Rajesh Mota, gauravrajesh.mota@studenti.unipd.it Advisor: Prof. Nicola Pompeo, University Roma Tre, nicola.pompeo@uniroma3.it
Co-advisor: Prof. Enrico Silva, University Roma Tre, enrico.silva@uniroma3.it
EleMeas Lab,
Head: Enrico Silva
Electrical, magnetic and electrodynamic
Measurements Laboratory
CryoEQS Lab,
Head: Nicola Pompeo
Cryogenics and Electrodynamics of Quantum solids
and Superconductors Laboratory

## Slide 2
AIM
Agenda
Study of Surface Impedance of super conductors in High Magnetic fields.
Motivation
High field Super conductivity
Experimental Techniques
Research activities
School & Course

## Slide 3
Haloscopes for Axions
The FCC-hh beam screen
Motivation

## Slide 4
FCC Beam screen
Gaussian bunches ( protons, 8 cm long)
Image currents flow on the beam screen GHz
The e.m. perturbation generates instabilities
Grow rate of instabilities
Surface resistance
With copper - at 50 K ; beam lost after 25 turns
We need materials with lower
High Temperature Superconductor (HTS)! (High frequency, fields, temperature)?
Beam screen
Particle bunch
Image current
S. Calatroni, IEEE Trans. Appl. Supercond., 26(3), 2016

## Slide 5
The search for Axions
(QCD) Axions:
– Peccei-Quinn proposed solution for Strong CP problem
– Candidate for (at least a part of) Cold Dark Matter
Detection scheme
– Axion photon photon-coupling:
: axion-photon coupling)
– Conversion in microwave photon collected in a high-Q cavity
→ haloscopes for the detection of galactic halo axions
Signal power:
Alimenti, A.; Torokhtii, K.; Di Gioacchino, D.; Gatti, C.; Silva, E.; Pompeo N., Instruments 2022, 6, 1.

## Slide 6
Electrodynamics of High Field
Super Conductivity

## Slide 7: Vortex Motion
Surface impedance
superconductors in a dc field in Mixed state Vortices
→ moving vortices → e.f.m. → finite complex resistivity
: Complex Resistivity
As function of frequency , applied dc field H & temperature T.
: basic physics, material property, applications...
microwave
e.m. field
F
Jmw

## Slide 8: Frequency Dependence
flow
pinning
&creep
Frequency dependence of vortex motion : many models, one equation
Re(vm)
Im(vm)
Pompeo, Silva, Phys. Rev. 78, 094503 (2008)
Focus on
NEED FOR WIDEBAND EXPERIMENTAL TECHNIQUES

## Slide 9
Wide Band Measurement Techniques
Dielectric-loaded resonator
Corbino disk

## Slide 10: Dielectric-Load Resonator

## Slide 11: Corbino Disk Measurements
The can be determined by impedance of the sample by following relation
where is impedance of the coaxial line (std: 50)
Where d is thickness of the sample , is complex conductivity
and , are inner and outer diameter of the corbino disk.
Rs/Z0 in the superconducting MgB2 thin film
E. Silva, N. pompeo et al Supercond. Sci. Technol. 24 (2011) 024018

## Slide 12
Current Research Activities

## Slide 13
TE011 - mode
TE021 - mode
DR: Measurement and Data Analysis

## Slide 14
14
Line calibration
Calibration of coaxial line under cryo temp and magnetic fields
: measured reflection coefficient
: actual reflection coefficient
: Error parameters Determine with 3 measures with std.
: signal reaches the detector directly without interacting with the sample
: damping and phase shift in the transmission line
:signal coming from the sample is reflected again and adds to the signal
approaching the sample
# standard calibration procedure not feasible with cryostat
Need for proper calibration suitable in the required conditions.
E. Silva, N. pompeo et al Supercond. Sci. Technol. 24 (2011) 024018

## Slide 15: Calibration scheme used wide
Method | Advantages | Limitations | Used For
OSL (Open-Short-Load) (Scheffler & Dressel, RSI 2005) | Simple, standard method Compatible with Corbino | Sensitive to parasitic effect Connector repeatability 4 days for one set of mes. | Broadband Corbino thin-film superconductors
DBC (Data-Based Calibration) (Wang et al., arXiv 2021) | No physical standards needed Works in dilution fridges | Requires reproducible data Sensitive to modelling errors | Cryogenic setups with limited access
TRL + Cryo Switches (Anlage Lab, RSI 2013) | In-situ calibration Reduces thermal drift | Expensive switches Complex integration | Multi-port precision superconducting devices
Superconducting Normal-State Ref. (Zingler et al., arXiv 2018) | Eliminates geometry mismatch Material-specific | Only works if Tc well-defined | Thin films with sharp superconducting transitions

## Slide 16: Motivation & Methodology
Vector Network analyzer
Corbino cell
Calibration std. OSL
Temp
sensors
VNA
1
2
Calibration performed
(Cable –end to end )
Standards and DUT
Reference plane
RF cables affect measurement accuracy:
– Phase delay
– Insertion/return loss
– Connector repeatability
Even 'identical' cables may differ due to manufacturing tolerances.
Goal: check if cables are interchangeable in calibration chain.
1-port OSL calibration performed for cable.
4 identical cables tested on two VNA channels (CH1 & CH2).
For each cable: extract error parameters and apply correction

## Slide 17: Cable 1 CH 1

## Slide 18: Comparison of Error para. for all cables

## Slide 19
Cross Characterization
Applying error parameters from one cable to DUT data measured with another test's transferability of calibration.
Selected EP - Selecting error parameters of different cable which show similar behaviour that as od the cable measuring DUT
Any divergence shows sensitivity to cable-specific effects.
Correction did not work; it means cables are not effectively interchangeable.
VNA Cal
Model Cal
Selected EP
Random selection of EP

## Slide 20: Motivation to changed Approach
4-Cable Calibration Issues:
- Each cable introduces unique insertion loss and phase delay.
- Calibration is time-consuming and requires repetition for each cable.
- Not interchangeable: error correction not transferable.
Microwave Switch (MM5230-03NDB-TR) Advantages:
- MEMS switch with ultra-low insertion loss (~0.3 dB) up to 26.5 GHz.
- Enables repeatable, automated calibration with fixed reference plane.
- Switching time ~20 µs, >3 billion cycle lifetime.
- Potentially suitable for cryogenic environments and compact integration.
Outcome: Improved calibration stability, reduced human error, and faster experimental cycles.
Vector Network analyzer
Corbino cell
Microwave switch
Circuit board
Calibration std. OSL
Temp
sensors

## Slide 21
Corbino Disk Cell Design old
Design Requirements @cryo T and High H
Mech. Stabilities
Size of magnet bore
Size of samples
Contact btw coaxial line & sample
Stable thermalization (Heaters and T sensors)
Top cover
Female 4-Hole Flange Launcher
Superconducting sample
Sample holder
Temp sensor
central screw to ensure good contact b/n sample and launcher
Bottom support

## Slide 22: Updated Corbino Disk Cell Design
Vector Network analyzer
Corbino cell
Calibration std. OSL
Temp
sensors
Microwave switch
Circuit board
Top cover
Female 4-Hole
Flange Launcher
Superconducting sample
Sample holder
Temp sensor
central spring to ensure good contact b/n sample and launcher
central screw to extract sample holder
“The redesign aims to reduce thermal drift, improve contact repeatability, and enable integration of microwave switching circuit for using single cable from VNA to the circuit board.
Slot for adding heating coil

## Slide 23
Electrical Characterization of Ti Thin Film
using Van der Pauw Method

## Slide 24: Introduction
Thin metal films (like Ti) are crucial in electronics, superconductivity studies, and as reference standards.
Titanium films can act as resistive standards, contact layers, or be tested as candidate calibration loads.
The d.c. electrical resistivity tells us about the film’s quality, morphology, and suitability for applications.
Unlike simple 2-probe methods, Van der Pauw is geometry-independent and reduces contact resistance effects

## Slide 25: Theory
4 small contacts placed at the edges of the thin-film sample.
Works for any shape if:
– Film thickness is uniform, no holes/isolated regions.
– Contacts are small relative to sample area.
Advantage: avoids geometry dependence and contact resistance errors.
Van der Pauw relation:
Symmetric case: (Ω/sq)
Resistivity: (Ω·m)
Where t (m) is thickness of the film.

## Slide 26
Sample Details
Material: Titanium thin film deposited on silicon substrate(500um).
Thickness TI films : 50 nm & 100 nm.
Deposition method: sputtering .
Geometry: Square (4mm*4mm)
Contact pads (0.5mm*0.5mm)(of Ti)
Contact pads: Four small contacts at corners A, B, C, D (with silver paste).
Sample
Contact pads

## Slide 27: Results
Rs(T) exhibit not metallic behaviour, but reminiscent of semiconductors
Is the Si substrate probed in the measurement?
New set of dielectrics to be precured and tested.
Ti film 50 nm
Ti film 100 nm
New connections made with contact pads
1st measurements
Sample isolated from sample holder using a Teflon sheet
T.G. Berlincourt Phys. Rev. 114, 969 (1959)

## Slide 28: Conclusion & Future Work
4-cable calibration tests showed that cables are not effectively interchangeable → error correction not transferable.
Redesign Corbino disk cell with integrated microwave switch for automated, repeatable calibration.
Next steps: • Test switch-based calibration in cryogenic setup. • Begin systematic measurements of Ti films on new set of dielectrics to be used for calibration loads .
Final goal: link experimental results with vortex dynamics models, providing reliable broadband impedance data for superconductors under relevant physical conditions.

## Slide 29: School and Course
Joint Universities Accelerator School 2025 (9 CFU).
Course 2: The Technology & Applications of Particle Accelerators (118h15)
Attended full curriculum Lectures (71h15), Seminars (9h), Workshops (21h), Visits (17h)
Advance scientific programming in Matlab (6 CFU)

## Slide 30
Thank you
GAURAV RAJESH MOTA