## Slide 1: R&D ongoing activities
Bartlomiej Rachwal, Wander Baldini
4-th SuperB Collaboration Meeting, March 31 – June 4 2012, La Biodola (isola d’elba)

## Slide 2: R&D: Ongoing Activities
The choice of BIRO readout for the barrel requires some additional R&D studies: Light Yield studies for 10cm scintillator bars:
  Light yield vs number of fibers
  ITEP scintillator, length=25cm, wrapped in aluminum foil
  Kuraray Y11(300) fibers, f=1.2mm, length ≈ 40cm
  3x3 mm2 MPPC, 50x50 microns pixels
Preparation of the irradiation test at the GELINA facility @ IRMM (Geel Belgium):
  Very useful visit to IRMM on May 8 to make arrangements and get more informations
  Data taking on July 9-20

## Slide 3: Preliminary results on light yield with 10 cm scintillator bar

## Slide 4: Calibration
Working conditions with cosmics
Gain = 10.5 ADCch/p.e.
Vbias= 72.5V
T=22.5 C

## Slide 5
Preliminary
17.2 p.e.
35.7 p.e.
60.8 p.e.
10.5
little bit too low

## Slide 6
Preliminary
Between the two measurements the fibers were removed, the sensor was cleaned from the optical grease and put back again
60.8 p.e.
61.1 p.e.

## Slide 7: Preparation of the GELINA test

## Slide 8: The radiation issue: Expected Neutron Rates
Expected maximum rates ~ 500 n/cm2 ~1.5x1010 n/cm2 year  10 years x safety factor 5: ~ 7.5 x 1011 n/cm2 in 10 years
This with the present simulations (no shielding between magnet and IFR-L0), but also with 3 cm tungsten shielding (no 4.5cm).
With the new production, rates might change
New results should be presented at this meeting
Expected rates on L0 barrel
Energy distribution
1eV
1MeV

## Slide 9: The radiation issue: Irradiation test at Gelina
The facility: 100 MeV linac, electrons on Uranium target + moderator to obtain a white spectra neutron beam
The main goals are:
  Measure the Dark current/noise with dose
  Monitor the gain and the signal waveform with dose (illuminating the devices with a LED)
  Estimate the neutron flux reduction power of various amounts of Boron loaded Polyethylene (2 ,5, 10 cm)
  Integrate a total dose of ~ 7.5 x 1011 n/cm2 (approximately 10 years x safety factor 5)

## Slide 10: The setup
Neutron beam
(≈ 4cm diameter)
Dark Currents
Dark Rates
Waveform
Temperatures
Bias Voltages
Target Hall
Neutron flux
monitor #1
Neutron flux
monitor #2
25 m downstream
Boron loaded Polyethylene (take data with and without it )
BF3
6Li
Space for 24 Devices (3 on
each of the 8 wedges)
LEDs in the irradiated area and outside (need light guide)
SiPMs

## Slide 11: Conclusions
R&D studies on 10cm scintillator bar are ongoing
Very preliminary results shows that with 3 fibers we collect ≈ the same amount of light of 5cm bars with 1 fiber (≈ 60 p.e.)
It Will be interesting to see the amount of light collected with 1,2,3 fibers on a 5cm scintillator bar
If with one fiber we are collecting already almost all the produced light we shouldn’t see a big increase
If this is the case maybe it’s not worth to put 3 fibers…. on 5 cm bars…
Preparation of the GELINA test is ongoing