## Slide 1: Measurements and cross checks on OPERA timing equipments
G. Sirri
INFN Bologna
On behalf of the OPERA Collaboration
LNGS, 28 March 2012

## Slide 2: Contents
NB Details about what is discussed here can be found in the “Report of the OPERA Collaboration to the Scientific Committees”, 12 March 2012
OPERA Timing
Time delay from the external LNGS Lab to the OPERA Master Clock
Time Delay from the OPERA Master Clock to the TT sensor
OPERA Master Clock Oscillator Frequency
Conclusions

## Slide 3: OPERA Event Time Stamping
FPGA
FPGA
10 MHz
O/E
ESAT GPS2
OPERA Master Clock
GPS Antenna
T10
Ts
ROC
PMT
TRIGGER
DAQ RESET
Every 0.6 s
MC PPmS
TMC
N*10 ns
M*0.6 s
20 MHz
HERTZ
Link to CERN
100 MHz
DAQ Time Stamp
121 …
122… 123…
ESAT PPmS
8 km
External Lab
Underground
Triggers from Target Tracker : time-stamped by an FPGA (100MHz)
DAQ cycle : 0.6 s ; Reset and Clock provided by OPERA Master Clock (hall C)
FPGA increments two counters:
M: coarse counter incremented by the DAQ RESET .
N : fine counter which increments every 10 ns
When a trigger is issued  the FPGA assigns to it the reading of the two counters.
DAQ Time-Stamps delayed by the time to transfer the GPS information to the FPGA (red path).
Master Clock oscillator: 10 MHz VECTRON OC-050 (stability 10-12/s)
5.9 ns

## Slide 4: Time delay from the external Lab to the OPERA Master Clock
Patch Panel
8 km
ESAT GPS2
LASER TX
ESAT
PPmS
Patch Panel
Splitter
HERTZ
Fiber 23
LNGS External Lab
LNGS Underground
OPERA
Master Clock
5 m Fiber
MC PPmS
5.9 ns
To FEBs
GPS Antenna
LVD
BOREXINO
O/E
O/E
tB
tA
ICARUS
Time delay tA between 2 reference points:
HERTZ reference of the GPS2 ESAT UTC time.
MC PPmS generated by the OPERA Master Clock, synchronous with the ESAT PPmS
The “two-way” technique
measure : tA – tB and tA + tB
see the “Report”

## Slide 5: Time delay from the external LNGS Lab to the OPERA Master Clock
6-8 Dec 2011 ( 41074.6 ± 0.5 ) ns r.m.s. ~ 6 ns
(*) The value quoted in ArXiv:v2 is 40996 and is obtained by subtracting 5.9 ns
20 Jul 2006 ( 41001.4 ± 0.3 ) ns r.m.s. ~ 3 ns
13 Dec 2011 ( 41000.0 ± 0.3 ) ns r.m.s. ~ 3 ns
(73.2 ± 0.6 stat ) ns larger than 2006
(*)
confirmed in 2007
consistent with 2006
Notes: Aimed to perform a cross-check and to look for second order effects (e.g. , temperature, earthquake,…)
Controversial puzzling results …
Further investigation involving everything were planned.
The 3 ns (6ns) jitter due to the Opera MC latching of the GPS signal. No proof of additional “eigenstates”

## Slide 6: Study on the fiber optical connectors
The delay changes by acting on both ends of a 5 m fiber
up to >100 ns according to the connection
FC
FC2
Any change of the input light signal intensity affects the time delay
Measurements of the time delay at an early stage of the amplification chain after the photodiode  next slide
Test Point
Photodiode
ETX100
Pre-Amplification Amplification
Reference
Voltage
Comparator 5V
Optical
GPS signal
Numerical
GPS signal
3.3V Buffer
OPERA Master Clock
5 m Fiber
MC PPmS
To FEBs
From External Lab
ESAT PPmS

## Slide 7: Delay measured at an early stage of the amplifier circuit
The effect is related to the charging up of the photodiode capacitance.
The Reset signal sent to the sensors is delayed by the same amount as the MC PPmS.
 Underestimation of the neutrinos ToF
FC connector correctly plugged and screwed
FC connector plugged in a "70 ns delay position"
Yellow : HERTZ
Blue : Amplifier out
Purple : MC PPmS
200 ns
1 V
Notes: The rise time and plateau of the amplifier output signal depend on the input light intensity

## Slide 8
Pictures of the Rack hosting the OPERA Master Clock
(pictures not aimed at checking cable connection)
FC fiber connector incorrectly plugged already by October 2011.
No earlier pictures available.
Since 13 Dec 2011 connector properly plugged
13 October 2011
6 December 2011
14 December 2011
8
Notes: it was not possible to determine for how long that faulty connection had been present.

## Slide 9: Time Delay from the OPERA Master Clock to the TT sensor
OPERA Master Clock
5.9 ns
From External Lab
ESAT PPmS
T10
FPGA
DAQ RESET
Every 0.6 s
Pulse Generator
Adjustable Delay Unit
Scope
7.15 ns
ROC
PMT
300.7 ns
A
(A) PPmS
RESET (T10)
Oscilloscope measurements
Delay between RESET (at T10) and Master Clock reference point: Δt(scope) + Δtlemo+ ΔtAB- Δtprobe= 4260.6 ns
Compatible with “official” 4262.9 ns
B
REFERENCE POINT
Δt(scope) = 3961 ns
MC PPmS
FPGA
DAQ Time Stamp
Reset/Clock fanout
121 …
122… 123…

## Slide 10: Time Delay from the OPERA Master Clock to the TT sensor
OPERA Master Clock
ESAT PPmS
T10
Reset/Clock fanout
FPGA
DAQ Time Stamp
Ts
Pulse Generator
Adjustable Delay Unit
ROC
PMT
TRIGGER
300.7 ns + 4500 ns
B
Charge injection using a capacitor
Scope
Δt(scope) = 582.8 ns
5.9 ns
A
REFERENCE POINT
+4500 ns
Charge injection to the PMT socket using a capacitor triggered by the MC PPmS.
1 Trigger / 1ms
Oscilloscope Test delay between RESET and 1st Trigger : 582.8 ns Compatible with the “official numbers”(*): 5.9+300.7+4500 + 37-4262.7 = 580.9 ns
DAQ Test
Expected Time-Stamps from “official numbers”(*) : 5.9+300.7+4500 + 37-4262.7-24.5 = 556.4 ns
(*) 37 ns is the “official” delay between the PMT connector and point Ts, 24.5 is the FPGA latency
DAQ RESET
Every 0.6 s
MC PPmS
121 …
122… 123…
From External Lab
Notes: A 2.3 pF capacitor was added in order to inject at the PMT connector level a signal equivalent to 100 p.e.

## Slide 11
1st triggers after DAQ cycle start
all triggers
Distribution of all triggers is broader than that of 1st Trigger
1st Triggers : Mean value : 560.1 ns compatible with “official” 556.4 ns
1 Trigger / 1ms DAQ cycle: 0.6 s
Time distribution of triggers as recorded by the DAQ (keeping only the ns digits of the DAQ time)

## Slide 12
time delay between 2 triggers: 1ms + 0.123 ns
at the end of the DAQ cycle ⇒ 74 ns more
⇒ time drift between the OPERA MC and ESAT GPS2 oscillators
Time of the DAQ cycle (ms)
Time delay from 1st trigger of the DAQ cycle
(keeping only the ns digits of the DAQ time)
Delay (ns)
1 Trigger / 1ms DAQ cycle: 0.6 s

## Slide 13: OPERA Master Clock: Frequency Measurements
Timing sources:
Cs Frequency Standard (10 MHz)
Rb Frequency Standard (5 MHz)
ESAT GPS Receiver (10 MHz, generated by the Rb with adjustments based on GPS data)
OPERA Master Clock Frequency (Vectron OC-050 , 10 MHz)
Results:
Rb+ESAT frequency consistent with Cs within specifications (<10-9).
OPERA Master Clock Oscillator frequency > 0.124 ppm wrt nominal value.
Overestimation of OPERA event time-stamps: depending on the event time within the 0.6 s DAQ cycle : (124.08 +/- 0.08) ns/s.
 overestimation of the neutrinos ToF CNGS neutrinos not randomly distributed in the DAQ cycle  See M. Sioli’s talk
Notes: CNGS Standard run : ~200 ms (mean) CNGS bunched beam : ~97 ms

## Slide 14: Conclusions
OPERA timing measurements at LNGS from Dec-2011 to Mar 2012
Found two unknown instrumental effects :
OPERA Master Clock card : time delay dependence from the input light signal  n-ToF underestimation. calibration close to the BB run likely indicating ~70 ns
MC oscillator frequency: > 0.124 ppm wrt nominal value  n-ToF overestimation depending on event time within the 0.6 s DAQ cycle : (124.08 +/- 0.08) ns/s .
Notes: MIGLIORARE IL PHRASING

## Slide 15: For the measurements reported here we are indebted to
LNGS Computing Service
LNGS Electronic Service
Electronic Services of IPHC Strasbourg, IPNL Lyon, INFN Bologna

## Slide 16: SPARES

## Slide 17: Statement from OPERA Collab.
The OPERA Collaboration, by continuing its campaign of verifications on the neutrino velocity measurement, has identified two issues that could significantly affect the reported result. The first one is linked to the oscillator used to produce the events time-stamps in between the GPS synchronizations. The second point is related to the connection of the optical fiber bringing the external GPS signal to the OPERA master clock.
These two issues can modify the neutrino time of flight in opposite directions. While continuing our investigations, in order to unambiguously quantify the effect on the observed result, the Collaboration is looking forward to performing a new measurement of the neutrino velocity as soon as a new bunched beam will be available in 2012. An extensive report on the above mentioned verifications and results will be shortly made available to the scientific committees and agencies.
23 February 2012