## Slide 1: ST + IT @ CNAO2025
Giacomo Traini, Giorgio Splendiani , Christian Fink, Riccardo Ridolfi,Marco Toppi, Giacomo Ubaldi, Eleuterio Spiriti

## Slide 2: ST amplitude

## Slide 3: Time of flight resolution
Next future: check for possible noise to cancel with a FFT filter

## Slide 4: IT setup summary
Unsuccessful replacement of datalink 111: sensors read-out from that board seems to be broken (to be fully investigated after the data taking @ LNF). The IT has been operated using 27/32 sensors.
The additional De10 board used to set the delay on the trigger line has some trouble in distributing the busy signal (wrong connection? firmware bug? ). The delay has been sent using a different system as done @ LNF
Thresholds: due to lack of time, to ensure a smooth data acquisition during the 1st night the IT has been precautionary operated @ 12σ. The 2nd night thresholds has been optimized with a noise scan
Delay
(NIM)
V1495
(VME)
Trig(NIM)
DAQ
IT
Trig delayed (NIM)
busy
clk
r_clk
Trig delayed (LVDS)

## Slide 5: IT setup summary
p 200 MeV, thr=12σ, delay=360 μs
p 150 MeV, thr=12σ, delay=360 μs
p 70 MeV, thr=12σ, delay=360 μs
C 200 MeV/u, thr=12σ, delay=360 μs, no TGT, B on
C 200 MeV/u, thr=12σ, delay=360 μs, no TGT, B on
C 200 MeV/u, thr=12σ, delay=440 μs, no TGT, B on
C 200 MeV/u, thr=12σ, delay=440 μs, no TGT, B off
C 200 MeV/u, thr=12σ, delay=440 μs, TGT, B off
C 200 MeV/u, thr=custom, delay=440 μs, TGT, B off
C 400 MeV/u, thr=custom, delay=440 μs, TGT, B off
C 200 MeV/u, thr=custom, delay=440 μs, TGT, B on
Delay
(NIM)
V1495
(VME)
Trig(NIM)
DAQ
IT
Trig delayed (NIM)
busy
clk
r_clk
Trig delayed (LVDS)

## Slide 6: Noise and thresholds (I)
Automatic pedestal scan performed @ CNAO between th 1st and 2nd night.
x 128
No significant differences wrt LNF in the pedestal mean values, while a slightly smaller sigmas have been observed (different temperature?)
CNAO LNF

## Slide 7: Noise and thresholds (II)
We performed a noise scan with random trigger, stdalone acquisition to set thresholds (2nd night)
x 128
We set 128 thresholds (32 sensors x 4 sectors) imposing a noise of 0.5 pixel/sector/event
Distribution of set thresholds @ CNAO
“custom thresholds”

## Slide 8: Delay effectiveness: IT frame distribution
run 8010, C @ 200MeV/u, B on, delay=360 μs
run 8016, C @ 200MeV/u, B on, delay=440 μs
Final choice

## Slide 9: IT vs VT frame distribution
run 8016, C @ 200MeV/u, B on,
VT
run 8016, C @ 200MeV/u, B on, IT, delay=440 μs
Expected difference according to the measurement perfomed @ LNF, but still a not understood behaviour

## Slide 10: Delay effectiveness: efficiency on C ions
VT
IT
Selection: 1BMtrack, 1Trk in VT with BM match (norm)
We project the VT track on the IT and we search and count the clusters with dist<1mm
preliminary
custom thresholds
Run 8039, C200, TGT, B off
Due to inactive sensors?

## Slide 11: Single sensor efficiency
We suspect some small geometry issues (+ alignment accuracy)
The cluster may be in the adjacent sensor if the sensor assignment is wrong

## Slide 12: Single sensor efficiency
The cluster may be in the adjacent sensor if the sensor assignment is wrong
When we exclude the events where the track impinging the sensor boundaries thins are better

## Slide 13: Beam map on IT
broken sensors
Keep in mind that sensors have different depths! In principle we should have at least one hit per particle…
with the B field we move towards the missing sensor…
Run 8039, C200, TGT, B off

## Slide 14: IT vs VT correlation
Run 8022, alignment run, C200, no TGT, B off
In x correlation seem worse than y…

## Slide 15: IT vs VT correlation
Run 8016, C200, no TGT, B on
… may be due to residual B? check in the MC
Particles
deflection
Spread

## Slide 16: IT response
run 8016, C200, Thr = 12σ
run 8038, C200, Thr = custom

## Slide 17: Conclusions
We do not know if, how (and when) the lost sensors could be recovered. Fortunately, the other sensors seems to perform as expected
Wrt the schedule agreed before the data-taking we did not perform the delay scan due to the De10 problem.
However, according to the shown results (even really preliminary), I would privilege a smooth data-acquisition to maximise data collection keeping the setup as it is, postponing all the studies after the data-taking (a test with electrons is scheduled @ BTF in Feb 2026)

## Slide 18: Delay effectiveness: efficiency on C ions
VT
IT
Selection: 1BMtrack, 1Trk in VT with BM match (norm)
We project the VT track on the IT and we search and count the clusters with dist<1mm
preliminary
VT tracks not crossing sensor14
Run 8039, C200, TGT, B off
custom thresholds