## Slide 1
motivation: laser ion acceleration at Garching for
bio-medical applications (CALA project)
 requires online monitoring and dosimetry
 measure prompt g rays from nuclear reactions
development of a Compton Camera with electron
tracking capability:
 design specifications via simulations
 prototype laboratory setup and characterization tests
Peter G. Thirolf, LMU Munich
Development of a Compton Camera
for online range monitoring
of laser-accelerated proton beams
via prompt-g detection

## Slide 2
Range Monitoring / Online-Dosimetry
within phantom/patient
Detection of prompt g-rays
phantom/
sample
Laser-driven Proton Acceleration:
Studies of (pre-) clinical applicability
aperture
mini-
quadrupoles
D+ D-
aperture or
spectrometer plate
DLC foil
ATLAS laser
(60 J, 20 fs, 3 PW, 1-5 Hz)
Kapton foil
I.

## Slide 3
Prompt gamma emission from
proton beam on biomedical sample
irradiation of water phantom
with 100 MeV protons
(FLUKA simulation):
ca. 0.042 g / primary proton
ca. 108 protons (16 pC) per laser pulse
 ca. 4.2.106 g/laser pulse
4.4 MeV 12C
5.2 MeV 15N, 15O
6.1 MeV 16O
Counts / GeV/ incident p
50 cm
Ø 10 cm
p
H2O
0 2 4 6 8 10 12 14 16 18
Energy [MeV]
key issue:
localization of Bragg peak
within sample
approach: prompt g emission
from nuclear reactions
advantage of hadron therapy:
Bragg peak
Notes: 3
3

## Slide 4
Principle of a Compton Camera
 tracking:
 and electron tracking:
exploit kinematics of Compton scattering:
measure DEg,E’g and directions eg of scatter + absorption process:
 g source origin can be restricted to a cone surface
additional electron tracking of direction ee and energy Ee :
 origin of g can be restricted to an arc
 reconstruction of incompletely absorbed g events possible


e , ee

Scatterer/
Tracker
Absorber
Compton
cone
Compton
arc
E‘e
 ≥ 1 MeV
E‘e

D
Notes: 4
4

## Slide 5
g point source:
isotropic,
 and electron tracking:
Scatterer/Tracker:
DSSSD
Absorber:
LaBr3
(+ PMT)
10 mm each
50 mm
85 mm
z
x
y
E: 0.5 - 6 MeV
e-

Compton Camera Design
Scatterer/Tracker:
double-sided silicon strip
detectors (DSSSD)
active area 50 x 50 mm2
thickness :
300 mm or 500 mm
128 strips on each side
pitch size 390 mm
Absorber:
LaBr3 scintillator
 active volume:
50 x 50 x 30 mm3
multi-anode PMT:
Hamamatsu 9500
256 Pixel (3 x 3 mm2);
presently combined to
64 Pixel (6 x 6 mm2)

## Slide 6
simulation for tracker/absorber specifications:
Compton Camera Design Simulations
simulation/reconstruction tool: MEGAlib
 Monte Carlo simulation (based on Geant4)
 event reconstruction
 image reconstruction (LM-ML-EM algorithm)
6x6 mm2  3x3 mm2 pixel:
spatial resolution improves by ≥50 %
nDSSSD=6
full absorption
g tracking ; source - tracker: 50 mm
d=500 mm + electron tracking:
 improved efficiency
5 MeV
30 iterations,
50 events
64 pixel (6x6mm2)
256 pixel (3x3mm2)
300mm Si
500mm Si
300mm Si
500mm Si
g + e tracking, 500 mm
g + e tracking, 300 mm
g tracking, 500 mm
g tracking, 300 mm
g
g +e

## Slide 7
Reconstruction Properties
- e ≈ 10-3 – 10-5 (@ 1- 5 MeV for optimum resolution)
- angular resolution ≈ 2o – 2.5o (@ 2-6 MeV)
reconstruction efficiency:
6x 500 mm, 3x3 mm2 pixel
angular resolution:
source – tracker: 50 mm
Notes: 7
7

## Slide 8
LaBr3
PMT
64 channels
LaBr3 crystal (Saint Gobain):
 active area 50 x 50 x 30 mm3.
PMT: Hamamatsu H9500.
 256 fold multi-anode .
Compton Camera: Absorber
signal processing:
start version: 64 pixel (6x6 mm2)
upgrade: 256 pixel (3x3 mm2)
 fast amplifier + CFD
(Mesytec MCFD-16, 16 ch.)
 charge-sensitive digital
converter (Mesytec, 32 ch. QDC)

## Slide 9
light distribution: 137Cs (662 keV), collimated source (Ø 1 mm)
: source position
(background corrected, gain matching: PMT uniformity, electronics)
Spatial Resolution
pixel, X
pixel,Y
Charge [a.u.]
r [mm]
pixel, x
pixel, X
pixel,Y
pixel, X
pixel,Y
Point Spread Function:
(FWHM of radial
charge distribution)
final goal:
- scan with 0.5 mm grid
- resolution ca. 1 mm

## Slide 10
Gassiplex (4x16 ch. ASIC):
 charge-sensitive preamp  shaper
 digital discriminator
 track & hold-stage
multiplexed ADC
Compton Camera Scatter Detectors
Scatterer/Tracker:
double-sided silicon strip
detectors (DSSSD)
active area 50 x 50 mm2
thickness : 500 mm
128 strips on each side
pitch size 390 mm
stack of 6 modules
Adapter
board
1
2
3
4
DSSSD
Gassi-plex (64 ch.)
Gassi-plex
Gassi-plex
Gassi-plex
AC
AC
Wafer
AC
A
C
2x 64 strips/
side
VME readout controller
6x 256 channels

## Slide 11
Conclusions and Outlook
Compton camera for online monitoring and dosimetry via prompt g ’s:
- system designed to allow for (g + electron) tracking
- spatial resolution: ~ 1.7 mm @ 5 MeV and 50 mm distance to source
- reconstruction efficiency (optimum resol.): ~ 10-5 @ 5 MeV
next steps: - laboratory commissioning of scatter/tracker part
- online tests of complete Compton camera (Ep=20-150 MeV)
prototype setup and characterization:
- absorber characterized:
LaBr3 with multi-anode PMT (6x6mm2 pixel, upgrade: 3x3mm2)
- scatterer designed, prototype just received:
6x DSSSD, (500 mm, 50x50 mm2, 2x128 ch.)

## Slide 12
Thanks to …
LMU Munich: C. Lang, S. Aldawood, D. Habs, M. Gross,
R. Lutter, H. van der Kolff, J. Bortfeld,
K. Parodi
Thank you for your attention !
TU Munich: L. Maier, R. Gernhäuser
SSL Berkeley: A. Zoglauer