## Slide 1
M. Bruzzi1,2*, M. Bucciolini2,3, G. A. P. Cirrone4, C. Civinini2, G. Cuttone4, D. Lo Presti5,6, S. Pallotta2,3, C. Pugliatti5,6, N. Randazzo5, F. Romano4,7, V. Sipala8,9, M. Scaringella1, C. Stancampiano5,6, C. Talamonti2,3, M. Tesi1, E. Vanzi10, M. Zani2,3
1Physics Department, University of Florence, Italy
2INFN - Florence Division, Italy 3Dipartimento di Scienze biomediche, sperimentali e cliniche, University of Florence, Florence, Italy 4INFN - Laboratori Nazionali del Sud, Catania, Italy, 5INFN - Catania Division, Italy 6Physics Department, University of Catania, Catania, Italy 7Centro Studi e Ricerche e Museo Storico della Fisica, Rome, Italy 8Chemistry and Pharmacy Department, University of Sassari, Sassari, Italy 9INFN Cagliari Division, Italy
10 SOD Fisica Medica, Azienda Ospedaliero-Universitaria Careggi, Firenze, Italy
* Spokesperson
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy
Proton Computed Tomography system: recent results and upgrade status

## Slide 2: Outline
Introduction: proton therapy and proton imaging
Proton imaging apparatus
  Tracker
    Single module architecture
    Detector
    VLSI front-end
  Calorimeter
    Crystal YAG:Ce
    Electronic readout
    Trigger system
First results with proton beam
Conclusions
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 3: Proton therapy
The proton therapy is a good clinical treatment for cancer as it permits to obtain a dose distribution extremely conform to the target volume.
The Bragg peak shape ensures that healthy tissues in front of and beyond the tumor are not damaged.
Through the weighted superposition of proton beams of different energies it is possible to deposit a homogenous dose in the target region using only a single proton beam direction.
(Spread Out Bragg Peak -SOBP).
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 4: proton Computed Tomography: why?
In proton therapy treatments is important to know:
Patient positioning:
Currently performed using X-rays radiography in previous phase
pCT allows better accuracy and
single phase positioning / treatments
Dose Calculation:
Currently performed using X-rays computed tomography
Problem: protons and photons have
different interaction with matter
pCT uses protons directly for dose calculation
The error intrinsic in this conversion (due to e,Z) dependency on atomic number and electron density) is the principal cause of proton range indetermination (3%, up to 10 mm in the head)‏
[Schneider U. (1994), Med Phys. 22, 353]

## Slide 5: Parameters of pCT
Problem: protons, because the multiple Coulomb scattering, don’t move in straight line.
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 6
THE “PRoton IMAging (PRIMA) PROJECT” pCR apparatus oriented to pCT
  PROGRAM
  Design and manufacture a high-performance prototype for proton radiography and tomography.
  Develop suitable imaging algorithms;
  Validate the pCR/pCT system with pre-clinical studies
  Conceive an upgraded prototype for pre.clinical studies
  - Hardware and data acquisition;
  - Reconstruction algorithms.
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 7: Concept
proton Computed Radiography and Tomography
Reveal the track of the single proton using a silicon telescope
Measure the residual energy of the proton using a calorimeter
Reconstruct the most likely path of the single proton
pCR for different projections with a rotating gantry
Single event information can be processed by reconstruction algorithms to produce tomographic images.
P1
P2
P3
P4
z
x
y
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 8: Architecture of the pCT apparatus
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 9: Tracker module architecture
To achieve a read-out rate of 1 MHz a fully parallel digital strip readout system has been developed
Eight 32-channel VLSI front-end chips acquire the detector signals and sends data in parallel to an FPGA (Xlinx Spartan-3AN) which performs zero suppression and moves data to a buffer memory (~5x105 events).
An Ethernet commercial module is use both for data transfer to the central acquisition PC and to control the tracker module DAQ parameters
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 10
Silicon strip detector
Manufactured by Hamamatsu Photonics
53x53 mm2
p+-on-n strips
256 ch, 200 m pitch
200 m thickness
(To reduce the multiple scattering
in the detector while keeping
a good sensitivity to protons)
Vfd = 65-75V
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy
Notes: 10
10

## Slide 11: VLSI front-end description
AMS 0.35u CMOS Technology
1.6 mm x 6 mm
32 channels
Power dissipation = 14,5 mW @ chan
Vcc = +3.3 V
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 12: Calorimeter
YAG:Ce properties | 
PHYSICAL PROPERTIES | 
Density [g/cm3] | 4.57
Hygroscopic | No
Chemical formula | Y3Al5O12
LUMINESCENCE PROPERTIES | 
Wavelength of max. emission [nm] | 550
Decay constant [ns] | 70
Photon yield at 300k [103 Ph/MeV] | 40-50
4 YAG:Ce scintillating crystals
Each crystal 30 x 30 mm2 x 100mm
4 Photodiode 18 mm x 18 mm
4 commercial front-end (Charge Sensitive Amplifier & shaper )
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 13: pCT system designed and manufactured by the PRIMA INFN CSN5 collaboration
4 x-y TRACKER MODULES
1 CALORIMETER
Entry and Exit
position and direction
Residual Energy
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 14
First tests: one x-y plane with YAG:Ce calorimeter (I)
INFN LNS p 60MeV
GOAL: to test the functionality of the architecture.
A x-y tracker plane has been coupled with the calorimeter. The test has been performed at LNS with 60MeV proton beam.
Collimator was used ( Ø 5mm ). Rate 1-50kHz
Counting map ( 1 crystal not connected)
Data projection
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 15
Map of events triggered by crystals II (red) and IV (black)
Correlation between tracker and calorimeter data
 collimator removed from the beam pipe
Data projection
Good correlation !
First tests: one x-y plane with YAG:Ce calorimeter (II)
INFN LNS p 60MeV
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 16
Calorimeter
Tracker: 4 x-y planes
Beam
Alignment test
Radiography of a non homogeneous phantom
radiographies at different angles for first tomographic reconstruction
First tests: Complete apparatus
INFN LNS p 60MeV
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 17
First tests: Radiography of a walnut
INFN LNS p 60MeV
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 18
One cluster in each plane
Cut on difference between “entrance“ and “exit” angles
Sigma=0.170 ± 0.09mm
FWHM = 0.400± 0.212mm
Residual energy (MeV)
P1 P2 P3 P4
x
y
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy
First radiography tests with complete apparatus
INFN LNS p 60MeV

## Slide 19
pCT prototype mounted on
180 MeV protons beam line
Radiography without filtering
Hole thickn. (cm)
1
1,5
2,3
3,3
4,5
6
8
10
Pmma cylindrical phantom with holes
14 cm total thickness
Not to scale
First radiography tests with complete apparatus
Svedberg Laboratory Uppsala p 180MeV
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 20
  PMMA phantom: diameter 2 cm, height 4 cm – holes f 4&6 mm, length 2 cm 36 projections (10 degrees steps)
2 cm
First Tomography tests with complete apparatus
INFN LNS p 60 MeV
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 21
Definition of the tomographic equation (Wang, Med.Phys. 37(8), 2010: 4138)
«projection»
Unknown stopping power distribution (at E0)
Evaluation of the “projection” term (through numerical integration starting from NIST tables and using the measured Eres)
Eres
Wang projection
Mean measured residual energy plotted at P3 plane
Corresponding Wang projections
Vanzi E. et al. The PRIMA collaboration: preliminary results in FBP reconstruction of pCT data, in press on Nucl. Instrum. Meth A.
Data Analysis: Tomographic equation
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 22
We defined a plane parallel to detector’s planes passing through the phantom axis. Plane is sampled in a 256 x 256 matrix, 200µm pixel size. For each event, associated projection bin is determined by the intersection of the line connecting P2 and P3 impact points with the plane.
P2
P3
Data Analysis: Rebinning

## Slide 23
Butterworth filter: order 2,
cut-off 20/128 of the Nyquist freq.
Good quality images without cuts: possibility of reducing acquisition times and dose in patient procedures
Tomographic reconstruction: first results
Vanzi E. et al. The PRIMA collaboration: preliminary results in FBP reconstruction of pCT data, in press on Nucl. Instrum. Meth A.
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 24
The edge of the phantom is fitted with an erf function on 20 slices in the homogeneous region.
The mean FWHM of the LSF over the 20 slices was evaluated and used to quantify resolution.
The derivative of the erf function is a Gaussian function that describes the Line Spread Function (LSF) of the imaging system.
FWHM=0.9±0.1mm
Noise=2.4%
Mara Bruzzi Riunione RDH 22 Aprile 2013 Roma
Resolution Evaluation

## Slide 25
Increase 4x of the active area
On-line data acquisition
Data rate up to 1 MHz
A new prototype based on the same technology, to meet pre-clinical demands now under development
Tracker:
Calorimeter:
4x1 silicon detectors
5x20 cm2 active area
7x2 YAG crystals
6x21 cm2 active area
Rectangular shape to perform tomography in slices
X1 FE
X2 FE
X3 FE
X4 FE
top
Scaringella M. et al. The PRIMA collaboration:, in press on Nucl. Instrum. Meth A.
Upgrade: the pre-clinical prototype
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy

## Slide 26
7x2 YAG:Ce scintillating crystals 30 x 30 mm2 x 100mm each. Photodiodes assembled . Now under test (RDH)
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy
Upgrade: the upgraded calorimeter

## Slide 27
A pCT apparatus has been designed and built by the PRIMA Collaboration INFN CSN5 based on:
  Silicon microstrip detectors 5x5 cm2 active area
  YAG:Ce cristal calorimeter
The prototype has been tested under proton beams with energies up to 180 MeV with promising results.
First tomographic images were reconstructed with FBP with encouraging results. Even using the information on two planes only and without cuts on events good quality images were obtained.
A new prototype with larger area (5x20 cm2) is under development for pre-clinical validation.
Future progress in WP3 RDH Project INFN CSN5.
Conclusions
M. Bruzzi,, International Nuclear Physics Conference INPC2013: 2-7 June 2013, Firenze, Italy