## Slide 1: Imaging devices for medicine and security
Andy Boston
ajboston@liv.ac.uk

## Slide 2: Outline of presentation
Focus on ionising radiation and gamma-rays
  What are the challenges?
  What detector technology can we consider?
  Select example projects & links to fundamental research
  The future prospects

## Slide 3: KNOWLEDGE EXCHANGE
Scientific Research
Applications
e.g. AGATA
Gamma-ray Tracking
Imaging
Medical
Security
Environment

## Slide 4: What are the challenges?
In Nuclear Medicine:
  Know the energy
  Want the location over a small field of view
  Need to cope with high count rates
  Multimodality applications (eg PET/CT)
  Image fusion

## Slide 5: What are the challenges?
In Nuclear Security :
  Don’t know the energy & a broad range
  Want the location over a large field of view
  Need to cope with wide range of count rates
  Image fusion

## Slide 6: What are the detector requirements?
Ideally would want:
  Good energy resolution (Good light yield/charge collection) < few %
  High efficiency (High Z)
  Position resolution
  Timing resolution
Detector materials:
  Semiconductors (Si, Ge, CdZnTe)
  Scintillators (LaBr3, CsI(Tl), NaI(Tl), BaFl, BGO, LYSO…)

## Slide 7: What are the detector requirements?
Need to know the location of the radiation:
  Use a mechanical collimator (Anger Camera)
  Use positron annihilation for LoRs
  Use other electronic collimation
Range of energies:
  Medical 141 keV – 511 keV
  Security 60 keV – 2 MeV
Operating environment:
  B-fields? Microphonics? High temperature?

## Slide 8: Medical Imaging
Focus on SPECT
How it works and the opportunities

## Slide 9: What is SPECT?
Functional imaging modality
8 mSv typical dose

## Slide 10: What SPECT Radionuclides?
141 keV
t1/2=65.94h
2.1105y
t1/2=6.01h
>99%
910-5%
stable

## Slide 11: Tomographic Imaging
The sinogram is what we aim to measure
  - Measure of intensity as a function of projection, θ and position, r
  - Often seen plotted as a 2d grey scale image
Underlying source distribution
“Shepp-Logan Phantom”
Measured result – “Sinogram”
(256 projections, 363 positions per projection)
Note : We measure from 0 to 180°
θ
r
p(r , θ)
x
y
f(x ,y)

## Slide 12: SPECT : Problems/Opportunities
Technical
Collimator Limits Spatial Resolution & Efficiency
Collimator is heavy and bulky
Energy of radioisotope limited to low energy
NaI:Tl Dominant for >40 Years...
MRI  Existing PMTs will not easily operate
Would like to be able to image a larger
fraction of events.
Common radionuclides: 99mTc, 123I, 131I

## Slide 13
ProSPECTus
Next generation Single Photon Emission Computed Tomography
Nuclear Physics Group, Dept of Physics, University of Liverpool, Nuclear Physics & Technology Groups, STFC Daresbury Laboratory, MARIARC & Royal Liverpool University NHS Trust

## Slide 14: ProSPECTus: What is new?
ProSPECTus is a Compton Imager
Radical change  No mechanical collimator
Utilising semiconductor sensors
Segmented technology and PSA and digital electronics (AGATA)
Image resolution 7-10mm  2-3mm
Efficiency factor ~10 larger
Simultaneous SPECT/MRI

## Slide 15: What’s different?
Conventional SPECT
Compton camera
Gamma rays detected by a gamma camera
Inefficient detection method
Incompatible with MRI
2D information
Gamma rays detected by a Compton camera
Positions and energies of interactions used to locate the source
3D information.
Source
E0
Factors that limit the performance of a Compton Imager:
Energy resolution, Detector position resolution, Doppler Broadening

## Slide 16
Research : Compton Imaging
Compton Cones of Response projected into image space

## Slide 17
Research : Compton Imaging
Compton Cones of Response projected into image space

## Slide 18
Research : Compton Imaging
Compton Cones of Response projected into image space

## Slide 19
Research : Compton Imaging
Compton Cones of Response projected into image space

## Slide 20
Research : Compton Imaging
Compton Cones of Response projected into image space

## Slide 21: System Configuration
1cm
2cm
2cm
5cm
141keV
Si(Li)
Ge
Total Coincident ~3.49%
SPECT ~ 0.025% (typical value)
Factor of ~140
Event Type | %
Single / Single | 2.23
Single / Multiple | 0.33
Multiple / Single | 0.61
Multiple / Multiple | 0.04
Not absorbed | 0.28
GEANT4 simulations L. Harkness
Ortec & SemiKon

## Slide 22: Ingredients of g-Tracking
Pulse Shape Analysis to decompose recorded waves
Highly segmented
HPGe detectors
·
·
·
·
Identified
interaction points
(x,y,z,E,t)i
Reconstruction of tracks
e.g. by evaluation of permutations
of interaction points
Digital electronics
to record and process segment signals
g
1
2
3
4
reconstructed g-rays
Notes: What has to be done to perform tracking
What are the ingredients
Detectors highly segmented, encapsulated, cryostats
Electronics Digital
PSA to extract Energy Time and Position
Algorithm development particularly for position
Area where effort is required Thorsten Kroell. See him.
Tracking algorithms, reconstruction of the events for full array
Obtain full energy
Many technical development in all areas.

## Slide 23: Utilise characterised sensors Am-241 AC x-y surface intensity distribution
Uniformity of response for planar germanium crystals.
AC01 AC12
DC12
DC1

## Slide 24: Utilise unique pulse shape response
AC signals
DC signals
DC signals
AC signals
“superpulse” pulse shapes for 137Cs (662 keV) events versus depth

## Slide 25: Image Reconstruction Algorithms
Sensors have excellent energy & position information.
Uniformity of sensor response
Optimise existing:
  Analytical
  Iterative
  Stochastic
Requirement for GPU acceleration

## Slide 26: Compton Imaging
~7º Angular Resolution FWHM, central position
2cm source separation
152Eu E = 1408 keV
22Na E = 1274 keV
152Eu
Multi-nuclide imaging
No PSA (5x5x20)
Cone back projection

## Slide 27: MRI compatibility & Status
Test existing gamma-ray detector in an MRI scanner
Does the detector cause distortions in the MRI image? No
Does the MRI system degrade the detector performance? In certain positions (which can be minimised)
Encouraging results!

## Slide 28: What are the next steps?
Immediate priorities
We have an integrated Compton Gamma camera optimised for <500keV
Demonstrate sensitivity with phantoms
Trials including clinical evaluation
For the future:
Consider electron tracking Si scatterer
Possible use of large CZT analyser (requires large wafer material with 1cm thickness)

## Slide 29: Security Imaging
SNMs and other threats
Coded aperture systems (low energy)
Focus on wide FOV and variety of stand off distances
Compton cameras

## Slide 30: Location and Identification…
The ability to locate and identify radioactive material with high precision
Quantification of waste into low/intermediate/high brackets
Wide range of activities from ~37kBq -> MBq
There are many open challenges and opportunities
Courtesy K. Vetter LBL (work @ LLNL)

## Slide 31: Si(Li) + Ge Cryogenic solutions
Mechanically cooled
Battery powered
Work in collaboration with Canberra

## Slide 32
CZT Room temperature: PorGamRayS
A portable gamma-ray spectrometer with Compton imaging capability (60keV – 2MeV)
Gamma-ray spectroscopy/imaging with CZT detectors. Pulse Shape Analysis to refine spatial resolution and correct charge collection issues

## Slide 33: 137Cs example image with CBP
370 kBq @ 5cm standoff 19mm FWHM

## Slide 34: 133Ba example image with CBP
340 kBq @ 5cm standoff
19mm FWHM

## Slide 35: Compton Imaging
Typical measurements:
10μCi 152Eu
6 cm from SPET 1
Source rotated
Zero degrees in 15º steps up to 60º
Detector separation
3 – 11cm in 2cm steps
Gates set on energies
2 sources 152Eu and 22Na
at different x and y
Use of the SmartPET detectors in
Compton Camera configuration

## Slide 36
FWHM ~ 8mm
6 cm source to crystal
30 mm crystal to crystal
E = 1408 keV, 30 keV gate
Compton Camera measurements (Ge/Ge)
No PSA (5x5x20)
Iterative reconstruction

## Slide 37: Stereoscopic Optical Image Fusion
1.5m standoff
A Compton Camera provides 3D source location
Utilise a 3D optical imager
Bubblebee 3 camera head

## Slide 38: The next steps
A number of preclinical and field trails are under way.
Utilise e-tracking to reduce the uncertainty in the final image
Improved image reconstruction
Simplified electrode structure
Further use of room temperature sensors

## Slide 39: Imaging devices for medicine and security
Andy Boston
ajboston@liv.ac.uk