## Slide 1: Nuclear Structure with Gamma-ray Tracking Arrays
Dino Bazzacco
INFN Padova

## Slide 2: Challenges in Nuclear Structure
Neutron-rich heavy nuclei (N/Z → 2)
Large neutron skins (rn-rp→ 1fm)
New coherent excitation modes
Shell quenching
132+xSn
Nuclei at the neutron drip line (Z→25)
Very large proton-neutron asymmetries
Resonant excitation modes
Neutron Decay
Nuclear shapes
Exotic shapes and isomers
Coexistence and transitions
Shell structure in nuclei
Structure of doubly magic nuclei
Changes in the (effective) interactions
48Ni
100Sn
78Ni
Proton drip line and N=Z nuclei
Spectroscopy beyond the drip line
Proton-neutron pairing
Isospin symmetry
Transfermium nuclei
Shape coexistence

## Slide 3: Experimental Conditions and Challenges at Radioactive Beam Facilities
Low intensity for the nuclei of interest
High background levels
Large Doppler broadening
High counting rates
High g-ray multiplicities
High efficiency
High sensitivity
High throughput
Ancillary detectors
Need of advanced general-purpose instrumentation
Beyond the capability of the best Compton-suppressed Detector Arrays

## Slide 4: Effective Energy Resolution
Intrinsic
Opening
Recoil
Eg 1 MeV
DElab 2 keV
b(%) 5±0.01 20±0.005
DQ(deg) 8 2
DEg/Eg (%)

## Slide 5: Motivation of g-ray tracking
eph ~ 10%
Ndet ~ 100
too many detectors needed to avoid summing effects
opening angle still too big for very high recoil velocity
Smarter use of Ge detectors
segmented detectors
digital electronics
timestamping of events
analysis of pulse shapes
tracking of g-rays
Compton Suppressed
Ge Sphere
Tracking Array
eph ~ 50%
Ndet ~ 1000
W ~ 40% q ~ 8º
50% of solid angle taken by the AC shields
large opening angle  poor energy resolution at high recoil velocity
eph ~ 50%
Ndet ~ 100
W ~ 80% q ~ 3º
W ~ 80% q ~ 10º
Pulse Shape Analysis qeff ~ 1º Gamma-ray Tracking Neff ~ 10000
from Calorimetric to Position Sensitive operation mode

## Slide 6: Position-sensitive Operation Mode and Gamma-ray Tracking
Pulse Shape Analysis of the recorded waves
Highly segmented
HPGe detectors
Identified
interaction points
(x,y,z,E,t)i
Reconstruction of g-rays from the hits
Synchronized digital electronics to digitize (14 bit, 100 MS/s) and process the 37 signals generated by crystals
Analysis of gammas
and correlation with other detectors
·
·
·
·
·
·
·
·
g
Readout Raw Data (10 kB/evt/crystal)
Event building of time-stamped hits and ancillaries
Global level
Local level
HARDWARE
SOFTWARE

## Slide 7: Reconstruction of gammas rays
Mg = 30
High-multiplicity simulated event
Efficiency depends on position resolution
Position resolution (FWHM, mm)
100 keV
10 MeV
g-ray energy

## Slide 8: Two Suitable Geodesic Configurations
120 crystals  GRETA
180 crystals  AGATA
Configuration | “small” | “big”
# of hexagonal crystals | 120 | 180
# of crystal shapes | 2 | 3
# of clusters | 30 | 60
Covered solid angle (%) | 78.0 | 78.4
Germanium weight (kg) | 230 | 370
Centre to crystal-face (cm) | 18.5 | 23.5
Signal channels | 4440 | 6660
Efficiency at Mg = 1 (%) | 36.4 | 38.8
Efficiency at Mg = 30 (%) | 22.1 | 25.1
P/T at Mg = 1 (%) | 51.8 | 53.2
P/T at Mg = 30 (%) | 43.4 | 46.1
Monte Carlo and simulations by Enrico Farnea
Notes: Enrico a precious colleague and fantastic person departed/deceased on April 14

## Slide 9: Status after ~10 years of development pursued by the AGATA and GRETA collaborations
Germanium detectors
Electronics and DAQ
  Fully digital systems with common clock and time-stamping
  Real time trigger (timestamp based in AGATA)
  Coupling to EDAQ of Auxiliary detectors based on timestamp
Pulse Shape Analysis
Gamma-ray Tracking
Problems encountered
  Cross Talk  solved
  High counting rates  solved by digital electronics
  Neutron damage  solved by PSA
Early implementations: AGATA Demonstrator and GRETINA
  Performance
  Evolution

## Slide 10: AGATA detectors
Volume ~370 cc Weight ~2 kg
(shapes are volume-equalized to 1%)
AGATA capsules Manufactured by Canberra France
AGATA Asymmetric Triple Cryostat
Manufactured by CTT
80 mm
90 mm
6x6 segmented cathode
Energy resolution
Core: 2.35 keV
Segments: 2.10 keV
(FWHM @ 1332 keV)
A. Wiens et al. NIM A 618 (2010) 223
D. Lersch et al. NIM A 640(2011) 133
Cold FET for all signals

## Slide 11: GRETINA Detectors (Canberra/France)
B-type
A-type
36 segments/crystal
4 crystals/module
148 signals /module
Cores: Cold FETs
Segments: Warm FETs

## Slide 12: Scanning of Detectors
662keV
374keV
288keV
<010>
<110>
T30
T60
T90
Region of Interest
Ge Energy
NaI Energy
374 keV
288 keV
U. Liverpool
920 MBq 137Cs source
1 mm diameter collimator

## Slide 13: Pulse Shape Analysis concept
B4
B5
B3
C4
C5
C3
CORE
A4
A5
A3
791 keV deposited in segment B4
measured

## Slide 14: Pulse Shape Analysis concept
B4
B5
B3
C4
C5
C3
CORE
A4
A5
A3
C4
D4
E4
F4
A4
B4
x
y
z = 46 mm
(10,10,46)
measured calculated
791 keV deposited in segment B4

## Slide 15: Pulse Shape Analysis concept
B4
B5
B3
C4
C5
C3
CORE
A4
A5
A3
C4
D4
E4
F4
A4
B4
x
y
z = 46 mm
(10,15,46)
measured calculated
791 keV deposited in segment B4

## Slide 16: Pulse Shape Analysis concept
B4
B5
B3
C4
C5
C3
CORE
A4
A5
A3
C4
D4
E4
F4
A4
B4
x
y
z = 46 mm
(10,20,46)
measured calculated
791 keV deposited in segment B4

## Slide 17: Pulse Shape Analysis concept
B4
B5
B3
C4
C5
C3
CORE
A4
A5
A3
C4
D4
E4
F4
A4
B4
x
y
z = 46 mm
(10,25,46)
measured calculated
791 keV deposited in segment B4

## Slide 18: Pulse Shape Analysis concept
B4
B5
B3
C4
C5
C3
CORE
A4
A5
A3
C4
D4
E4
F4
A4
B4
x
y
z = 46 mm
(10,30,46)
measured calculated
791 keV deposited in segment B4

## Slide 19: Pulse Shape Analysis concept
Result of Grid Search Algorithm
B4
B5
B3
C4
C5
C3
CORE
A4
A5
A3
C4
D4
E4
F4
A4
B4
x
y
z = 46 mm
(10,25,46)
measured calculated
791 keV deposited in segment B4

## Slide 20: Pulse Shape Analysis Algorithms
Computation Time/event/detector
ms
s
hr
Position resolution (mm FWHM)
2
0
4
6
8
Singular Value Decomposition
Genetic algorithm
Wavelet method
Full Grid Search
Least square methods
Artificial Intelligence (PSO, SA, ANN, ...)
Adaptive Grid Search
now

## Slide 21: Examples of signal decomposition
1 A 6 1 B 6 1 C 6 1 D 6 1 E 6 1 F 6 CC
Eg = 1172 keV net-charge in A1
Eg = 1332 keV net-charge in C4, E1, E3
x10
x10
Tomography of interactions in the crystal: non uniformities due to PSA

## Slide 22: Position resolution (GRETINA)
Decomposition program, ORNL, LBNL
coincidence
s = 1.9 mm (average of 18 crystals)
sx = 1.2 mm, sy = 0.9 mm
S. Paschalis et al, NIMA 709 (2013) 44–55

## Slide 23: Position resolution (AGATA)
P.-A.Söderström, F.Recchia et al, NIMA 638 (2011) 96
12C(30Si,np)40K at 64 MeV v/c = 4.8 %
Two target positions: 5.5 and 23 cm (-16 cm and +1.5 cm re center of array) to remove systematic errors
Spectrum at short distance and used peaks
Position resolution of first hit (fwhm)
s ~ 2 mm at 1 MeV
Eg (keV)
Ep1 (keV)

## Slide 24: Cross talk correction: Results

## Slide 25: Correction of Radiation Damage
B.Bruyneel et al, Eur. Phys. J. A 49 (2013) 61
Line shape of the segments of an AGATA detector at the end of the experimental campaign at Legnaro (red) The correction based on a charge trapping model that uses the positions of the hits provided by the PSA restores the a Gaussian line shape (blue)

## Slide 26: First Gamma Tracking Arrays
AGATA Demonstrator
GRETINA
15 crystals (out of 180); 5 Triple Clusters
Commissioned in 2009 at LNL (with 3 TC)
Experiments at LNL in 2010-2011
Now at GSI, working with 20-25 crystals
S. Akkoyun et al, NIMA 668 (2012) 26–58
28 crystals (out of 120); 7 Quadruple Clusters
Engineering runs started early 2011 at LBNL
Experiments at LBNL in 2011
Now at MSU, working with 28 crystals
S. Paschalis et al, NIMA 709 (2013) 44–55
LBNL, 2011
LNL, 2011

## Slide 27: Doppler Correction 220 MeV 56Fe  197Au (AGATA + DANTE)
Doppler correction using center of crystals
FWHM ~20 keV
Detector FWHM = 2.2 keV
Doppler correction using center of hit segments
FWHM = 7 keV
Doppler correction using PSA (AGS) and tracking FWHM = 3.5 keV
(3.2 keV if only single hits)
Eg (keV)
v/c 8%
E(2+) = 846.8 keV
56Fe 2+  0+
846.8 keV
56Fe 4+  2+
1238.3 keV
4.8 keV FWHM
Au recoils also seen by Dante
Original
Corrected

## Slide 28
GRETINA at BGS
GRETINA – BGS coincidence
Data acquired using separate systems
Use time stamps to correlate data
September 7, 2011 – March 23, 2012
I-Yang Lee

## Slide 29: Doppler corrected spectra
Corrected for 58Ni
58Ni 2+ 1454 keV
FWHM = 14 keV
Corrected for 136Xe
136Xe 2+ 1313 keV
FWHM = 8 keV
Coulomb excitation: 58Ni(136Xe,136Xe’)58Ni
42
42
2+0+
58Ni
2+0+
136Xe
2+0+
58Ni
2+ 0+
136Xe
Particle–g angle (deg.)
Particle–g angle (deg.)

## Slide 30: AGATA at the GSI-FRS in-flight RIB
HECTOR
AGATA
Lund-York-Cologne
CAlorimeter (LYCCA)
12 weeks of beam
New FRS tracking detectors (>106 s-1 at S2, 105s-1 at S4)
New LYCCA-0 particle identification and tracking system
Higher SIS intensities and fast ramping 109(U) to 1010 (Xe, Kr) ions/spill
IKP-Cologne Plunger (under construction)
AGATA
GSI-FRS
ADC
Double Cluster
First part of GSI campaign ended 21/11/2012
Four experiments performed , using up to 19 crystals:
- Coulomb Excitation of n-rich Pb, Hg and Pt isotopes
- Pygmy resonance excitation in 64Fe,
- Isomer Coulex in 52Fe
- Lifetimes in the heavy Zr-Mo region
+ M1 excitation in 85Br, 131In + studies of HE background
Courtesy H-J. Wollersheim

## Slide 31: Doppler-Correction of Uranium X-Rays Technical Commissioning
Doppler-
shift
correction
Au target
X-rays
Au target
X-rays
U beam
X-rays
U beam
X-rays
AGATA
Position
Information
+
LYCCA
particle tracking
Courtesy Norbert Pietralla (INPC2013 talk, session B2); Analysis by Michael Reese
U beam on Gold Target:
thickness 400 mg/cm2
U velocity at Target position:
v/c ≈ 0.5
U-atoms have x-rays around 100 keV
Doppler shift to 100 – 150 keV

## Slide 32: gg capabilities 135 MeV 32S  110Pd (6 AGATA crystals)
138Sm
6 gates on: 347keV,545keV,686keV,775keV,552keV,357keV
871 keV 22+ - 20+
The performance of AGATA using g-ray tracking is comparable to conventional arrays with a much larger number of crystals

## Slide 33: 64Ge g-g, from 65Ge on 9Be at v/c=0.4
plain singles
tracked
Reduction of Compton background by tracking allows – for the first time – gamma spectroscopy with fast beams with spectral quality comparable to arrays with anti-Compton shields.

## Slide 34: Imaging of Eg=1332 keV gamma rays AGATA used as a big Compton Camera
F. Recchia, Padova
Far Field Backprojection
Near Field Backprojection
All 9 detectors
One detector
All 9 detectors
One detector
Source at 51 cm  Dx ~Dy ~2 mm Dz ~2 cm

## Slide 35: Polarization with AGATA crystals
Coulex test experiment with 2 AGATA clusters (6 crystals)
  12C (32 MeV) on 104Pd (2+ at 555.8 keV) and 108Pd (2+ 433.9 keV)
  angular efficiency normalized on 137Cs source (666.6 keV)
Similar study done at TU-Darmstadt using one AGATA crystal; hits placed at center of fired segments (no PSA)  B.Alikani et al. NIMA 675(2012)14
Large dataset taken at the end of the Legnaro campaign by P.G.Bizzeti with 2 facing AGATA triple-clusters at 3 different distances to study the entanglement of 511 keV photons from the b+ 22Na source.
AGATA-Demonstrator experiment Non-yrast octupole bands in the actinides 220Ra and 222Th by J.F.Smith and D.Mengoni
More details: D.Mengoni, Session I5, Thursday
Presented by B.Melon Session A1, Monday

## Slide 36: From the Demonstrator to AGATA 1p Plans for the next few years
LNL: 2009-2011 15 crystals (5TC)
Total Eff. ~6%
GSI: 2012-2014 25 crystals (5DC+5TC)
Total Eff.~10%
GANIL: 2014-2016 45 crystals (15 TC)
Total Eff. ~15%
AGATA+VAMOS
Demonstrator + PRISMA
AGATA + FRS
Talk by D.Mengoni session I5, thursday
Talk by N.Pietralla session B2, monday

## Slide 37: GRETINA Science Campaigns
ANL FMA
NSCL S800
July 2012 - June2013
2013 - 2014
Single particle properties of exotic nuclei – knock out, transfer reactions.
Collectivity – Coulomb excitation, lifetime, inelastic scattering.
24 experiments approved for a total of 3351 hours.
Structure of Nuclei in 100Sn region.
Structure of superheavy nuclei.
Neutron-rich nuclei – CARIBU beam,
deep-inelastic reaction, and fission.

## Slide 38: Science campaign at NSCL: July 2012 – June 2013
GRETINA
S800
GRETINA
electronics
24 experiments approved: 3351 hours
GRETINA at target position of S800 spectrograph
Talk by I-Y. Lee I5, Thursday

## Slide 39: Summary
AGATA (first phase) and GRETINA are constructed and commissioned.
Several problems have been successfully solved.
Number of detectors will increase over the next years.
Performance expected to improve over the years due to progress in electronics and data processing algorithms.
Physics Campaigns have been performed at LNL, GSI, LNBL, MSU
Physics Campaigns planned ANL, GANIL, MSU for the next years.
AGATA and GRETA/GRETA will be major instruments for the next generation of facility such as FAIR, FRIB, SPES, SPIRAL2, …
Gamma-ray tracking arrays will have a large impact on a wide area of Nuclear Physics.