## Slide 1: 3D diamond detectors
Felix Bachmair on behalf of 3DD/RD42 collaboration
RD13, Florence, 3. – 5.6.2013
Notes: 3DD:
University of Manchester,
CEA Saclay,
CERN,
ETH Zurich,
Ohio State University

## Slide 2: Overview
Motivation
Fabrication
Properties of conducting columns
Test beam results
Conclusion & Outlook

## Slide 3: Motivation – Why 3D?
Drift distance shorter than in planar (≈ 100μm vs. 500μm)
Comparable to mean free path of charge carriers in irradiated diamond (few 1015 p/cm2)
Results in more efficient charge collection
Planar
3D
150 μm

## Slide 4: Fabrication – Surface treatment
100µm
Surface
Treatment
Column
Formation
Metallization
Notes: OHIO: acid boiling -> remove surface dirt
not a hydrogen terminated surface, terminate with oxygen before drilling holes

## Slide 5: Fabrication – Column formation
Surface
Treatment
Column
Formation
Metallization
Femto second laser
  Wavelength = 800 nm
  Repetition rate = 1kHz
  Pulse duration = 100 fs
  Spot size = 6 μm
  Pulse Energy:
    E = 1 – 23 μJ/pulse
    Φ = 1 – 30 J/cm2
Notes: In Manchester
run with plexi glas
20 mum per second --> 30 sec/ column
PICTURE LASER Beam goes through the diamond moves upwords

## Slide 6: Fabrication – Column formation
Surface
Treatment
Column
Formation
Metallization
Notes: In Manchester
run with plexi glas
20 mum per second --> 30 sec/ column
Hole is put from buttom to the top
Different connection at bottom and top 

## Slide 7: Fabrication
Surface
treatment
Column
Formation
Metallization
Notes: Ohio: Titanium, platinum, gold

## Slide 8: Conducting Columns
Seed Side

## Slide 9: Conducting Columns
Exit Side

## Slide 10: Conducting Columns – micro machining
 | Low Power | High Power
Low Speed
High Speed
Seed Side

## Slide 11: Conducting Columns – micro machining
 | Low Power | High Power
Low Speed
High Speed
Exit Side

## Slide 12: Conducting Columns – micro machining
 | Low Power | High Power
 | 92.2 ± 1.4 % | 78.7 ± 2.1%
 | 93.3 ± 1.3 % | 87.6 ± 1.7 %
Low Speed
High Speed
371 columns for each parameter set
Assuming binominal distribution for error calculation
Notes: Power Speed Fails
Low High 25
Low Low 29
High High 46
High Low 79

## Slide 13: Missing Columns

## Slide 14: Column Resistance: Contact
Metallization & Column OK: ~ 90 %
Resistivity of columns:
Notes: How does this measurement work:
Metallization of the columns on back with small pads,
Measure the resistance

## Slide 15: CERN Test beam
Three structures:
Strip detector as a reference to compare to with backplane (bias voltage)
3D mask layout (bias and readout from the same side)
  Without machined columns to understand the electric field
  With machined columns
Strip
3D mask no columns
3D mask with columns
Blue: front side

## Slide 16: CERN Test beam
Three structures:
Strip detector as a reference to compare to with backplane (bias voltage)
3D mask layout (bias and readout from the same side)
  Without machined columns to understand the electric field
  With machined columns
Strip
3D mask no columns
3D mask with columns
Blue: front side
Brown: backside

## Slide 17: CERN Test beam
Using Strasbourg strip telescope in H6A (NA-SPS) with 120 GeV protons (4X, 4Y planes, Resolution: a few μm)
Triggering on coincidence of 2 scintillators
≈ 20 cm

## Slide 18: CERN Test beam - Clustering
After pedestal subtraction: Clustering
  Seed  5 σ
  Hit  3 σ
  At least one Seed needed to form a cluster
Seed Cut
Hit Cut

## Slide 19: CERN Test beam – Telescope Resolution
σ =
2.04μm
σ =
1.94μm
σ =
2.13μm
σ =
2.11μm
Silicon strip detectors
Using eta corrected hit position for each detector

## Slide 20: Diamond alignment
Using strip detector for alignment: highest hit position
Before alignment
After alignment

## Slide 21: Results
Showing results for 3D single crystal CVD (batch2)
  Strip detector:
    Bias Voltage VBias = +500 V
  3D detector
    Bias Voltage VBias = +25 V
  1.1M events collected

## Slide 22: CERN Test beam – Fiducial regions
Requiring one and only one cluster in each telescope plane
Plotting average telescope position for selecting rough fiducial region
Requiring at least one Diamond cluster
Mean charge of first cluster

## Slide 23: Fiducial Regions
Mean charge of cluster with the highest pulse height
In diamond detector:

## Slide 24: Fiducial Regions
Overlay with pattern of mask
Mean charge of cluster with the highest pulse height

## Slide 25: Fiducial Regions
Fiducial regions for the three different patterns
Mean charge of cluster with the highest pulse height

## Slide 26: CERN Test beam
 Strip @ 500V
3D with columns @ 25V 
Strip
3D no columns
3D
with columns
500
1000
1500
Charge of diamond cluster in ADC counts

## Slide 27: PH in strip detector
Similar pulse height in every channel
Landau distribution, MP @ ~1000

## Slide 28: PH in 3D without columns
Influence of strip structure (500V vs. 25V), which is at lower channels
Strip Detector @ 500 V
3D with columns @ 25 V
Effect of Strip
Detector @ 500V

## Slide 29: PH in 3D with columns
Gaussian tail to lower pulse heights
Small ‘bump’ at low pulse height ~ 400ADC

## Slide 30: Comparison: 3D with columns vs. Strip detector
3D with
Columns
@ 25 V
Strip
@ 500 V

## Slide 31: Strip vs. 3D
Peaking at the same pulse height
@ 500 V
@ 25 V

## Slide 32: Strip vs. 3D
Peaking at the same pulse height
@ 500 V
@ 25 V
Missing
Readout columns?

## Slide 33: Y Alignment
3715 um
Predicted from strip detector alignment
3715 μm

## Slide 34: Y Alignment
5530 um
3715 um
Predicted from strip detector alignment
3715 μm

## Slide 35: Pulse height distribution with overlaid grid
8 cells which seems to have a broken readout column (predicted 7 [p.12] )

## Slide 36: Area without missing readout column
Notes: Can be increased to a bigger section

## Slide 37: Pulse height for this area
good agreement of strip detector and 3D detector

## Slide 38: Conclusion
Produced conductive columns within diamond bulk
Achieved ≈ 90 % success rate in column formation
Success in metallization and column contact
Test beam results  columns collects full charge with much lower bias voltage

## Slide 39: Outlook
Batch 3: new Layout
  Possibility of better alignment in y
  Better isolation between structures
  Bias voltage from the back
  Cubic and hexagonal cells
  Isolated cells
Radiation tests to be done
More groups to start with 3D diamonds within RD42, e.g. Gottingen

## Slide 40: backup

## Slide 41: Why diamond?
 | Si | diamond
Band Gap [eV] | 1.12 | 5.45
Electron Mobility [cm2/V s] | 1450 | 2200
Hole Mobility [cm2/V s] | 500 | 1600
Saturation Velocity [cm / s] | 0.8 · 107 | 2 · 107
Breakdown Field [V / m] | 3 · 105 | 2.2 · 107
Resistivity [Ohm] | 2.3 · 105 | > 1011
Dielectric Constant | 11.9 | 5.7
Displacement Energy [eV] | 13-20 | 43
e-h creation energy [eV] | 3.6 | 13
Average e-h pairs per MIP per μm | 89 | 36
Charge Collecting Distance [μm] | ~full | ~ 250-300
Thermal conductivity [W/cm K] | 1.5 | 22
Low leakage current
Low capacitance
Operation at high voltages
Operation without cooling
Radiation tolerant
Low signal
Single crystals small in size
Still quite expensive
Notes: Once if it is irradiated: Mean free path is longer than the drift distance

## Slide 42: TCAD Simulation
+40V
Slice @ Z=250μm
(25,25)
MIP
42
150µm

## Slide 43: Side View

## Slide 44: Motivation – Single Crystals
Cross polarized image
www.diamondcvd.com
Notes: Cross polarized image with microscope
Defects are changing polarization
Show up as bright regions
Star == point defect
remove

## Slide 45: Motivation – Polycrystalline Diamond
P.W.May, DRM, V19, I5-6,p369-396, 2009
Notes: remove

## Slide 46: Fiducial Regions
Overlay with pattern of mask and Fiducial cuts for each pattern

## Slide 47: Mean pulse height for each cell