## Slide 1
GEM Digital (KLOE) and Analog (COMPASS) readout
&
BesIII readout requirements (DAQ & FEE)

## Slide 2: Outline
GEMs - main parameters
  Gain
  Discharges
  Cluster multiplicity
GEM Digital Readout (KLOE IT)
  readout electrodes
  layer instrumentation
  HV sectors
  HV distribution network
  readout plane details
  FEE integration
  HV & FEE chain
  readout chain
GEM Analog Readout (COMPASS)
  readout electrodes
  chambers
  position accuracy
  efficiency vs track multiplicity
  time measurements accuracy
  APV 25 front-end
GEM operation in magnetic fields
  charge broadening (simulation)
  Test beam setup
  resolution
  efficiency
  cluster size
BES IT DAQ & FEE requirements
  DAQ requirements
  2D readout
  CSTRIPS effect on SNR
Conclusions

## Slide 3
GEMs

## Slide 4: GEM: Triple GEM

## Slide 5: GEM: gain and discharge probability
TOTEM GEM detectors for tracking and triggering
Leszek Ropelewski CERN PH-DT2-ST & TOTEM

## Slide 6: GEM: strips signal
TOTEM GEM detectors for tracking and triggering
Leszek Ropelewski CERN PH-DT2-ST & TOTEM

## Slide 7: GEM: strips induced signal
TOTEM GEM detectors for tracking and triggering
Leszek Ropelewski CERN PH-DT2-ST & TOTEM

## Slide 8
Digital Readout - KLOE IT

## Slide 9: KLOE2 IT - readout electrodes
4 independent tracking layer
XV readout strips
  25-30° stereo angle (depending on the layer)
  X strips (r-φ coordinate): 650 μm pitch - 250 μm width - 694 mm length  C ≈ 100 pF
  V strips (z coordinate): 650 μm pitch (V-pads connected through vias) - length range: 1÷773 mm  C ≈ 1-200 pF
  Overall resolution: σrφ= 200 μm - σz= 500 μm
≈ 30k readout channels
Strip signals readout through a 120-pin connectors (each connector collects 40 X strips and 80 V strips)
X pitch 650µm
 ResX 190µm
V pitch 650µm
 ResY 350µm
XV strips on the same plane
3 foils spliced without overlap: kapton strips are glued on the back of head-to-head joints
1k strips
1M pads
~1mm
dead zones

## Slide 10: KLOE2 IT - layer instrumentation

## Slide 11: KLOE2 IT - X strips and HV sectors

## Slide 12: KLOE2 IT - V strips and HV sectors

## Slide 13: KLOE2 IT - HV Distribution Network

## Slide 14: KLOE2 IT – Readout Plane (Detail)

## Slide 15: KLOE2 IT - FEE Integration
Readout (XV) plane connectors
HV macro-sector connector
Readout boards
HV max ≈ 5 kV !!
GND
HV distribution must be carefully designed

## Slide 16: KLOE2 IT - HV & FEE chain (I)
HV & FEE boards
HV & signals cables
Signals cables collected by means of special boards on the back side of a custom crate
Custom HV distributors
Concentrator (hidden by fibers)
A. Aloisio et al., “An FPGA Based General Purpose DAQ Module for the KLOE-2 Experiment” Journal of Physics: Conference Series 331 (2011) 022033

## Slide 17: KLOE2 IT - FEE chain
BARI
Uppsala Institute
(Pawel Marciniewski)
5 mt Twisted pairs
Optical Links
Transition Board
(collects cables from 8 Gastone boards)
GIB Board
(manages up to 8 Gastone Boards)
Data download
Lev1 distriution
Gastone boards parameters setting
1024 channels
GIB crate
(back view)
GIB crate
(front view)
128
channels
Concentrator Board
manages 16 GIB boards
Event building
D
A
Q
TB – GIB interface
LNF
A. Balla et al., “GASTONE: A new ASIC for the cylindrical GEM inner tracker of KLOE experiment at DAFNE”, ,Nucl. Inst. & Meth. A, vol. 604, no. 23, 2009

## Slide 18
Analog Readout - COMPASS

## Slide 19: COMPASS: readout electrodes
Due to diffusion the charge cloud collected on the readout board is bigger than the strip width (≈ 3.5 x pitch) and a weighting method is used for calculate the exact track position in two dimensions
Every single strip versus the other readout coordinate acts as a plate capacitor. With the permittivity ε=3.9 of Kapton and an area of 2.27⋅10-1cm2, this capacitance is 15.7 pF
50 µm kapton

## Slide 20: COMPASS: chambers

## Slide 21: COMPASS: position accuracy

## Slide 22: COMPASS: time measurement accuracy

## Slide 23: COMPASS: efficiency

## Slide 24: COMPASS: front end (I)
Anode and cathode connected to GND
NO DC into preamplifier
Charging up of strips avoided by diode leakage currents
Running Experience with the COMPASS Triple GEM Detectors – Bernard Ketzer

## Slide 25: COMPASS: front end (II)

## Slide 26
GEM operation in magnetic field

## Slide 27: Magnetic field: charge broadening (sim.)
Design and Construction of a cylindrical GEM detector as the Inner Tracker device of the KLOE-2 experiment - Gianfranco Morello on behalf of Frascati, Bari and Cosenza groups

## Slide 28: Magnetic field: test beam setup
Design and Construction of a cylindrical GEM detector as the Inner Tracker device of the KLOE-2 experiment - Gianfranco Morello on behalf of Frascati, Bari and Cosenza groups

## Slide 29: Magnetic field (X-V angle = 40°): resolution
Design and Construction of a cylindrical GEM detector as the Inner Tracker device of the KLOE-2 experiment - Gianfranco Morello on behalf of Frascati, Bari and Cosenza groups
1T

## Slide 30: Magnetic field (X-V angle = 40°): efficiency
Design and Construction of a cylindrical GEM detector as the Inner Tracker device of the KLOE-2 experiment - Gianfranco Morello on behalf of Frascati, Bari and Cosenza groups
Gain = 6x104

## Slide 31: Magnetic field (X-V angle = 40°): cluster size
Design and Construction of a cylindrical GEM detector as the Inner Tracker device of the KLOE-2 experiment - Gianfranco Morello on behalf of Frascati, Bari and Cosenza groups
threshold = 3.5 fC

## Slide 32
BES IT
DAQ & FEE requirements

## Slide 33: BES IT: DAQ requirements
1) Level-1 trigger rate (average): 1 a 2 kHz
2) Dead time: 4 μs
3) Latency: 8 μs +/- 0.5 μs
4) Event size: 12 kB (including ZDD)
5) System (distributed clock) ≈ 42 MHz (3 bunch crossings - 3x8ns)
  Counters to check events alignment ? Timestamp ?
6) Offline t0 reconstruction

## Slide 34: BESIII IT: 2D readout
Main requirements:
  Minimize electrode capacitance to reduce series noise contribution
    Small width strips
    Special strips design to reduce X-V CCOUPLING
    GND plane far from the anode
  Minimize XV views crosstalk
    Special strip design to reduce X-V coupling
    GND plane near to the anode
STRIPS | COMPASS |  | KLOE IT |  | BES IT (I) |  | BES IT (X view) | 
Width | 350μ | 80μ | 250μ |  | 350μ | 80μ | 350μ | 80μ
Pitch | 400μ |  | 650μ |  | 650μ |  | 650μ | 
Capacitance | ≈ 17 pF |  | X ≈ 100 pF | V ≈ 1 ÷ 200 pF | ≈ 41 pF |  | XSS ≈ 41 pF | XSG≈ 41 pF
Ground distance | ≈ 3 mm (?) |  | ≈ 200μ |  | ≈ 3 mm |  | ≈ 200μ | 

## Slide 35: BESIII IT: CSTRIPS effect on SNR (I)
Example 1: COMPASS + APV25
  CSTRIP COMPASS ≈ 17 pF
  ZIN APV25 ≈ 900 Ω (?!)
  APV25 shaping time ≈ 50 ns
  APV25 input noise ≈ 396 e + 59.4 e/pF (deconvolution mode)  Noise ≈ 0.26 fC
rise time ≈ 33 ns
OK
QSTRIP ≈ 6 fC (average, NO gas gain fluctuations)
  Ar/CO2 (70/30)
  G = 8000
  Factor 1/2 because the two view
  Factor 1/3 because the strip multiplicity
SNR ≈ 23

## Slide 36: BES IT: CSTRIPS effect on SNR (III)
QSTRIP ≈ 6 fC (average, NO gas gain fluctuations)
  Ar/CO2 (70/30)
  G = 8000
  Factor 1/2 because the two view
  Factor 1/3 because the strip multiplicity
Example 2: BES IT + APV25
  CSTRIP BESIII ≈ 41 pF
  ZIN APV25 ≈ 900 Ω (?!)
  APV25 shaping time ≈ 50 ns
  APV25 input noise ≈ 396 e + 59.4 e/pF (deconvolution mode)  Noise ≈ 0.45 fC
rise time ≈ 81 ns
NO
SNR < 13 (ballistic deficit not considered !!)
Example 3: BES IT + VFAT2
  CSTRIP BESIII ≈ 41 pF
  ZIN VFAT2 ≈ 120 Ω
  VFAT2 shaping time ≈ 22 ns
  VFAT2 input noise ≈ 400 e + 40 e/pF  Noise ≈ 0.33 fC
rise time ≈ 21 ns
NO/OK
SNR < 18 (ballistic deficit not considered !!)
Example 3: BES IT + GASTONE
  CSTRIP BESIII ≈ 41 pF
  ZIN GASTONE ≈ 120 Ω
  GASTONE shaping time (CD < 40 pF) ≈ 30 ns
  GASTONE input noise ≈ 800 e + 40 e/pF  Noise ≈ 0.4 fC
rise time ≈ 21 ns
OK
SNR ≈ 15

## Slide 37: BES IT: readout conclusion (if any …) and open questions
The experience with the KLOE IT has shown that the preamplifier and the shaper must be carefully designed to maximize readout SNR for CD > 20 pF
A custom device is required for analog strips readout (mainly because the ≈ 8 μs DAQ latency and the foreseen value of CD )
  64/128 channels board
    discharge protection network
    charge measurement - resolution ?
    time measurement - resolution ?
    latency buffer up to 8 μs
    Serial data readout (analog or digital)
XV anode strips readout electrode must be designed to minimize parasitic capacitance
XV crosstalk must be minimized as well
  new strip size/geometry?
Radiation background ? Do we need rad-tol devices near to the detector ? Probably NO, but …

## Slide 38: BES IT: Example of 2D optimized readout
A.Bondar et al., Performance of the triple-GEM detector with optimized 2-D readout in high intensity hadron beam
minimize strip capacitance
decrease efficiency plateau starting point
decrease discharge probability
decrease HVGEM working point

## Slide 39
Spares

## Slide 40: Spares: KLOE DAQ timing
≈ 200 ns
Data readout (dead time)
  ≈ 2 μs (initial design)
  ≈ 4 μs (actual dead time)
Discriminated signals are moved to the readout buffer
Data download is started
If Lev2 GIB: data are routed to the concentrator boards

## Slide 41: Spares: VFAT chip
CMS GEM Upgrade Workshop III 18-20 April 2012
- Stefano Colafranceschi

## Slide 42: Spares: KLOE DAQ timing

## Slide 43: Some results from COMPASS
Hit-residuals distribution
Events arrival time distribution (signal amplitudes measured in three successive time bins of 25 ns)
Cluster size distribution for 80 μm and 350 μm wide strips