## Slide 1: LKr status
R. Fantechi

## Slide 2: Lkr components
Detector
  Ok, sleeping for 2 years..
  Cold since 1998
HV/LV
Readout
  Preamplifiers, transceivers, CREAMs
Trigger
Calibration
  Hardware, software

## Slide 3: High voltage – Low voltage
HV ready to be switched on
  Fix one channel of the current monitoring system which seems to be broken
  Start with the standard 3000 V value
LV for preamplifier and transceiver almost ok
  Need to change fuses to single transceivers and check one regulator card for transceivers
  Likely no transceiver to be replaced
DCS for both subsystems ready

## Slide 4
Preliminary data using only 75% of channels

## Slide 5: CREAM installation

## Slide 6: A fully cabled crate
Black Eth data cables
Gray Eth sum cables
Red choke/error cable
Yellow TTC fiber
Green Eth DCS cable

## Slide 7: Current issues
18 CREAMs still to be delivered at CERN
  Faulty modules repaired at CAEN
  In some cases, need to redo the production
  Additional 5 prototypes available if needed (FW not loadable via VME)
  Enough modules to be able to lend some to Mainz for the MUV readout
Firmware refinements under way
  Proper handling of choke/error conditions
  Fix a problem with TSL handling

## Slide 8: Baseline sigma distributions
Preamplifiers /transceiver off
Preamplifiers /transceiver on

## Slide 9: LKr trigger infrastructure
Crates and power supplies arrived, PS to be installed
Cabling done for racks 6,7,8
Two weeks and half needed to complete the cabling
  From 15/9

## Slide 10: LKr trigger infrastructure

## Slide 11: LKr trigger electronics
Teldes available at CERN now
  To be ent back to PG for testing
Mezzanine receiver cards
  2 preproduction cards assembled and tested
Mezzanine transmitter cards
  2 prototype tested, 4 preproduction PCB being produced this week, 1 week needed and then to be assembled
Cables in preparation
Test bench available for TX/RX cards
Request for a LKr trigger test bench at CERN

## Slide 12: Plans for the early days of the run

## Slide 13: Calibration hardware
Calibration driver electronics unchanged wrt NA48
  Only some cosmetics in connections
The in/out interface for the programming has been changed
  From CAMAC/NIM to VME/TALK board
  Patch panel for connections
VME software rewritten with C++ classes
Calibration sequence software integrated with the run control
Ready to go

## Slide 14: Calibration software
Work going on by Alan and Michal
  Migrate NA48 Fortran code to NA62 C++ one
  Be ready for the 2014 run
  Crosscheck the procedures with 2012 data
  Fundamental not only to get calibration constants but also for a continuous check of the performances of the single channels
  Eventually be able to prepare digital filter constants for the CREAMs

## Slide 15: Strategy

## Slide 16: Time alignment of pulses
Double Fermi-Dirac fit

## Slide 17: Current status