## Slide 1: The WaveDAQ System for the MEG II Upgrade
Stefan Ritt, Paul Scherrer Institute, Switzerland
Luca Galli, Fabio Morsani, Donato Nicolò, INFN Pisa, Italy

## Slide 2: The WaveDAQ System for the MEG II Upgrade
Stefan Ritt, Paul Scherrer Institute, Switzerland
Luca Galli, Fabio Morsani, Donato Nicolò, INFN Pisa, Italy

## Slide 3: DRS4 Chip
Shift Register
Clock
IN
Out
Storage
capacitors
Inverter “Domino” ring chain
0.2-2 ns
FADC
33 MHz
Switched Capacitor Array (Analog Memory) developed at PSI
5 GSPS / 11.5 bits SNR, 9 channels on 5 mm x 5 mm chip, 40 mW / chn.
Used at ~200 locations worldwide
2012: “Gigahertz Waveform Sampling: An Overview and Outlook”
Pile-up rejection O(~10 ns)
Time measurement O(10 ps)
Charge measurement O(0.1%)

## Slide 4: MEG & MEG II
MEG Experiment 1999-2013
Separated DAQ & Trigger
3000 Channels DRS4 (0.8 GSPS / 1.6 GSPS)
1000 Channels Trigger (100 MSPS)
5 Racks
PMT
Active
Splitter
Trigger
DRS
DAQ
HV
Trigger Bus
Clock distribution
S2: M. de Gerone: An extreme high resolution Timing Counter for the MEG II Experiment
S2P: M. Simonetta: Test and characterization of SiPMs intended as detector for the MEG timing counter
S5P: D. Nicolò: An FPGA-based trigger for the MEG II Experiment
S5P: A. Pepino: A high performance Front End Electronics for Drift Chamber readout in the MEG II Experiment
S7P: M. Grassi: A new cylindrical drift chamber for the MEG II Experiment
S7P: G. Rutar: A Dedicated Calibration Tool for the MEG and MEG II Positron Spectrometer
S7P: L. Galli: MEG II drift chamber prototype characterization with the silicon based cosmic ray tracker at INFN Pisa
S7P: M. Venturini: Ageing tests for the MEG II drift chamber
S9P: D. Nicolò: A liquid hydrogen target for the calibration of the MEG and MEG II liquid xenon calorimeter
S9P: K. Ieki: Upgrade of the MEG liquid xenon calorimeter with VUV-light sensitive large area SiPMs
SiPM
TDAQ
MEG II Experiment 2014-
9000 Channels
Same rack space Combine DAQ & Trigger
Power

## Slide 5: Crate Options
Feature | VME | ATCA | ???
Transfer speed O(100 MB/s) | ✔ | ✔ | ✔
Dual-Star Topology with Gbit links | ✗ | ✔ | ✔
Shelf management | ✗ | ✔ | ✔
Fast trigger distribution | ✗ | ✗ | ✔
Low-jitter precision clock O(ps) | ✗ | ✗ | ✔
200 V SiPM biasing | ✗ | ✗ | ✔
< 2000 EUR per crate including power | ✗ | ✗ | ✔

## Slide 6: WaveDAQ System
Crate Management Board
Power supply 24V / 300W
Fan / Temp. control
Power cycle each slot
FPGA Firmware upload
Ethernet remote control
Fans blow from back to front
16 WaveDREAM boards (256 channels)
Data Concentrator Board
Trigger Concentrator Board

## Slide 7: WaveDREAM Board (WDB)
Drs4 based REAdout Module
Spartan 6

## Slide 8: Preamp
Gain | BW3db (MHz) | Noise (mV)
1 | 940 | 0.37
10 | 880 | 0.40
100 | 300 | 1.2
100 | 500 | 1.7
100 | 800 | 3.3
LMH6629
LMH6629
PE4215
PE4215
ADC904
LTC6409
Different compensations
3.3 mV at output = 33 mV at input

## Slide 9: WaveDREAM2 HV
+/- 1mV
Ripple < 10 mV
~ 5 EUR / channel
Microblaze @
Spartan

## Slide 10: Temperature Sensor Extension
accuracy ±0.5∘, 3 EUR / sensor
1-16 sensors per WD2 board with only one coaxial cable
Automatic HV adjustments with temperature changes

## Slide 11: Trigger Concentrator Board (TCB)
Receives serial links (SERDES) from WD boards
Computes crate local trigger
Send trigger via serial links to global trigger in dedicated crate
FCI Densishield cables
Crate local trigger
Global trigger
KINTEX 7
KINTEX 7

## Slide 12: Ancillary system
Contains master clock
Distribute clock (jitter < 12 ps measured)
Distribute trigger
4 diff. pairs for
  Clock
  Trigger
  Busy
  (Sync)

## Slide 13: DAQ Concentrator Board (DCB)
Receive Gbit links from WDB
Use SERDES instead GTX (lower latency)
Waveform preprocessing in Zynq CPU
Output via Gbit Ethernet (10 Gbit optional)
Board under design
Tests with Zed-Board and “Backplane Simulator”

## Slide 14: Half Height Backplane

## Slide 15: SPI configuration

## Slide 16: Gbit links for DAQ & Trigger
Event Builder PC
Global Trigger

## Slide 17: Trigger Bus & HV

## Slide 18: Clock Distribution

## Slide 19: WaveDAQ Clock Distribution
master clock
DCB
FCI Densishield Cable
WD
WD
WD
WD
Crate
LMK03000 Jitter Cleaner
Goal: < 5 ps clock jitter at system level

## Slide 20: Pin Assignment

## Slide 21: Minimal System
Power-over-Ethernet

## Slide 22: One-crate system
220 V
Gbit Ethernet
Up to 256 Channels
Detector

## Slide 23: MEG II System
35 crates x 256 channels = 9000 channels
. . .

## Slide 24: WaveDAQ Performance
Trigger resolution 10 ns (100 MHz clock)
Trigger bandwidth 8 Gbit / s
Trigger latency <380 ns *) (9000 channels)
DAQ bandwidth 2 Gbit / s
DAQ time measurement 10 ps *)
DAQ dead time 3 - 35 ms / event
MEG II: 7 x 107 m/s, DAQ eff. > 95% @ 30 Hz *)
*) projected

## Slide 25: Conclusions
WaveDAQ system has been designed to fulfill needs of MEG II experiment
System has huge potential for many others (costs: ~130 EUR / channel)
Status: Crate fully working, trigger board and WaveDREAM board successfully tested, firmware to be finished, DCB under design
First full crate test end of 2015, full system (35 crates) in 2016
DRS5 chip (no dead-time) planned for 2017+