## Slide 1
DOM acceptance and calibration
Sites: INFN Napoli , INFN Catania, ECAP, Nikhef, other
Man Power: All site
Time Plan: DOM Assembly with all pre-qualified probes
See PRR DOM time plan
Testing and calibration
15000 PMTs / 31 (PMT/DOM) in 4 sites = 121 DOMs per site  6 months (20 days/month)
Instrumentation: Green Box
Mechanical Frame € 4.000
Laser source (available at INFN Naples and Catania) € 15.000
Optical splitter 1x4 : 5 PMT + 1 power meter € 3.000
Improved Optical Power Meter (fibre monitor) € 5.000
UPS for the laser source € 1.500
1 full shore station
GPS receiver
acoustic emitter, signal generator (GPS-PPS trigger) € 2.000
rotating dish + reference € 2.000

## Slide 2
Measurements and time details:
DOM integration 12 h
Gel solidification 12 h
Green box closing and pre-setting 30 min
Set Gain (PMT on at Nominal HV for pulse mode operation)
Recover correspondance of number of PMT/base and mounting position from the Database
Wait conditiong of PMTs 1 h
Functioning test of the DOM using laser source 30 min
Wait or not 8 h? depending on integration schedule
A:Dark Current Rate (with glass and gel) 10 min
Dark rate on all PMTs  verify the specs of the PMT (if not long dark adaption observe only large deviations of few PMts)
Mapping and relative time calibration using 40 K coninc (see http://pi1222.physik.uni-erlangen.de/Qualification/24)
TOT check using single photo-electrons from 40K
B: Laser calibration (absolute) 4 reference PMTs (far one from the other) - 2 min
Laser triggered by GPS and amplitude set to 10 spe and 20 kHz
Check absolute time on reference PMTs (TT or delta-T)
Check relative to absolute time for PMTs close to the fiber
Redo mapping and check with previous one
C:Functionality test of nano beacon 5 min
maximum amplitude @ 20 kHz
minimum amplitude @ 20 kHz
Check the nanobeacon functionality
Redo mapping for PMTs close to the nano-beacon and check with previous ones
D: Functionality test of acoustic reception (20,30,40 kHz, 2 amplitudes, acoustic emitter trigger GPS-PPS) 5 min
Check functionality of the piezo
rough estimate to sensitivity to different frequencies and measure of noise
check time correlation signal – master clock
D:Compass Calibration using rotating dish 15 min
readout of AHRS > 1Hz, PC with RS232 to drive the step motor encoder
store compass data and AHRS data  recover the AHRS calibration matrixes
DOM Packing 20 min

## Slide 3
Parameters to be saved:
Histo coincidences for mapping (to be defined)
Mapping matrix + indication on PMT (green, yellow, red)
TOT (@ spe) – green box (histo, mean, sigma fit, or RMS)
Dark rate (specifying dark adaption time)
Histo TOT (dark rate counts)
31 Histo Delta-T (TtT)

## Slide 4
Acoustic emitter
Holder
Holder
Rotating dish
(encoded motor below)
VEOC-DOM: extra length for 1 complete 360° turn
Base used also for assembly

## Slide 5
Optical splitter for time calibration
Soft flange
For VEOC-DOM

## Slide 6
Acoustic emitter
Holder
Holder
Rotating Dish
Encoded motor below)
Optical splitter for time calibration
Soft flange
For VEOC-DOM