## Slide 1: Integrazione PMT e BaseBoards: procedure, test e calibrazione
Carlos Maximiliano Mollo
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 2
Introduzione
Integrazione PMT e BaseBoard e prime misure
Studio sull’andamento anomalo dei dark counts
Procedura di coating
Progettazione della stazione di test per i PMT di KM3NeT
Lavorazioni meccaniche
Elettronica e test
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 3: Test setup
Carlos Maximiliano Mollo
INFN (Naples)
15 cm
Dark Box
PC
NI-USB to I2C
MiniOctopus
Power Supply @3.3V
Lecroy Wavepro SDA 760Zi 40 GS/s
LVDS ToT signal
LabVIEW
Software
PMT+Base Board
USB
I2C
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 4: Dark Counts
Carlos Maximiliano Mollo
INFN (Naples)
Base Board without conformal coating
PMT ZB7845
From Hamamatsu Datasheet:
HV=-1270V (Hex code 0xB9)
DC= 370 cps
TTS= 3.578 ns
First Measurement
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 5: Dark Counts
Carlos Maximiliano Mollo
INFN (Naples)
ZB7845
Chart: Dark counts vs time
| 0.0 | 15.0 | 30.0 | 95.0 | 125.0 | 175.0 | 245.0 | 1050.0 | 1410.0 | 2350.0 | 2520.0 | 2980.0 | 3350.0
 | 10000 | 4300 | 3500 | 2200 | 2078 | 1950 | 2150 | 11500 | 18000 | 48000 | 52000 | 49000 | 47430
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 6: Dark Counts
Chart: Dark counts vs time
| 0.0 | 15.0 | 30.0 | 95.0 | 125.0 | 175.0 | 245.0 | 1050.0 | 1410.0 | 2350.0 | 2520.0 | 2980.0 | 3350.0
 | 10000 | 4300 | 3500 | 2200 | 2078 | 1950 | 2150 | 11500 | 18000 | 48000 | 52000 | 49000 | 47430
Carlos Maximiliano Mollo
INFN (Naples)
ZB7845
PROBLEMA
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 7: Dark Counts: solution
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 8: Dark Counts: solution
Dark count measurements with different holder materials
1) 2 layers of insulating tape on the PMT and neoprene
2) Silastene on the PMT and 2 layers of insulating tape
3) 4 Layers of insulating tape
1
2
3
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 9
Carlos Maximiliano Mollo
INFN (Naples)
And the winner is…
Chart: Dark Counts vs Time after power-up
Oleg's data | 53600 | 41700 | 22500 | 8800 | 6600 | 6200 | 4730 | 4440 | 2610 | 2040 | 950
3 hours
8 hours
T= 17.8 °C
24 hours
Dark Counts
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 10
Understanding of the DC anomaly
One of the possible explanations for the DC anomaly is the creation of discharges in air due to gaps between the glass surface of the PMT and the holder.
First Setup
UUT: PMT ZB7789 No insulating tape,
no coating (HV:-1100 V DCHamamatsu: 592 cps)
PROBE: PMT ZB7858 with insulating tape
(HV:-1250 V DCHamamatsu: 1180 cps)
If the UUT is noisy and the effect is due to light emission we will see in the PROBE high dark counts when UUT is ON.
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 11
Understanding of the DC anomaly
As expected the UUT is noisy: NO INSULATING MATERIAL BETWEEN THE PMT AND THE HOLDER
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 12
Adaption
test
Understanding of the DC anomaly
Instead, the PROBE PMT reaches fast the nominal DC (1180 cps) using insulating tape
After the adaption we powered up the UUT …
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 13
UUT:ON
UUT:ON
UUT:ON
UUT:OFF
UUT:OFF
UUT:OFF
UUT:OFF
Understanding of the DC anomaly
We powered up and down the UUT and we observed changes in the PROBE DC rate.
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 14
Understanding of the DC anomaly
Is really the correlation between the PROBE and the UUT due to photons?
Second Setup
UUT: PMT ZB7789 No insulating tape,
no coating (HV:-1100 V DCHamamatsu: 592 cps)
PROBE: PMT ZB7858 with insulating tape
(HV:-1220 V DCHamamatsu: 1180 cps)
Black paper between the PMTs
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 15
UUT:OFF
UUT:ON
UUT:ON
UUT:ON
UUT:OFF
UUT:OFF
UUT:OFF
Understanding of the DC anomaly
The correlation disappeared. A noisy PMT generates photons!
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 16
Understanding of the DC anomaly
But are really the photons due to the scintillation between the PMT and the holder or are they produced by the normal functioning of the dynodes?
(Light Emission from Light Detectors, Max L. Knötig et al , Max Planck Institute fur Physik)
UUT: AA0347
PROBE: ZB7858
Third Setup
UUT: PMT AA0347 coating on the tube
(HV:-990 V)
PROBE: PMT ZB7858 with insulating tape
(HV:-1250 V DCHamamatsu: 1180 cps)
There is a transparent material with a high
electric impedance?
YES!!! The same coating used for the bases!!!!!!!
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 17
Chart: Dark Counts vs Time AA0347 scintillation test
Understanding of the DC anomaly
The UUT (PMT AA0347) with insulating coating on the tube reaches quickly a good DC rate.
No anomalies on the DC behavior
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 18
Chart: Dark Counts vs Time ZB7858 scintillation test
UUT:OFF
UUT:OFF
UUT:OFF
UUT:OFF
UUT:ON
UUT:ON
UUT:ON
Understanding of the DC anomaly
As shown in the plot the DC of the PROBE does not change when the UUT is powered up/down
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 19
Since the holder could not perfectly stick to the PMT glass surface, there could be air gaps with distances below 0.2 mm where electrical discharges occur.
In the case of good insulating material covering the glass tube the external surface of the coating is at 0V. Therefore, electrical discharges cannot occur.
Understanding of the DC anomaly
The higher the negative HV the most evident the effect! As shown in the last meeting
at Leiden.
A good insulation material in perfect contact with the PMT surface avoid the effect.
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 20: Prova definitiva sull’origine dell’andamento anomalo dei dark counts
PMT ZB7789
Nessun isolamento elettrico fra supporto e
Fotomoltiplicatore
Flusso di SF6 (pressure 2 bar) in corrispondenza
Del punto di contatto fra fotomoltiplicatore
e supporto
Se l’origine dell’andamento anomalo dei DC è data da micro scariche fra il vetro
Del PMT ed il supporto, la presenza di un gas fortemente elettronegativo come
L’SF6 dovrebbe abbattare i DC
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 21: Prova definitiva sull’origine dell’andamento anomalo dei dark counts
La presenza di SF6 altera i DC. L’unica spiegazione: Riduzione di scariche esterne
al PMT.
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 22
An innovative method to control the dark counts
THE LIQUID COATING CAN BE USED TO IMPROVE THE DC RATES
INSULATING
TAPE
COATING
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 23
An innovative method to control the dark counts
CHINESE PRODUCTION
DC RATE TREND OVER A PERIOD LONGER THAN 1 DAY
GOOD STABILITY!
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 24
An innovative method to control the dark counts
THIS SOLUTION IS:
FAST: the coating on the PMT will be done simultaneously with the base board coating
CHEAP: no additional costs or man power is required
EASY: the PMT coating procedure is very simple and is not different with the respect to the base board coating procedure.
0 IMPACT: the coating is transparent, the coating layer is very thin and do not interfere with the holder mechanic design.
Ok but: may the coating interfere with the optic gel?
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 25
Compatibility tests, measurements and procedure
Liquid coating and GEL compatibility
In order to check the gel-coating compatibility we poured the gel around a PMT whose glass tube was covered with coating. Note that, given the goal of this test, we did not follow the procedure to eliminate the bubbles from the gel.
it was not sticky in any place close or far the glass where the coating was applied.
The situation after about 72 hours
The gel was well polymerized
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 26
Compatibility tests, measurements and procedure
In order to be effective the coating must be placed between the PMT and the O-Ring in the DOM holder.
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 27
Compatibility tests, measurements and procedure
THE LIQUID COATING APPLICATION PROCEDURE
Pour the liquid in a container large enough to immerse the PMT up to the position indicated by the red arrow in the previous slide.
Immerse the PMT in the coating with the fingers on the longer diameter. Move the PMT slowly to let the liquid coating enter between the various components of the base board and wait 2-3 seconds. DO NOT immerse the connector in the coating! Make sure that the connector of the base board remains clean. Of course, part of the flat cable between the base board and the connector will be immersed in the coating
Remove the PMT from the container and tilt it of 45° so that the coating excess drips. During this operation is recommended to rotate the PMT to uniform the layer of coating ;
Wait until it stops to drip;
Place the PMT in a holder with a diameter slightly less than the maximum diameter of the PMT;
Wait at least 48 hours (depending on the room temperature) so that the coating is dry.
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 28
Compatibility tests, measurements and procedure
Testing PMTs hosted in the PMT support structure
PMT ZB7858: NHV=1250 V, DC= 1180, insulating tape used as insulator for the tube.
PMT AA0347: NHV= 990 V, DC= 250 conformal coating used as insulator for the tube.
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 29
Testing PMTs hosted in the PMT support structure
Compatibility tests, measurements and procedure
Chart: DC (cps) vs Setup
| A | B | C | D
DC (Hz) | 250 | 350 | 480 | 530
cps
setup
Setup A: PMT AA0347 placed in our “homemade” holder with liquid coating
Setup B: PMT AA0347 placed in the DOM holder with
Insulating tape in the region near the O-Ring and the
O-ring.
Setup C: PMT AA0347 placed in the DOM holder with
O-Ring and no insulating tape or coating between the PMT and the O-ring.
Setup D: PMT AA0347 placed in the DOM holder without O-Ring and no insulating tape or coating.
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 30
Compatibility tests, measurements and procedure
Testing PMTs hosted in the PMT support structure with and without light concentrator
PMT ZB7858: NHV=1250 V, DC= 1180 cps
Setup 1: With concentrator ring
Setup 2: Without concentrator ring
Setup 1: final DC rate about 1300 cps (concentrator ring not in contact with the PMT)
Setup 2: final DC rate about 1150 cps
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 31
Compatibility tests, measurements and procedure
Concentrator ring in contact with the PMT
PMT ZB7858: NHV=1250 V, DC= 1180 cps
Setup 1: With concentrator ring
Setup 2: Without concentrator ring (not in contact)
Setup 1: final DC rate about 1300 cps
Setup 2: final DC rate about 1150 cps
Setup 3: Without concentrator ring (in contact)
Setup 3: final DC rate about 1150 cps
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 32
Compatibility tests, measurements and procedure
Concentrator ring in contact with the PMT
It seems that if the concentrator ring is in contact with the PMT the DC trend is better.
WHY?
The concentrator ring is a conductor, if it is directly in contact with the PMT its potential is the same as the PMT surface.
No micro-discharges will occur.
The measurements was done using the ZB7858 that works at
-1250V. Ones of the highest nominal HV.
All effects will be most evident for those PMTs with high negative voltage.
Conclusion: the contact between the PMT and the concentrator ring is not critical
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 33: Dark box for PMTs assembly in view of the mass production
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 34
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box components
LASER source
PiLas Mod. EIG2000DX | 
Repetition rate (internal trigger) | 50 Hz – 1 MHz
Repetition rate (external trigger) | Single shot – 1 MHz
External trigger input | TTL & VAR upt 120 MHz
External trigger pulse width | Typ. ≥ 4 ns
Synchronization output pulse width | typ. ≥ 4 ns for external triggering 50% duty cycle for internal oscillator
Typical jitter between synchronization trigger output and optical signal | typ. ≤ 4 ps
Warm-up time | < 5 minutes
LASER head | Wavelength 470 +/- 10 nm, spectral width < 7 nm, pulse width (FWHM) typ < 45 ps
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 35
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box components
Dual Channel Power Meter
Newport Mod. 2936-C | 
Full-Scale Current | 2.5 nA – 25 mA
Resolution | 10 fA
Accuracy | 0.2%
SNR | 100 dB
Bandwidth | 1.2 Hz – 400 kHz
Minimum Power Resolution | 11 fW
Newport Mod. 918D-SL-OD3 | 
Material | Silicon
Spectral Range | 400 to 1100 nm
NEP (detector noise equivalent power) | 0.55 pW/√Hz
Power Meter Probes
Minimum Detectable Power 0.60 pW
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 36
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box components
1 to 32 and 1 to 2 Optical splitters
Opal Diffusing Glass
Fiber-to-fiber U-bench for filters
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 37
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box components
Near Lambertian Distribution
Opal diffusing Glass
Chart: FC/PC LASER OUTPUT POWER DISTRIBUTION
| 0.0 | 5.0 | 10.0 | 15.0 | 20.0 | 25.0 | 30.0 | 35.0 | 40.0 | 45.0 | 50.0 | 55.0 | 60.0 | 65.0 | 70.0 | 75.0 | 80.0 | 85.0 | 90.0 | 95.0 | 100.0 | 105.0 | 110.0 | 115.0 | 120.0
 | 0 | 0 | 12.12 | 56.4 | 185.3 | 468.1 | 1023 | 1922 | 4150 | 6220 | 7190 | 8556 | 8736 | 8620 | 7110 | 6500 | 4040 | 2030 | 956 | 465.6 | 166.7 | 43.8 | 5.44 | 0 | 0
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 38
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box design
Optical scheme
62 PMTs under test In parallel
Single LASER source
Dual channel Power Meter for LASER monitoring
2 high resolution, low power probes
1 splitter 1 to 2
2 splitters 1 to 32
62 opal diffusing glass disks
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 39
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box concept design
PMT Hamamatsu
PMT holder unit
D: 83 mm
D: 54 mm
D: 73 mm
lateral
front
back
O-Ring
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 40
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box concept design
Two trays of 31 PMTs each
top
bottom
57 cm
66 cm
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 41
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box concept design
Two trays of 31 PMTs each
lateral
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 42
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box concept design
The Dark Box
120 cm
88 cm
58 cm
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 43
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box concept design
The Dark Box
62 dark cylinders:
each PMT is optical insulated
with the respect of the others
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 44
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box concept design
The Dark Box
Opal diffusing
disk
FC/PC Pannel
Connector
Optical fiber
15 cm
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 45
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box concept design
Trays positioning
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 46
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box concept design
Trays positioning
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 47
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box concept design
Trays positioning
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 48
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box concept design
Guide hole
Tray pin
Trays positioning
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 49
Carlos Maximiliano Mollo
INFN (Naples)
PMTs + bases TEST – Dark Box
  Dark Box concept design
The Dark Box Electronics
2 DOM electronics:
4 Octopus boards
2 CLBs
2 Power Boards
1.5 meters extension
Cable Between
Octopus Boards
and PMT bases
(Marco Circella)
KM3NeT – IT Meeting – Catania 28 Novembre 2014

## Slide 50
lavorazione del vassoio
Disegni tecnici esecutivi di Michele Costa
Lavorazione Meccanica: INFN Napoli – INFN Catania - LNS
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 51
Lavorazione della struttura interna
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 52
Dark box
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 53
Lavorazione della struttura interna
4 vassoi: durante il test è possibile inserire nei vassoi i PMT che verranno testati alla fine del
Test.
In questo modo prevediamo almeno 2 test al giorno per un totale di 124 PMT testati al giorno
60 giorni necessari per testare i 7500 PMT
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 54
Elettronica
white rabbit switch
CLBs, octopus boards (Long e short) e power board
Laser Controller
WR
CLB0
CLB1
OBL
OBL
OBS
OBS
Alle Base Boards
DB
PC
DAQ e DB
V. Kulikovskiy
G. Riccobene
C. Bozza
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)
LASER
Controller

## Slide 55
DAQ Test
Presentazione di V. Kulikovskiy
Elettronica
Sistema di raffreddamento per le CLB e Power Boards
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)

## Slide 56
Grazie!
KM3NeT – IT Meeting – Catania 28 Novembre 2014
Carlos Maximiliano Mollo
INFN (Naples)