## Slide 1: Calibration of Beam Position Monitors (BPMs)
Excellence in Detectors and Instrumentation Technologies
INFN - Laboratori Nazionali di Frascati, Italy
October 20-29, 2015
Yellow: Your Tasks T1 – T8

## Slide 2: Comparison of BPM Types
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015
BPM Type | Application | Precaution | Advantage | Disadvantage
Shoe-Box | p-synchrotrons heavy-ion accelerators | long bunches fRF < 10 MHz | very linear no x-y coupling sensitive for broad beams | complex mechanics capacitive coupling between plates
Button | p-linacs all e-accelerators | fRF > 10 MHz | simple mechanics | non-linear x-y coupling possible signal deformation
Stripline | colliders p-linacs all e-accelerators | best for β ≈ 1 short bunches | directivity „clean“ signal large signal | complex 50 Ω matching complex mechanics
Cavity | e-linacs (e.g. FELs) | short bunches, special applic. | very sensitive | very complex high frequency
P. Forck, “Lecture Notes on Beam Instrumentation and Diagnostics”, JUAS 2011

## Slide 3: Beam Position Monitors
 | Number | Beam Pipe | Length | Type | Single Bunch Resolution (RMS) | Train Averaged Resolution (RMS) | Optimum Resolution Range | Relaxed Resolution Range | x/y Crosstalk | Bunch to Bunch Crosstalk | Trans. Alignment Tolerance (RMS)
 |  | mm | mm |  | μm | µm | mm | mm | % | μm | μm
Standard BPM | 219 | 40.5 | 200/ 100 | Button | 50 | 10 | ± 3.0 | ± 10 | 1 | 10 | 200
Cold BPM | 102 | 78 | 170 | Button/ Re- entrant | 50 | 10 | ± 3.0 | ± 10 | 1 | 10 | 300
Cavity BPM Beam Transfer Line | 12 | 40.5 | 255 | Cavity | 10 | 1 | ± 1.0 | ± 2 | 1 | 1 | 200
Cavity BPM Undulator | 117 | 10 | 100 | Cavity | 1 | 0.1 | ± 0.5 | ± 2 | 1 | 0.1 | 50
IBFB | 4 | 40.5 | 255 | Cavity | 1 | 0.1 | ± 1.0 | ± 2 | 1 | 0.1 | 200
short version of E-XFEL BPM specification
specified charge range: 0.1 – 1nC
courtesy: D.Nölle (DESY)
different BPM types to meet different requirements
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015

## Slide 4: Beam Position Monitor
most common: capacitive pickups
signal generation via beam electric field
operation principle
electric field induces image charge on pick-up
→ pick-up mounted isolated inside vacuum chamber
→ amount of induced charge depends on distance between beam and pick-up
popular design: button-type pickup
→ simple, cheap, …
→ moderate resolution
LHC button pickup
courtesy: R.Jones (CERN)
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015
button pickup: high pass characteristics
P. Forck, “Lecture Notes on Beam Instrumentation and Diagnostics”, JUAS 2011
not well suited for long bunches
→ especially: low energy hadron beams, i.e. heavy ion beams
→ small coupling between pickup and bunch
Zt

## Slide 5: Electrostatic Beam Position Monitor
-
-
-
-
-
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-
-
+
+
+
+
+
-
-
+
-
+
-
+
-
+
-
-
-
-
+
+
+
-
+
+
-
+
-
-
+
-
+
-
-
-
-
-
+
+
+
-
+
+
-
-
+
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
Rhodri Jones –CERN Beam Instrumentation Group
T1) How QBeam(t) looks like?

## Slide 6
BPM Signal Generation
induced charge on BPM button
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015
-a
z = c∙t
BPM button
bunch
+a
0
Qind(t)
image current
button voltage (distorted by Impedance)
signal behind cable
…
E
Q(t) = bunch length
to be measured

## Slide 7
A = (d/2)2 p = p 5.42 mm2
r = Ø/2 = 30.16 mm
Ce= 3.07 pF
R0= 50W
Beampipe: round: Ø = 60.325 mm
Button Diameter d=10.8 mm
ZT  0.52 W
Area A
r
Let’s simulate: Beam.mat
T2): for w>>wL
Calculate ZT and fL
Zt(w) = UBPM(w) / IBeam(w)
T3) Plot Z(w) vs w at f=1, 5, 10, 20 GHz
see simulation Beam.mat
Transfer Impedance ZT
for w>>wL
for round beam pipes

## Slide 8
UBPM(t)
UBPM(w)
QBeam(t)
QBeam(w)
Beam.mat

## Slide 9: Position Reconstruction
Two common monitor geometries
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015
x
y
beam
linac-type
x = Kx
P1 – P3
P1 + P3
y = Ky
P2 – P4
P2 + P4
storage ring-type
x = Kx
(P1+P4) – (P2+P3)
P1+P2+P3+P4
y = Ky
(P1+P2) – (P3+P4)
P1+P2+P3+P4
x
y
beam
difference in position reconstruction
requires knowledge of monitor constant Kx, Ky
Position Information
r
α
→ rule of thumb (circular duct)
linac-type
storage ring-type
Beampipe: round: r = 30.17 mm
Button Diameter d =10.8 mm
sin(a) = 0.175
= 0.177
Kx=15.16 mm (linac)
Kx = 21.45 mm (ring)
T4) Calculate Kx,
for both types

## Slide 10: Electrostatic Pick-up – Button
Non-linear
  requires correction algorithm when beam is off-centre
  vacuum chamber not rotational symmetric (SR-source)
Area A
r
Rhodri Jones –CERN Beam Instrumentation Group
courtesy: A.Delfs (DESY)
beam
pump channel
button pickup (cut)
PETRA-III BPM close to ID

## Slide 11: Monitor Constant Calculation
arbitrary geometries
→ no simple rule of thumb calculation
strategy for calculation
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015
start with position raster for point beam
for each beam position (x0, y0)
→ calculate induced charge onto buttons
e.g. via boundary element method:
build position map
→ for each (x0, y0) assign Δx,y/Σ values
plot values along (x,y)-axes
sensitivity Sx,y
sensitivity: slope at origin
monitor constant: Kx,y = Sx,y-1
T5) Calculate Kx from simulation Hybrid.mat
and compare with “rule of thumb”
= Kx * D/S

## Slide 12
Input at x and z
out W = x+z, out Y = x-z
180 deg Hybrids

## Slide 13
-100% to 100% means the ± radius => x=10 => x= 30.17 mm
Kx = 1/S = 2 * 3.017 / 0.4 = 15.085
y=S*x, xfull scale = position variation, S=y/x

## Slide 14: OLYMPUS @ DORIS (DESY)
two-photon exchange in lepton scattering
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015
R. Milner et al., „The OLYMPUS experiment“, Nucl. Instrum. Methods A741 (2014) 1
compare e+p and e-p elastic scattering
OLYMPUS Target Chamber
BPM
BPM

## Slide 15: Monitor Constant Measurement
BPM test stand
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015
OLYMPUS BPMs
RF generator
499.66 MHz
ca. 33 dBm
test stand
→ vertical scanning device
→ 2 precision movers (micro screws) on top
→ wire antenna (Ø 0.2mm) centered in BPMs,
stretched by weight, damped movement
electrical input signal
→ 500 MHz cw signal induced on wire
→ electrical wire length: ¾ λ of 500 MHz
(standing wave)
→ essential: solid ground connections
signal readout
U. Schneekloth et al., Proc. IBIC 2014, Monterey (Ca), USA, (2014) 324

## Slide 16: BPM Test Stand
wire antenna
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015
weight
nylon chord
solder tag
RF generator
power amplifier

## Slide 17: Tasks: BPMs
calculate BPM signals using the Simulation Tool
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015
get a first impression about BPM signal forms
→ chamber geometry influence
→ non-linearities
→ output impedance
calculate monitor constants for OLYMPUS BPMs
use rule-of-thumb formulae for both geometries
measure OLYMPUS BPM monitor constants (both geometries)
T6) define electrical center of both BPM bodies (origin) and calibrate movers
T7) perform 1-dim. scan along one axis → max. wire position: ± 15 mm (!!!)
→ 7a) measure signal amplitudes from each button
→ 7b) calculate Δ/Σ from measured signals
→ 7c) plot Δ/Σ versus wire position
T8) determine monitor constant from slope at origin and compare with simulation and rule-of-thumb
(measure 2-dim. position map)
Done

## Slide 18: OLYMPUS Target Chamber BPMs
target chamber BPMs:
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015
round: Ø 60.325 mm button diameter 10.8 mm

## Slide 19: Processing System Families
Legend:
/ = Single channel
Wide Band
Narrow band
Normaliser
Processor
Active
Circuitry
Heterodyne
POS = (A-B)
Automatic
Gain Control
on S
Multiplexed
POS = [log(A/B)]
= [log(A)-log(B)]
Logarithmic
Amplifiers
Individual
Treatment
Amplitude
to Time
POS = [A/B]
POS = [ATN(A/B)]
Amplitude
to Phase
POS = D / S
Heterodyne
Hybrid
D / S
Electrodes
A, B
Direct
Digitisation
POS = D / S
DIGITISER
No turn by turn
turn by turn
Synchronous
Detection
Limiter,
Dt to Ampl.
.
Limiter,
f to Ampl.
Homodyne
Detection
Differential
Amplifier
Passive
Normalisation
Down
Conversion
Rhodri Jones –CERN Beam Instrumentation Group

## Slide 21: Modern BPM Read-out Electronics
Based on the individual treatment of the electrode signals
  Use of frequency domain signal processing techniques
    Developed for telecommunications market
  Rely on high frequency & high resolution analogue to digital converters
    Minimising analogue circuitry
    Frequency down-conversion used if necessary to adapt to ADC sampling rate
    All further processing carried out in the subsequent digital electronics
Rhodri Jones –CERN Beam Instrumentation Group

## Slide 22: Normalising the Position Reading
To make it independent of intensity
3 main methods:
  Difference/Sum : (VA – VB) / (VA + VB) = D / S
  Phase : Arctan(VA / VB)
  Logarithm : Log(VA) – Log(VB)
Rhodri Jones –CERN Beam Instrumentation Group

## Slide 23: Narrowband Signals: Cavity BPM
cavity BPM: collect directly position information
bunch excites several resonating modes while passing a
pillbox-like cavity
→ short bunches deliver wide spectrum of frequencies
monopole mode TM01(0): beam intensity
→ maximum at center
→ strong excitation
dipole mode TM11(0): beam position
→ minimum at center
→ excitation by beam offset
→ slightly shifted in frequency wrt. monopole mode
Linac: high resolution BPM for short bunches and single pulses
requirement: increase in BPM signal strength
standard BPMs: intensity signals which have to be subtracted to
obtain position information
→ difficult to do electronically without some of the intensity
information leaking through
task: antenna design to couple out dipole mode
amplitude: position information → only absolute value !
phase (wrt. monopole mode): sign information → simultaneous measurement required !
t / ns
U / V
antenna signal: time domain
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015

## Slide 24: Cavity BPM
f / GHz
U / V
TM01
TM11
TM02
U~q
U~q∙|r|
U~q
cavity frequency spectrum
problem: monopole mode (TM01) leakage into dipole mode (TM11)
→ suppression of monopole mode required
q: beam charge, r: beam offset
courtesy: D.Lipka (DESY)
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015

## Slide 25: Cavity BPM
suppresion of monopole mode
dipole mode (TM11) signal coupled out via waveguide
→ choose outcoupling at position of large TM11 electric field amplitude
design waveguide with cutoff frequency above f01 (monopole mode) resonance
Monopole Mode
Dipole Mode
courtesy: D.Lipka (DESY)
influence of outcoupling waveguide
narrow-band electronics for signal processing
→ B.Keil, Proc. DIPAC’09, Basel (Switzerland) 2009, TUOC01, p.275
→ D.Lipka, Proc. DIPAC’09, Basel (Switzerland) 2009, TUOC02, p.260
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015

## Slide 26: Today’s State of the Art BPMs
Prototype BPM for ILC Final Focus
  Required resolution of 2nm (yes nano!) in a 6×12mm diameter beam pipe
  Achieved World Record (so far!) resolution of 8.7nm at ATF2 (KEK, Japan)
Courtesy of D. Lipka,
DESY, Hamburg
Courtesy of D. Lipka & Y. Honda
Rhodri Jones –CERN Beam Instrumentation Group

## Slide 27
Reserve

## Slide 28: BPM Signal Generation
induced charge on BPM button
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015
-a
z = c∙t
BPM button, Ø2a
e-bunch
+a
0
Qind(t)
chamber geometry
beam charge & bunch length
image current
button voltage
signal behind cable
…
E

## Slide 29: BPM Signals
observation (1): signals are short with small modulation
single bunch response → nsec or sub-nsec pulse signals
beam position information → amplitude modulated on large (common mode) beam intensity signal!
courtesy: M. Gasior (CERN)
analog signal conditioning
signal sampling (ADC)
digital signal processing
data acquisition and control system interface
trigger, CLK & timing signals
provides calibration signals or
other drift compensation methods
RF device, EM field
detection, center of
charge
symmetrically
arranged electrodes
or resonant structure
BPM building blocks
BPM Pickup
control
system
(LAN)
feedback bus
(if applicable)
Analog Signal Conditioning
Digital Signal Processing
Data Acquisition
Trigger & Timing
Control
Power Supply & Misc.
position
data
timing, trigger signals
CAL
ADC
CLK
Read-out Electronics
courtesy: M. Wendt (CERN)
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015

## Slide 30: BPM Signals
observation (2): nonlinearities
synchrotron radiation emission
→ pickups mounted out of orbit plane
vacuum chamber not rotational symmetric
→ εhor >> εvert (SR emission in hor. plane)
→ injection oscillations due to off-axis injection
(allows intensity accumulation)
courtesy: A.Delfs (DESY)
beam
pump channel
button pickup (cut)
PETRA-III BPM close to ID
correction of strong non-linearities in beam position required
Gero Kube & Kay Wittenburg, DESY / MDI
EDIT 2015, Frascati (Italy), 23./27.October 2015
especially BPMs for circular e-accelerators

## Slide 31
For Button with
Capacitance Ce & Characteristic Impedance R0
Transfer Impedance:
Lower Corner Frequency:
Area A
r
Rhodri Jones –CERN Beam Instrumentation Group
Electrostatic Monitor – Beam Response

## Slide 32: Electrostatic Monitor – Beam Response
VB
V
R
C
Frequency (Hz)
Response UBPM (V)
0
0
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
Rhodri Jones –CERN Beam Instrumentation Group