## Slide 1: Algebraic reconstruction algorithms applied to proton computed tomography data
M. Bruzzi1,2, M. Brianzi2, M. Carpinelli3,9, G.A.P. Cirrone4, C. Civinini2, G. Cuttone4, D. Lo Presti5,8,G. Maccioni3, S. Pallotta2,6,7, N. Randazzo8, M. Scaringella2, F. Romano4, V. Sipala3,9, C. Stancampiano5,8, C. Talamonti2,6,7, E. Vanzi10
Prima – RDH Collaboration
1Physics and Astronomy Department, University of Florence, Florence, Italy 2INFN - Florence Division, Florence, Italy
3INFN Cagliari Division, Cagliari, Italy
4INFN - Laboratori Nazionali del Sud, Catania, Italy
5Physics and Astronomy Department, University of Catania, Catania, Italy 6Department of Biomedical, Experimental and Clinical Sciences, University of Florence, Florence, Italy
7SOD Fisica Medica, Azienda Ospedaliero-Universitaria Careggi, Firenze, Italy 8INFN - Catania Division, Catania, Italy
9Chemistry and Pharmacy Department, University of Sassari, Sassari, Italy 10 Fisica Sanitaria, Azienda Ospedaliero-Universitaria Senese, Siena, Italy
Società Italiana di Fisica – Congresso Nazionale 2015
Roma 22/09/2015

## Slide 2: Introduction
Proton Computed Tomography (pCT): apparatus and principle of operation
Algebraic Reconstruction Technique (ART)
Proton Most Likely Path (MLP)
    ART results from 62 MeV beam test (INFN-LNS Catana line)
Two pCT systems:
  5x5cm2
  5x20cm2
    Under construction

## Slide 3: PRIMA collaboration: small area pCT apparatus
Four x-y silicon microstrip based
tracking planes
Yag:Ce calorimeter
Proton entry and exit positions and directions
Proton residual energy
CATANA beam line:
62 MeV protons
used to treat ocular tumors

## Slide 4: Tomografic set-up

## Slide 5: Algebraic Reconstruction Techniques
Iterative algorithm to reconstruct tomographic images (proton stopping power maps) from ‘projections’ (for pCT set of single proton events)
Starting point (S(x,y,E) stopping power):
Introducing the mass stopping power S/r:
E0 = a fixed energy (200 MeV or ... 60 MeV in our case)

## Slide 6: Algebraic Reconstruction Techniques
Dividing by S/r at energy E:
The left hand side doesn’t depend too much on the material composition (~2-4*10-3) and could be replaced by the one measured for liquid water (NIST pstar tables - http://physics.nist.gov/PhysRefData/Star/Text/PSTAR.html):

## Slide 7: Algebraic Reconstruction Techniques
Integrating along the proton path:
Ein is given by the accelerator, Eout by the calorimeter and the ‘path’ by the tracker (Most Likely Path)
Subdividing the object into a set of pixels, for the ith proton:
Where wij is the path length of proton i inside the pixel j
Wang, Med.Phys. 37(8), 2010: 4138

## Slide 8
Pixel 1
wij
Pixel j
Pixel N
Computational challenge: find the simplest (fastest) way to build the wij matrix (could have billions of elements, most of them equal to zero)
p in
200
MeV
p out 90 MeV
Phantom: 20 cm of water

## Slide 9: Algebraic Reconstruction Techniques
The problem is then to solve, for Sj, the following set of equations:
N = number of pixels; M number of protons
In our case:
  N = (250x250)=62500 pixels
  M ~ 36(angles)*106 events

## Slide 10: Algebraic Reconstruction Techniques
The system could be solved using an iterative formula:
Sk image vector at iteration k (stopping power)
wi ith track length in each pixel (vector)  Tracker
pi stopping power integral (number)  Calorimeter
lk relaxing factor (constant value or 0 as ~k-1)
S0 initial image: {0} or approx (i.e., from FBP reconstruction).
Gordon, R; Bender, R; Herman, GT J. Theor. Biol. (1970) 29 (3): 471–81.
Sk+1= Sk + lk {(pi - <wi,Sk>) wi} / ǁwiǁ2

## Slide 11: Most Likely Path in a pCT geometry
MLP example with 200MeV
kinetic energy protons in
20cm of water:
Entry: Y(0) = 0.2cm
Y’(0) = -10mrad
Exit: Y(20) = -0.1cm
Y’(20) = +10mrad
Silicon microstrip detectors:
320mm thick
200mm strip pitch
MLP error envelope plus contributions from detector position measurement error (~ pitch/√12) and MCS inside the silicon sensors 
The sensor thickness contribution affects only the MLP error at the edge of the phantom
s ~ 150-250mm
200MeV in
90MeV out
Starting from D.C. Williams Phys. Med. Biol. 49 (2004) and R.W. Shulte at al. Med. Phys. 35 (11) (2008)
5 cm of air have been inserted in front and behind the 20cm H2O phantom

## Slide 12: ART images from 62 MeV data
cm
x100mm
ART reconstruction:
14 iteration starting from {0}
~ 106 events per angle
(36 angles)
4mm vertical slice selected
(2D only)
~4’ per iteration CPU time
x100mm
PMMA
phantom
1 cm

## Slide 13: PMMA mass stopping power (from NIST database)
S/r(E0=60MeV)=
10.5 MeVcm2/g
rPMMA~1.19 gcm-3
S(E0=60MeV)=
12.49 MeV/cm
Compatible with the ART reconstructed value
60 MeV

## Slide 14: ART images from 62 MeV data
Starting image = {0}, still room to improve
Resolution [mm]
Small hole
Large hole

## Slide 15: FBP used as seed for ART
FBP initial image 
ART after 14 iterations starting
from FBP seed
x100mm
x100mm
x100mm
x100mm
Vanzi E. et al., Nucl. Instr. and Meth. A 730 (2013) 184-190

## Slide 16: ART Resolutions
External edges
Inner holes:
resolution affected by
multiple scattering
Resolution [mm]
Two different edge
Positions

## Slide 17: Conclusions
The Prima/RDH pCT ‘proof-of-principle’ apparatus has been tested at 62 MeV (INFN-LNS, Catania) and 175 MeV (Svedberg Laboratory, Uppsala)
FBP and Algebraic algorithms have been used to reconstruct tomographic images taken at 62 MeV
The use of ART improves the spatial resolutions obtained with FBP
ART is able to handle arbitrary proton paths (most suitable for pCT analysis) but should be carefully implemented to keep the reconstruction time at a reasonable level
A larger field of view pCT apparatus is under construction

## Slide 18: pCT upgrade (5x20cm2)
A system similar to the one already tested
  Microstrip tracker
  YAG:Ce calorimeter
But with a 50 x 200 mm2 field of view
On-line data aquisition
1 MHz capability
Rectangular aspect ratio to perform tomographies in slices
Beam pipe
Tracker planes
Phantom
Calorimeter
Silicon sensors

## Slide 19: Fully assembled Tracker plane
Master FPGA
Virtex6
Silicon mstrip
Slave FPGAs
Chip front-end

## Slide 20: YAG:Ce calorimeter
2x7 YAG:Ce Crystals Array
Size: 3x3x10cm3
CHASSIS NI PXIe-1071
RT Controller
NI PXI-8102
FlexRIO
NI PXIe-7962R
Ad.Mod.
NI-5751
14 Analog Channels
Tracker
7Dig I/O GEN
Dig. Trigger
Disable Trigger
Silicon Photodiodes 1.8x1.8cm2
Fast Charge Amplifier + Shaper
x14
Data Acquisition System
Parallel read-out
Sampling: 5MS/s
24 Samples x event