## Slide 1: The Proton Computed Tomography Apparatus developed by INFN (RDH-WP3)
M. Bruzzi1,2, C. Civinini2, G. Maccioni3, S. Pallotta2,4,5, F. Paulis3, N. Randazzo6, M. Scaringella7, V. Sipala3,8, C. Talamonti2,4,5, E. Vanzi9
1Physics and Astronomy Department, University of Florence, Florence, Italy 2INFN - Florence Division, Florence, Italy
3INFN Cagliari Division, Cagliari, Italy 4Department of Biomedical, Experimental and Clinical Sciences, University of Florence, Florence, Italy
5 SOD Fisica Medica, Azienda Ospedaliero-Universitaria Careggi, Firenze, Italy 6INFN - Catania Division, Catania, Italy
7Information Engineering Department, University of Florence, Florence, Italy 8Chemistry and Pharmacy Department, University of Sassari, Sassari, Italy 9 Fisica Sanitaria, Azienda Ospedaliero-Universitaria Senese, Siena, Italy
RDH Meeting
Roma 24/04/2015

## Slide 2: Introduction
Two pCT systems: 5x5cm2; 5x20cm2
Analysis results from the small area system
  Algebraic Reconstruction Technique
Large area apparatus configuration
  Tracker (some considerations on MLP)
  Calorimeter
  Present status
Plans for 2015/2016

## Slide 3: pCT apparatus
PARAMETER | VALUE
Proton beam kinetic energy | ~200 MeV
Proton beam rate | 1 MHz
Spatial resolution | < 1 mm
Electronic density resolution | <1%
Detector radiation hardness | >1000 Gy
Dose per scan | < 5 cGy
P1
P2
P3
P4
z
x
y
Single particle proton tracking: silicon strip detectors → MLP
Residual energy measurement: crystal calorimeter → energy loss
A set of single event information can be processed by appropriate reconstruction algorithms (ART, FBP) to produce tomographic images.

## Slide 4: PRIMA collaboration: small area pCT apparatus
Four x-y silicon microstrip based
tracking planes
Yag:Ce calorimeter
Proton entry and exit positions and directions
Proton residual energy
First test at INFN-LNS:
May 2011
CATANA beam line:
62 MeV protons
used to treat ocular tumors

## Slide 5: Tomografic set-up

## Slide 6: Algebraic Reconstruction Techniques
Iterative algorithm to reconstruct tomographic images from projections (or single events)
Iterative formula:
xk image vector at iteration k (relative stopping power)
ai ith event track length in each pixel  Tracker (straigth line, cubic spline or most likely path (MLP) + rasterizing algorithms)
bi ith event stopping power integral  Calorimeter
lk relaxing factor (constant value or 0 as ~const-k)
x0 initial image: {0} or approx.
Gordon, R; Bender, R; Herman, GT J. Theor. Biol. (1970) 29 (3): 471–81.
xk+1= xk + lk {(bi - <ai,xk>)ai / ǁaiǁ2}

## Slide 7: ART images from 62 MeV data
cm
x100mm
ART reconstruction:
14 iteration starting from {0}
~ 106 events per angle
(36 angles)
4mm vertical slice selected
(2D only)
~4’ per iteration CPU time
No MLP is used  to be done
x100mm
PMMA
phantom
1 cm

## Slide 8: ART images from 62 MeV data
Starting image = {0}, still room to improve
Resolution [mm]

## Slide 9: FBP used as seed for ART
FBP initial image 
ART after 14 iterations starting
from FBP seed
x100mm
x100mm
x100mm
x100mm

## Slide 10: ART Resolutions
External edges
Inner holes:
resolution affected by
multiple scattering
Resolution [mm]
Small hole
Large hole
Two different edge
Positions

## Slide 11: Geant4 sim + ART spline reco
As an intermediate step toward the MLP we used a fit to a cubic function to approx. the proton path inside the phantom
The four parameters of the cubic are determined by the two track segments extrapolation point on the phantom envelope and their angles
The MLP takes better into account the multiple scattering but it is much more CPU-time consuming then a simple cubic fit
  Cubic: simple constant 4x4 matrix inversion
  MLP: point-to-point variable matrix algebra
Spline reco: 15’ pre-processing run + 1’ per iteration on a 62.5K pixel [200x200mm2] image  1 hour per image

## Slide 12: 180 MeV proton run simulation
With the 5x5cm2 prototype we are forced to use metals to allow for sufficient proton energy loss inside the small phantom:
45mm diameter Aluminum phantom with f=2-6mm iron or copper inserts
ART image after 50 iterations (spline proton trajectory approx.)
x200mm
x200mm

## Slide 13: 180 MeV proton run simulation
Phantom external edges (two different
positions):
Resolution after 50 iterations: 350-450mm
Phantom internal hole (two different
positions):
Resolution after 50 iterations: 570-630mm

## Slide 14: pCT upgrade (5x20cm2)
A system similar to the one already tested
  Microstrip tracker
  YAG:Ce calorimeter
But with a 50 x 200 mm2 field of view
On-line data aquisition
1 MHz capability
Rectangular aspect ratio to perform tomographies in slices
Beam pipe
Tracker planes
Phantom
Calorimeter
Silicon sensors

## Slide 15: Fully assembled Tracker plane
Master FPGA
Virtex6
Silicon mstrip
Slave FPGAs
Chip front-end

## Slide 16: Silicon microstrip detectors
Chart: Bias Voltage (V)
| 5.0 | 10.0 | 15.0 | 20.0 | 25.0 | 30.0 | 35.0 | 40.0 | 45.0 | 50.0 | 55.0 | 60.0 | 65.0 | 70.0 | 75.0 | 80.0 | 85.0 | 90.0 | 95.0 | 100.0 | 105.0 | 110.0 | 115.0 | 120.0 | 125.0 | 130.0 | 135.0 | 140.0 | 145.0 | 150.0 | 155.0 | 160.0 | 165.0 | 170.0 | 175.0 | 180.0 | 185.0 | 190.0 | 195.0 | 200.0
 | 16.679 | 22.76 | 28.8 | 34.8 | 38.78 | 41.4 | 43.65 | 45.63 | 47.42 | 49.06 | 50.51 | 51.89 | 53.12 | 54.27 | 55.3 | 56.25 | 57.13 | 57.96 | 58.76 | 59.54 | 60.36 | 61.05 | 61.78 | 62.5 | 63.17 | 63.86 | 64.53 | 65.15 | 65.8 | 66.41 | 67.02 | 67.58 | 68.18 | 68.72 | 69.26 | 69.82 | 70.34 | 70.87 | 71.4 | 71.89
 | 19.48 | 27.89 | 35.74 | 43.89 | 48.79 | 51.89 | 54.51 | 56.81 | 58.89 | 60.76 | 62.43 | 64.03 | 65.37 | 66.62 | 67.77 | 68.97 | 69.74 | 70.65 | 71.5 | 72.33 | 73.4 | 73.94 | 74.72 | 75.47 | 76.17 | 77.1 | 77.54 | 78.21 | 78.88 | 79.51 | 80.17 | 80.78 | 81.41 | 81.99 | 82.57 | 83.14 | 83.7 | 84.25 | 84.81 | 85.32
 | 15.89 | 21.96 | 27.65 | 33.58 | 37.29 | 39.75 | 41.79 | 43.57 | 45.21 | 46.66 | 48 | 49.25 | 50.36 | 51.39 | 52.36 | 53.25 | 54.07 | 54.86 | 55.64 | 56.36 | 57.13 | 57.82 | 58.51 | 59.19 | 59.81 | 60.45 | 61.07 | 61.69 | 62.31 | 62.89 | 63.47 | 64.02 | 64.58 | 65.14 | 65.64 | 66.18 | 66.66 | 67.19 | 67.7 | 68.21
 | 20.94 | 30.21 | 38.5 | 47.42 | 52.67 | 55.92 | 58.67 | 61.15 | 63.36 | 65.3 | 67.1 | 68.7 | 70.12 | 71.43 | 72.57 | 73.6 | 74.58 | 75.44 | 76.3 | 77.1 | 77.95 | 78.66 | 79.42 | 80.15 | 80.86 | 81.57 | 82.22 | 82.89 | 83.58 | 84.18 | 84.86 | 85.44 | 86.05 | 86.68 | 87.24 | 87.83 | 88.38 | 88.94 | 89.5 | 90.02
 | 16.047 | 22.3 | 28.18 | 34.3 | 38.27 | 40.74 | 42.91 | 44.9 | 46.61 | 48.14 | 49.81 | 50.88 | 52.1 | 53.16 | 54.21 | 55.12 | 55.99 | 56.82 | 57.64 | 58.38 | 59.21 | 59.9 | 60.61 | 61.33 | 61.95 | 62.62 | 63.25 | 63.88 | 64.52 | 65.11 | 65.7 | 66.26 | 66.87 | 67.4 | 67.93 | 68.5 | 69.02 | 69.53 | 70.07 | 70.56
 | 27.65 | 41.81 | 54.42 | 67.94 | 74.86 | 79.29 | 82.83 | 86.04 | 88.79 | 91.3 | 93.52 | 95.49 | 97.25 | 98.79 | 100.25 | 101.4 | 102.49 | 103.51 | 104.47 | 105.4 | 106.38 | 107.23 | 108.08 | 108.9 | 109.67 | 110.43 | 111.19 | 111.95 | 112.71 | 113.4 | 114.14 | 114.8 | 115.48 | 116.17 | 116.8 | 117.48 | 118.1 | 118.76 | 119.4 | 119.99
 | 21.09 | 30.7 | 39.32 | 48.63 | 53.67 | 57.18 | 60.02 | 62.52 | 64.76 | 66.74 | 68.54 | 70.16 | 71.63 | 72.94 | 74.11 | 75.16 | 76.09 | 77 | 77.87 | 79.01 | 79.58 | 80.51 | 81.11 | 81.86 | 82.56 | 83.29 | 83.97 | 84.69 | 85.37 | 86 | 86.65 | 87.24 | 87.87 | 88.47 | 89.05 | 89.65 | 90.21 | 90.9 | 91.4 | 91.93
 | 16 | 21.98 | 27.42 | 33.3 | 37.22 | 39.71 | 41.8 | 43.65 | 45.33 | 46.86 | 48.26 | 49.52 | 50.69 | 51.77 | 52.76 | 53.68 | 54.55 | 55.36 | 56.14 | 56.9 | 57.71 | 58.4 | 59.11 | 59.81 | 60.47 | 61.14 | 61.77 | 62.42 | 63.06 | 63.64 | 64.24 | 64.8 | 65.38 | 65.94 | 66.47 | 67.02 | 67.53 | 68.08 | 68.61 | 69.11
 | 16.895 | 23.53 | 30.16 | 36.92 | 41.42 | 44.28 | 46.72 | 48.82 | 50.74 | 52.46 | 54.06 | 55.49 | 56.81 | 58.07 | 59.16 | 60.2 | 61.13 | 62.06 | 62.92 | 63.75 | 64.64 | 65.35 | 66.12 | 66.86 | 67.54 | 68.26 | 68.89 | 69.59 | 70.25 | 70.87 | 71.52 | 72.12 | 72.74 | 73.35 | 73.91 | 74.51 | 75.05 | 75.63 | 76.22 | 76.75
 | 15.534 | 21 | 26.63 | 32.68 | 36.26 | 38.71 | 40.77 | 42.58 | 44.24 | 45.72 | 47.07 | 48.34 | 49.49 | 50.54 | 51.51 | 52.44 | 53.3 | 54.09 | 54.89 | 55.68 | 56.47 | 57.19 | 57.88 | 58.56 | 59.21 | 59.88 | 60.51 | 61.15 | 61.77 | 62.35 | 62.93 | 63.48 | 64.07 | 64.63 | 65.17 | 65.73 | 66.22 | 66.75 | 67.27 | 67.75
 | 16.125 | 21.7 | 27.2 | 32.99 | 36.91 | 39.42 | 41.8 | 43.42 | 45.11 | 46.62 | 48.03 | 49.3 | 50.47 | 51.57 | 52.57 | 53.48 | 54.35 | 55.19 | 55.96 | 56.73 | 57.55 | 58.24 | 58.95 | 59.66 | 60.3 | 60.97 | 61.59 | 62.23 | 62.87 | 63.47 | 64.06 | 64.6 | 65.17 | 65.74 | 66.26 | 66.79 | 67.32 | 67.9 | 68.38 | 68.89
 | 18.68 | 26.44 | 34.12 | 42.99 | 46.97 | 50.1 | 52.67 | 55.01 | 57.04 | 58.9 | 60.6 | 62.12 | 63.52 | 65.13 | 65.98 | 67.06 | 68.06 | 69 | 69.89 | 71.02 | 71.7 | 72.46 | 73.25 | 74.03 | 74.74 | 75.48 | 76.17 | 76.87 | 77.55 | 78.21 | 78.88 | 79.51 | 80.15 | 80.77 | 81.39 | 82.01 | 82.59 | 83.19 | 83.77 | 84.32
36 p on n single-sided silicon microstrip detectors (HPK):
  51.2 x 51.2 mm2 active area
  <100> crystal type
  320 mm thickness
  200 mm pitch
  256 channels
  I(Vfd)~2-3nA/cm2
No bad strips for all 36 detectors (9216 strips)
Chart: 
| 75.0 | 80.0 | 85.0 | 90.0 | 95.0 | 100.0
Vfd | 0 | 3 | 28 | 1 | 4 | 0
Full Depletion Voltage (V)
10 more silicon microstrip sensors have been ordered to FBK (already produced, now under test) to allow for spares leading to, eventually, a 5° tracker plane

## Slide 17: Read-out group
Front-end chips
FPGA slave
Silicon microstrip detectors
Power regulators

## Slide 18: Tracker DAQ architecture
Central acquisition
Virtex 6
ML605 board
Calorimeter DAQ
Trigger
enable
Trigger
7
gen
SS 1
SS 6
.
.
.
256
256
SM
serial link data/clk
200 Mbs
SSD 1
SSD 6
256
256
Front end 1
Front end 6
tracker plane
SS 1
SS 6
.
.
.
256
256
SM
serial link data/clk
200 Mbs
SSD 1
SSD 6
256
256
Front end 1
Front end 6
tracker plane
SS 1
SS 6
.
.
.
256
256
SM
serial link data/clk
200 Mbs
SSD 1
SSD 6
256
256
Front end 1
Front end 6
tracker plane
SS 1
SS 6
.
.
.
256
256
SM
6 readout links
200(x8) Mbs each
SSD 1
SSD 6
256
256
Front end 1
Front end 6
tracker plane
4/5 readout links 1.6 Gbs each
The front end of each detector will read-out by a Xilinx Spartan 6 FPGA (SS)
When a trigger occurs the read-out unit containing at least one hit will send the data to a central FPGA (SM)
The SM will asynchronously send the data to the central acquisition board
Trigger
gen
Trigger enable
7
1Gbs Ethernet
2GB DDR3 RAM

## Slide 19: Test results with high intensity 90Sr source
Source profile
Two not working strips:
one because of sensor
one because of front-end chip
(one ~hot)
# strip
# cluster
Number of clusters (electrons) per
event <Nclu>=4.2
2.5ms DAQ window  ~ 2 MHz of e-
Presently @350kHz max. trigger rate

## Slide 20: YAG:Ce calorimeter
2x7 YAG:Ce Crystals Array
Size: 3x3x10cm3
CHASSIS NI PXIe-1071
RT Controller
NI PXI-8102
FlexRIO
NI PXIe-7962R
Ad.Mod.
NI-5751
14 Analog Channels
Tracker
7Dig I/O GEN
Dig. Trigger
Disable Trigger
Silicon Photodiodes 1.8x1.8cm2
Fast Charge Amplifier + Shaper
x14
Data Acquisition System
Parallel read-out
Sampling: 5MS/s
24 Samples x event

## Slide 21: New YAG:Ce calorimeter data
ADC counts
60 MeV
Double protons=120 MeV
Noise = 10 MeV
s=3.9%

## Slide 22: 2015/2016 pCT plan
Analysis
  Implement MLP into ART (+3D reconstruction) and CPU time optimization: second half 2015
Small area system
  A 180 MeV run at Uppsala has been scheduled: 1st-5th June 2015  end of activities on the small area prototype
Extended view system
  Validate the tracker plane hardware at 60 MeV
    We have 16 hours of beam at LNS approved, we hope to have them during July 2015 (depends on Catana schedule)
  Production and assembly of the 4/5 final tracker planes (PCB and front-end chips) by the end of 2015, beginning of 2016
  Integration of the tracker and calorimeter read-out (second half of 2015)
  Data taking with the full system during 2016 (at least two runs)