## Slide 1: Preliminary results on the detection of Prompt Photons from He, C and O ions (HIT test beam of 2014)
G. Battistoni
G. Battistoni (INFN Milano), F. Collamati (Univ. Roma 1 and INFN-RM1), E. De Lucia (INFN-LNF), R. Faccini (Univ. Roma 1 and INFN-RM1), F. Ferroni (Univ. Roma 1 and INFN-RM1), M. P. Frallicciardi (Centro Fermi), M. Marafini (centro Fermi), I. Mattei (Univ. Roma 3 and INFN-LNF), S. Morganti (INFN-RM1), S. Muraro (INFN-MI), L. Piersanti (INFN-LNF), A. Rucinski (INFN-RM1), A. Russomando (Univ. Roma 1), A. Sarti (Univ. Roma 1 and INFN-LNF), A. Sciubba (Univ. Roma 1 and INFN-RM1), E. Solfaroli-Camillocci (IIT,INFN-RM1), M. Toppi (Univ. Roma 1 and INFN-LNF), G. Traini (Univ. Roma 1), C. Voena (INFN-RM1) and V. Patera (Univ. Roma 1, Centro Fermi and INFN-RM1

## Slide 2: Experimental Setup
θ = 0º
θ = 10º
θ = 30º
BGOa
BGOb
BGOc
STS2c
STS2b
STS2a
DCH
LYSO
LTS
PMMA
Rn
Rs
SC
STS1c
STS1b
STS1a
φ = 90º
SC = plastic scintillator; trigger for the DAQ
PMMA = phantom
Rn = 2 pixelated LYSO,side by side,1.6x5x5 cm3 each
Rs = 2 pixelated LYSO,side by side,1.6x5x5 cm3 each
(PET photons detectors)
LTS = plastic scintillator (charged particles TOF)
DCH = Drift Chamber (charged particles tracking)
LYSO = matrix of four LYSO crystals (prompt photons and charged particles detector)
STS1a,b,c = plastic scintillators
STS2a,b,c = plastic scintillators
(charged fragments TOF)
BGOa,b,c = BGO crystals (charged fragments detectors)
BEAM
DCH
LYSO
φ = 60º

## Slide 3: LYSO crystal(s)
Made of 4 different crystals
Calibrated with different radioactive sources and with hadron beams at HIT
An individual calibration curve obtained
for each crystal
Calibration curves inserted in simulation analysis together with energy resolution parameters

## Slide 4
1) raw Charge vs TOF
12C
Essential of Data Selection 1
Photons (+ some electrons)
Charged particles

## Slide 5
2) Slew time correction to Charge vs TOF
3) Background fitted and subtracted
4) Nphoton evaluated for each charge bin
Essential of Data Selection 2

## Slide 6: HELIUM BEAM

## Slide 7
HE102
HE102_60
HE125
HE125_60
HE145
HE145_60
FLUKA: v2013 vs v2015
FLUKA2013
FLUKA2015
photon energy at production

## Slide 8
He102
He102_60
He125
He125_60
He145_60
He145
FLUKA: v2013 & v2015
FLUKA2013
FLUKA2015
Measured Energy

## Slide 9
Measured Energy Spectra
Here normalized to peak at ~ 4 MeV
Photon spectra from He beams at the same energy and different angles are compatible in shape within the errors.

## Slide 10
 | DATA (10-6 Gamma(2-9MeV)/(Hecorr*epsDT) | FLUKA 2015 (10-6 Gamma(2-9MeV)/primary) | FLUKA 2013 (10-6 Gamma(2-9MeV)/primary)
He 90º 125 MeV/u | 4.70 ± 0.07 - 8.58 ± 0.14 | 4.8 ± 0.2 | 4.1 ± 0.1
He 90º 145 MeV/u | 5.92 ± 0.09 - 6.52 ± 0.10 | 6.4 ± 0.3 | 5.6 ± 0.1
He 60º 102 MeV/u | 1.44 ± 0.03 - 2.88 ± 0.04 | 4.0 ± 0.2 | 3.2 ± 0.1
He 60º 125 MeV/u | 3.36 ± 0.05 - 3.79 ± 0.05 | 4.3 ± 0.2 | 3.0 ± 0.1
He 60º 145 MeV/u | 5.00 ± 0.06 - 5.32 ± 0.07 | 5.8 ± 0.3 | 4.2 ± 0.1
RAW Yields
(Raw Yield)DATA = Ngamma(2-9MeV) / (Nprimarycorr*epsDT)
(Raw Yield)FLUKA = Ngamma(2-9MeV) / Nprimary
*in the 60º config, the PMMA-LYSO distance is larger wrt the 90º config

## Slide 11
RAW Yields
Ngamma/Nprimary
Ngamma/Nprimary
Ngamma/Nprimary
Ngamma/Nprimary
Ngamma/Nprimary
DATA
MC FLUKA
Absolute Comparison: N𝛾/primary

## Slide 12
 | DATA (10-2 Gamma(2-9MeV)/(Hecorr*epsDT*epsTOT) | FLUKA (10-2 Gamma(2-9MeV)/(primary*epsTOT))
He 90º 125 MeV/u | 5.7 ± 0.3 - 10.3 ± 0.5 | 5.7 ± 0.4
He 90º 145 MeV/u | 7.2 ± 0.3 - 7.9 ± 0.4 | 7.7 ± 0.5
He 60º 102 MeV/u | 1.68 ± 0.09 - 3.4 ± 0.2 | 4.7 ± 0.3
He 60º 125 MeV/u | 4.5 ± 0.2 - 5.0 ± 0.3 | 5.7 ± 0.4
He 60º 144 MeV/u | 6.7 ± 0.3 - 7.1 ± 0.3 | 7.7 ± 0.5
Yields
(after correction for Acceptance⨂Detection Efficiency)
(Yield)DATA = Ngamma(2-9MeV) / (Nprimarycorr*epsDT*epsTOT)
(Yield)FLUKA = Ngamma(2-9MeV) / (Nprimary*epsTOT)

## Slide 13: CARBON BEAM

## Slide 14
FLUKA: v2013 & v2015
FLUKA2013
FLUKA2015
photon energy at production
C120
C160
C180
C220

## Slide 15
FLUKA: v2013 & v2015
C120
C160
C180
C220
FLUKA2013
FLUKA2015
Measured Energy

## Slide 16
Measured Energy Spectra

## Slide 17
 | DATA (10-6 Gamma(2-9MeV)/(12Ccorr*epsDT) | FLUKA (10-6 Gamma(2-9MeV)/ primary)
12C 160 MeV/u | 6.3 ± 0.1 - 6.5 ± 0.1 | 8.6 ± 0.1
12C 181 MeV/u | 7.9 ± 0.1 - 8.2 ± 0.2 | 9.9 ± 0.1
12C 220 MeV/u | 6.7 ± 0.2 - 10.2 ± 0.2 | 10.8 ± 0.1
RAW Yields
(Raw Yield)DATA = Ngamma(2-9MeV) / (Nprimarycorr*epsDT)
(Raw Yield)FLUKA = Ngamma(2-9MeV) / Nprimary
In carbon runs (just 90º) the PMMA was always 10 cm long.

## Slide 18
RAW Yields
Ngamma/Nprimary
Ngamma/Nprimary
Ngamma/Nprimary
DATA
MC FLUKA
The gammas contribution at low energy is always larger in the simulation (v2013,v2015) with respect to data. We asked ourselves if we were actually subtracting from signal an higher amount of background, counting less gammas in that energy range, so we tried to change bkg model and so on. The bkg subtraction cannot explain such large signal/MC difference.
Absolute Comparison: N𝛾/primary

## Slide 19
 | DATA (10-2 Gamma(2-9MeV)/(12Ccorr*epsDT*epsTOT) | FLUKA (10-2 Gamma(2-9MeV)/(primary*epsTOT))
12C 160 MeV/u | 6.5 ± 0.3 - 6.8 ± 0.3 | 11.8 ± 0.7
12C 181 MeV/u | 9.7 ± 0.4 - 10.0 ± 0.5 | 13.6 ± 0.8
12C 220 MeV/u | 8.0 ± 0.3 - 12.1 ± 0.4 | 15.6 ± 0.6
In carbon runs (just 90º) the PMMA was always 10 cm long.
(Yield)DATA = Ngamma(2-9MeV) / (Nprimarycorr*epsDT*epsTOT)
(Yield)FLUKA = Ngamma(2-9MeV) / (Nprimary*epsTOT)
Yields
(after correction for Acceptance⨂Detection Efficiency)

## Slide 20: OXYGEN BEAM

## Slide 21
FLUKA: v2013 & v2015
FLUKA2013
FLUKA2015
O300
O260
O210
O210_60
photon energy at production

## Slide 22
FLUKA: v2013 & v2015
FLUKA2013
FLUKA2015
O210
O260
O300
O210_60
Measured Energy

## Slide 23
Measured Energy Spectra
Photon spectra from 16O at the same energy beams and different angles have different shapes.
normalized to peak at ~ 4 MeV

## Slide 24
Explanation of change of shape at 60º
There is a correlation of the energy of photons exiting the PMMA, produced by projectiles (and their fast fragments) de-excitations, with the emission angle for CosTheta ≳ 0.2/0.4 (Theta ≲ 78º) => energy boost (from a dedicated simulation: 4He/12C/16O beam impinging on a PMMA target)
The more intense gamma lines are due to target nuclei de-excitations (11C~2 MeV;12C~4.4 MeV; 11C,15O:~5 MeV; 16O:~6.13 MeV ) which are instead not affected by boost (a part a minor effect due to nucleus recoil)
4He 145 MeV/u
no projectile fragments
=> no effect
Egam
CosTheta
12C 220 MeV/u
60º
CosTheta
Egam
16O 300 MeV/u
60º
CosTheta
Egam
The boost effect on prompt photons energy for the 12C appears around 2-4 MeV; less evident than the effect due to the 16O

## Slide 25
As a result, the theory predicts an uniform angular distribution for the 4He beam (which has no gamma de-excitation lines), but not for the 12C and 16O beams
4He 145 MeV/u
12C 220 MeV/u
16O 300 MeV/u
CosTheta
CosTheta
CosTheta
Explanation of change of shape at 60º

## Slide 26
Is the shift of the peak around ~4 MeV in the Ox spectra at 60º due to the energy boost of photons emitted by *C fragments (gamma peak ~ 4.3/4.4MeV->~5MeV) and 15O nuclei (gamma peak ~ 5.2MeV->~6MeV) coming from the 16O beam and mixing with the ~4 MeV peak coming from photons emitted by *C target fragments?
16O 300 MeV/u
*C
15O
Explanation of change of shape at 60º

## Slide 27
Photon spectrum 16O
Etrue of gammas produced in PMMA

## Slide 28
 | DATA (10-6 Gamma(2-9MeV)/(Oxcorr*epsDT) | FLUKA (10-6Gamma(2-9MeV)/primary)
Ox 90º 210 MeV/u | 8.8 ± 0.1 - 9.9 ± 0.3 | 15.2 ± 0.6
Ox 90º 260 MeV/u | 12.0 ± 0.2 - 12.7 ± 0.2 | 19.9 ± 0.6
Ox 90º 300 MeV/u | 16.5 ± 0.3 - 19.7 ± 1.1 | 21.0 ± 0.7
Ox 60º 210 MeV/u | 10.7 ± 0.2 - 11.5 ± 0.2 | 17.4 ± 0.6
Ox 60º 260 MeV/u | 14.4 ± 0.2 - 15.4 ± 0.3 | 24.2 ± 0.8
Ox 60º 300 MeV/u | 17.9 ± 0.3 - 20.3 ± 0.3 | 30.5 ± 0.8
RAW Yields
(Raw Yield)DATA = Ngamma(2-9MeV) / (Nprimarycorr*epsDT)
(Raw Yield)FLUKA = Ngamma(2-9MeV) / Nprimary
*in the 60º config, the PMMA-LYSO distance is larger wrt the 90º config

## Slide 29
RAW Yields
DATA
FLUKA2015
Ngamma/Nprimary
Ngamma/Nprimary
Ngamma/Nprimary
Ngamma/Nprimary
Ngamma/Nprimary
Ngamma/Nprimary
Absolute Comparison: N𝛾/primary

## Slide 30
RAW Yields
DATA
MC FLUKA
Ngamma/Nprimary
Ngamma/Nprimary
Ngamma/Nprimary
Ngamma/Nprimary
As for the C ions case, the bkg subtraction cannot explain such large signal/MC difference.
Absolute Comparison: N𝛾/primary
Ngamma/Nprimary
Strange: are we really sure?

## Slide 31
 | DATA (10-2 Gamma(2-9MeV)/(Oxcorr*epsDT*epsTOT) | FLUKA 2015 (10-2 Gamma(2-9MeV)/(primary*epsTOT))
Ox 90º 210 MeV/u | 10.5 ± 0.4 - 11.85 ± 0.56 | 18.12 ± 0.93
Ox 90º 260 MeV/u | 13.95 ± 0.51 - 14.78 ± 0.53 | 23.19 ± 1.04
Ox 90º 300 MeV/u | 23.74 ± 0.83 - 26.89 ± 0.96 | 30.26 ± 1.32
Ox 60º 210 MeV/u | 11.16 ± 0.42 - 11.97 ± 0.44 | 18.12 ± 0.87
Ox 60º 260 MeV/u | 13.77 ± 0.45 - 14.77 ± 0.49 | 23.19 ± 0.94
Ox 60º 300 MeV/u | 17.72 ± 0.53 - 20.14 ± 0.60 | 30.26 ± 1.10
Yields
(Yield)DATA = Ngamma(2-9MeV) / (Nprimarycorr*epsDT*epsTOT)
(Yield)FLUKA = Ngamma(2-9MeV) / (Nprimary*epsTOT)
Yields
(after correction for Acceptance⨂Detection Efficiency)

## Slide 32: Some conclusions
We are close to analysis completion
Importance for INSIDE: these results, if confirmed,
demonstrate the capability of predicting prompt photon
fluxes in the irradiation performed with ion therapeutic
beams
There remains a suspicious mismatch in the range of
detected energy 2-4 MeV in part of Carbon cases and for
Oxygen. There could be different explanations
(the MC model, effects due to data selection,…) under
investigation