## Slide 1: Development of μ-PIC with resistive electrodes using sputtered carbon
Kobe Univ.,Tokyo ICEPPA
F.Yamane,
A.Ochi, Y.Homma S.Yamauchi, N.Nagasaka, H.Hasegawa, T.KawamotoA, Y.KataokaA, T.MasubuchiA,
Outline
Introduction
New resistive material
Pixel alignment
Gain measurement
Summery and further prospects

## Slide 2
Introduction
New resistive material
Pixel alignment
Gain measurement
Summery and further prospects

## Slide 3
2D gaseous imaging detector produced by PCB technology.
For many purpose...
Mascot of NEWAGE
"Daakumatan"
Dark Matter Search(NEWAGE)
K.Nakamura
PTEP (2015) 043F01
Talked by T.Ikeda(15/10)
Micro Pixel Chamber:μ-PIC
ETCC(Electron-Tracking Compton Camera)
T.Tanimori+
The Astrophysical Journal 810 (2015)
Talked by T.Takemura(14/10)
Neutron Imaging
J.D.Parker+
NIM A697 (2013) 23
A. Ochi+ NIM A471 (2001) 264
Space Dosimeter(PS-TEPC)
Y.Kishimoto
NIM A732 (2013) 591
and more application…

## Slide 4
Stable operation in high rate HIPs(Highly Ionizing Particles) environment
  ->μ-PIC needs to have spark tolerant
In this research
  We propose μ-PIC for ATLAS forward muon detector.
  Sputtered carbon is used as resistive cathodes for spark protection.
  The production improvement and fundamental measurement of resistive μ-PIC are reported.
Our purpose

## Slide 5
Physics motivation
Muon Tagger
  ATLAS new endcap forward muon detector considered to be installed nearby the beam line (2.7<|η|<4.0 ) after the long shutdown from 2023.
Very hard environment for detectors
  Multiple track are incident in very small area.
  High rate HIPs background ~100kHz/cm2 .
  Detector size is limited (<5cm thickness).
Requirements for Muon Tagger
  Granularity of ~100um for track separation.
  Stable operation with high gain in high rate HIPs.
  Thickness<5cm

## Slide 6: Why μ-PIC ?
  Properties
  Good position resolution(2D) & high rate capability.
  Isolated pixels are arranged by 400um pitch with 250um diameter hole.
  Each pixel has an anode pin and a surrounding cathode.
  2D readout by orthogonal anode/cathode strips without the loss of signals.
μ-PIC has no floating structures (wire, foil, mesh…).
  Made by PCB/FPC technology.
  Large detector(30cm) can be produced.
  Larger area available by arranging many μ-PIC to tile form.
    (It is possible because there is no flowing structure)
μ-PIC with resistive electrodes was developed for spark protection. (Next slide)
μ-PIC
μ-PIC
μ-PIC
μ-PIC
μ-PIC
μ-PIC
μ-PIC
μ-PIC
μ-PIC
10~
30cm

## Slide 7
μ-PIC with resistive cathode
Resistive cathodes for reducing sparks
  Strong spark reduction was shown at high gain(>10000) operation under irradiation of the fast neutron(a few MeV).
  Spark rate was 104 times less than normal μ-PIC
    Spark rate = Spark counts / Number of neutron
Capacitive readout without AC coupling
  Pickup electrodes are lying under resistive cathodes and insulator(polyimide).
  Charges are induced from resistive cathodes.
  This capacitive readout can be applied for anodes.
  Detector construction can be simplified.

## Slide 8
Requirements
  Granularity of ~100um
    The pixel pitch should be reduced (400um->100um).
      ->Resistive ccathodes are needed to be able to form precise pattern.
  Stable operation
    High resistivity: Strong tolerance for sparks but continuous voltage drop distort electric field
    Low resistivity: Sparks cannot be reduced.
    It is important that resistivity must not be too high and too low.
      ->Fine resistivity control is needed.
  Detector size
    No floating structure & capacitive readout -> Possible!
Previous resistive material: Carbon polyimide
  Cathode surface is not flat so much.
  Difficult to form more precise pattern.
  Fine resistivity control is difficult.
Other resistive material is needed to achieve requirements.
300um
Requirements

## Slide 9
Introduction
New resistive material
Pixel alignment
Gain measurement
Summery and further prospects

## Slide 10: New material for resistive electrodes
Sputtered carbon has been developed and studied in Kobe Univ. (2013~).
  Diamond like carbon is formed on the substrate
  Very precise pattern can be formed easily with lift off process
3D image of resistive μ-PIC
Left: Carbon polyimide, Right: Sputtered carbon
300um

## Slide 11
Fine and uniform pattern.
  Enough to achieve granularity of ~100um.
Wide range of resistivity control is available (50kΩ/sq.~3000MΩ/sq.).
  Thickness control
  Nitrogen doping
Uniform resistivity.
Strong attachment on substrate.
Large size available (>2m)
Chart: 
| 300 | 500 | 700 | 1000 | 1200 | 1800 | 2000 | 2400 | 3600 | 500 | 700 | 1000 | 2400
DLC | 3338.4 | 154.44 | 93.3 | 5.361 | 2.75 | 2.43 | 2.03 | 1.25 | 0.4914
N-DLC | 0.665 | 0.706 | 0.435 | 0.055
Surface Resistivity[MΩ/sq.]
Thickness [Ǻ]
N2 content in Ar is 3.2%
40min.
3 hours
Resistivity vs thickness
Pure C
Be-Sputter Co. Ltd.
(Kyoto Japan)
Properties of the sputtered carbon
Vacuum chamber
(with Ar + N2 gas)
Rotating drum
4.5 m round
Sample
Sputtering
target
A resistive strips foil for ATLAS NSW

## Slide 12
Introduction
New resistive material
Pixel alignment
Gain measurement
Summery and further prospects

## Slide 13
Anode pins are formed by 2 different processes on “Top substrate” and “Under substrate” (next slide).
Sometimes, there are misalignments of pixels...
  ->Anode pins contact to the inner pickup electrodes!
Alignment process should be improved!
Alignment problem
Misalignment of anode
Top substrate
Under substrate
Top anode pin
Misaligned under anode pin

## Slide 14: Previous production method
Top substrate: Polyimide 25um
Cathode patterning using double side mask
Cu plating on top surface
Etching substrate
Plating for anode pin
25um
Top anode pin
Manufactured by Raytech Inc.
Cathode pattern
Cathode pattern
Pickup electrode
Pickup electrode

## Slide 15: Previous production method
Top substrate: Polyimide 25um
Cathode patterning using double side mask
Cu plating on top surface
Etching substrate
Plating for anode pin
Under substrate: Polyimide
  Because the surface is covered by Cu that will become anode strips, the anode pin of top surface cannot be seen…
Etching substrate and plating for anode pin -> Misalignment!
25um
Top anode pin
Manufactured by Raytech Inc.
Cathode pattern
Cathode pattern
Pickup electrode
Pickup electrode

## Slide 16: Improved production
75um
Top anode pin
Under substrate: Transparent dry resist 75um
Exposure
25um
Manufactured by Raytech Inc.
Cathode pattern
Cathode pattern
Pickup electrode
Pickup electrode
Dry resist

## Slide 17: Improved production
75um
Top anode pin
Under substrate: Transparent dry resist 75um
Exposure
Developing
Cu sputtering for anode connection
Ni plating
25um
Manufactured by Raytech Inc.
Cathode pattern
Cathode pattern
Pickup electrode
Pickup electrode
Dry resist
Cu sputtering for
anode connection

## Slide 18: Improved production
Under substrate: Transparent dry resist 75um
Photo etching for anode pin.
Cu sputtering for anode connection
Ni plating
Carbon sputtering with lift off
25um
Top anode pin
75um
Cathode: Sputtered carbon
Pickup electrode
Pickup electrode
Cathode: Sputtered carbon

## Slide 19
Readout pitch:400um
Active area:10cm×10cm, 256strips
Pixels are well aligned in all region.
Carbon sputtered cathodes are well formed.
μ-PIC(RC33)
250um

## Slide 20
Introduction
New resistive material
Pixel alignment
Gain measurement
Summery and further prospects

## Slide 21: Gain curve
Gas gain measurement using 55Fe source
  Gas mixture: Ar:C2H6 = 90:10
  Drift field: 2kV/cm
  Cathode HV: -500V ~ -700V
  Anode: Ground
  Readout: Cathode pickup electrode
  Preamp: ASD
Gain of more than 10000 was achieved.
There were no discharges.
However, there are some problems…
  Signals cannot be read by Anode.
  Gaseous amplification was not observed in Anode HV & Cathode ground operation (conventional way).
  Compared to previous detector,
    - Operation HV is higher than about 100V.
    - Gain increases slowly with HV value.
  Something is wrong!
Cathode signal
Anode signals cannot be seen
Chart: Gain curve
| 500 | 550 | 600 | 650 | 700
RC33-1 | 2942.6399 | 4097.3622 | 6402.5771 | 7218.0733 | 7856.7658
RC33-3 | 1426.2738 | 2008.7106 | 4588.0134 | 4787.6577 | 8871.0604 | 16168.2287
Gain of old μ-PIC
with carbon polyimide
FWHM:19%

## Slide 22
0
The cause of problems
Disconnection between anode pins and anode strips.
  It is thought that anode connection by Cu sputtering was incomplete because under substrate was thick (75um).
  Cathode –HV & Anode Ground operation only available.
Gain reduction was seen because of charge up on anode pins.
-HV
-HV
Chart: Gain variation
| 450 | 500 | 550 | 600 | 650 | 700
Within 10min after source on | 1426.2738 | 2008.7106 | 4588.0134 | 4787.6577 | 8871.0604 | 16168.2287
More than 2 hours after source on | 629.4942 | 1161.9143 | 1856.4396 | 2827.5906 | 4870.5609 | 6539.621 | 7306.8981 | 8673.108 | 9138.3807 | 9193.3675
Electrons remain on anode pin

## Slide 23
Introduction
New resistive material
Production
Gain measurement
Summery and further prospects

## Slide 24
Summary and further prospects
Summary
μ-PIC with resistive cathode using sputtered carbon is developed.
Thanks to production improvement, pixels were well aligned in all active region(10cm×10cm).
Gain of more than 10000 was achieved with no discharge. (Ar:C2H6=90:10)
Connection between the anode strip and the anode pin is incomplete.
Further prospects
  Anode connection should be improved (now producing).
Spark torrent test of carbon sputtered μ-PIC (3/2016).
E-field simulation.
Reducing pixel pitch. 400um -> 200um
True 2D readout with SRS system.

## Slide 25
Thank you!